## SN029
Spray Air Calibration Log |  |  | If new aircal is +/- 1.5% from previous trend line then re-do aircal
[12 empty] | Optode reading (average readings over 1 minute after stable) |  |  |  | Std dev |  |  |  | Barometer Reading |  |  | Calculations
Mission | Pre/Post | Date | Datenum | Time | Operator | Spray ID | Optode S/N | pH S/N | K2 from Data Sheet | Test Tank Date | Test Tank K0 | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Atmospheric Pressure | Relative Humidity | Temp | pH2O @ 100% hum | pH2O real | T_O2soly | O2_soly (umol/kg) | Density water (kg/L) | O2_soly (uM) | O2raw,air (uM) | Gain | QF | QC_gain | Comments | Gain used
19402901 | [31 empty] | 103.4% | No air cal - copied 19702901 | 103.3%
19502901 | [31 empty] | 103.4% | No air cal - copied 19702901 | 103.3%
19502902 | [31 empty] | 103.4% | No air cal - copied 19702901 | 103.3%
19702901 | Pre | 2019-07-15 |  | 13:00 | YT/BJ/KC | 29 | [5 empty] | 17.0622 | 0.62148 | 6.083 | 21.3051 | 0.0030732 | 6.2149e-05 | 0.0027839 | 0.003173 | 1017 | 0.57 | 22.6 | 25.4046 | 14.4806 | -0.061669 | 277.255 | 0.99792 | 276.679 | 271.6643 | 103.4% | 2 | 103.4% |  | 103.3%
19702901 | Post | 2019-09-16 |  | 15:15 | JKW/BJ | 29 | [5 empty] | 17.234 |  | 6.2587 | 19.7525 | 0.001 |  | 0.00057735 | 0.00095394 | 1016 | 0.59 | 20.6 | 23.0861 | 13.6208 | -0.05079 | 285.7992 | 0.99825 | 285.2995 | 279.5095 | 103.3% | 2 | 103.3%
19A02901 | Pre | 2019-09-30 |  | 15:00 | JKW/BJ | 29 | [5 empty] | 17.2012 |  | 6.3148 | 19.6227 | 0.0020494 |  | 0.0021679 | 0.0033162 | 1016 | 0.55 | 20.5 | 22.9009 | 12.5955 | -0.04988 | 286.5356 | 0.99828 | 286.0422 | 282.0164 | 102.8% | 2 | 102.8% |  | 102.8%
19C02902 | Pre | 2019-12-12 |  | 13:30 | JKW | 29 | [5 empty] | 17.167 |  | 6.1347 | 20.6139 | 0.0021602 |  | 0.0016997 | 0.00020548 | 1026 | 0.63 | 20.8 | 24.3484 | 15.3395 | -0.056825 | 281 | 0.99807 | 280.4583 | 273.9718 | 104.6% | 2 | 104.6% |  | 104.6%
20202901 | Pre | 2020-02-10 |  | 1430 | BDJ | 29 | [5 empty] | 17.5078 |  | 6.7496 | 16.4952 | 0.000748 |  | 0.002245 | 0.010491 | 1016 | 0.41 | 18.3 | 18.8156 | 7.7144 | -0.027975 | 305.3888 | 0.99886 | 305.04 | 301.4344 | 102.6% | 2 | 102.6% |  | 102.6%
20702901 | PRE (100% Humidity) | 2020-07-21 |  | 934 | JKW/BDJ | 29 |  |  | -0.001026 | 2020-07-22 | -1.393 | 17.3452 |  | 6.1148 | 19.7722 | 0.0021354 |  | 0.0017205 | 0.00028705 | 1015 | 1 | 19.6 | 23.1144 | 23.1144 | -0.050928 | 285.6875 | 0.99825 | 285.1869 | 273.0845 | 104.6% | 2 | 104.6% |  | 104.6%
20702901 | POST (100% Humidity) | 2020-09-17 |  | 2037 | JKW | 29 |  |  | -0.001026 | 2020-09-17 |  | 17.6517 |  | 6.1555 | 18.5588 | 0.002586 |  | 0.0020616 | 0.0026875 | 1016 | 1 | 21 | 21.4318 | 21.4318 | -0.042428 | 292.7038 | 0.99849 | 292.261 | 274.9022 | 106.6% | 4 | 104.5% | Calculated based on fit of y = 0.00003235*x-0.3815 (from previous deployments)
20A02901 | PRE (100% Humidity) | 2020-10-06 |  | 1745 | JKW | 29 |  |  | -0.001026 | 2020-10-06 |  | 17.7036 |  | 6.154 | 18.3631 | 0.0026533 |  | 0.0018974 | 0.0063294 | 1017 | 1 | 20.3 | 21.1706 | 21.1706 | -0.041057 | 293.8652 | 0.99852 | 293.4316 | 274.8352 | 107.2% | 4 | 104.5% | Calculated based on fit of y = 0.00003235*x-0.3815 (from previous deployments) | 104.6%
20A02901 | POST (100% Humidity) | 2020-10-22 |  | 1445 | BDJ | 29 |  |  | -0.001026 | N/A |  | 17.292 |  | 5.932 | 21.0978 | 0.005099 |  | 0.0034641 | 0.0081097 | 1011 | 1 | 21.1 | 25.0838 | 25.0838 | -0.060216 | 278.3683 | 0.99797 | 277.8026 | 264.9207 | 104.6% | 2 | 104.6%
20C02901 |  | 2020-12-03 | [29 empty] | 104.6% | Calculated based on fit of y = 0.00003287x - 0.40459442 (from previous deployments) | 104.6%
21202901 | PRE (100% Humidity) | 2021-02-08 |  | 1100 | BDJ | 29 |  |  | -0.001026 | 2021-02-09 |  | 17.2268 |  | 5.9423 | 21.3182 | 0.0014625 |  | 0.0014907 | 0.0025184 | 1017 | 1 | 21 | 25.425 | 25.425 | -0.061761 | 277.1848 | 0.99792 | 276.6082 | 265.3822 | 104.6% | 2 | 104.6% |  | 104.6%
22602901 | Pre (100% Humidity) | 2022-06-07 |  | 930 | JKW/JSL | 29 | [5 empty] | 17.5227 | 17.0618 | 6.363 | 17.0618 | 0.002571 | 0.0057986 | 0.0018974 | 0.0055011 | 1019.5 | 1 | 18.65 | 19.5044 | 19.5044 | -0.031942 | 301.8073 | 0.99876 | 301.4333 | 284.169 | 106.7% | 2 | 106.7% |  | 106.7%
22902901 | Pre (100% Humidity) | 2022-09-01 |  | 1405 | JKW/JSL | 29 |  |  | -0.0010296 | 2022-09-01 | -1.4053 | 16.8566 |  | 6.0006 | 21.424 | 0.005023 |  | 0.0044175 | 0.00050169 | 1014 | 1 | 22 | 25.5903 | 25.5903 | -0.062502 | 276.6198 | 0.9979 | 276.0378 | 267.9863 | 103.1% | 4 |  | Calculated based on fit | 106.7%
23202901 | Pre (100% Humidity) | 2023-02-13 |  | 1240 | JKW | 29 | [5 empty] | 17.4801 | 231.8333 | 6.1718 | 18.0955 | 0.0035903 | 112.9287 | 0.0025219 | 0.0095974 | 1014 | 1 | 18.4 | 20.8182 | 20.8182 | -0.039182 | 295.4669 | 0.99857 | 295.0458 | 275.6301 | 107.1% | 2 |  |  | 107.1%
23202901 | POST (100% Humidity) | 2023-03-02 |  |  | JKW | 29 | [5 empty] | 17.4468 | 233 | 6.2095 | 18.0493 | 0.002681 | 103.958 | 0.0012247 | 0.011972 | 1020 | 1 | 18.7 | 20.7578 | 20.7578 | -0.038859 | 295.7453 | 0.99858 | 295.3263 | 277.3138 | 107.2% | 2 |  |  | 107.2%
23702901 | PRE (100% Humidity) | 2023-07-05 |  |  | JKW | 29 | [5 empty] | 17.4658 | 234.6667 | 6.1154 | 18.4143 | 0.0030103 | 111.4266 | 0.0022167 | 0.002502 | 1023 | 1 | 20.9 | 21.2387 | 21.2387 | -0.041416 | 293.5605 | 0.99851 | 293.1245 | 273.1133 | 108.4% | 2 |  |  | 108.4%
23C02901 | PRE (100% Humidity) | 2023-12-05 |  | 10:30 | JKW/JSL | 29 | [5 empty] | 17.4328 | 236.3333 | 6.1566 | 18.3498 | 0.0025612 | 115.136 | 0.0017436 | 0.0095891 | 1021 | 1 | 18.3 | 21.153 | 21.153 | -0.040964 | 293.9446 | 0.99853 | 293.5116 | 274.9513 | 107.6% | 2 |  |  | 107.6%
24802901 | PRE (100% Humidity) | 2024-08-07 |  | 11:00 | BW/JKW | 29 | [5 empty] | 17.4294 | 19.0775 | 6.0062 | 19.1993 | 0.0029889 | 0.0011643 | 0.0022509 | 0.0054929 | 1016 | 1 | 19.9 | 22.306 | 22.306 | -0.046914 | 288.9616 | 0.99836 | 288.4885 | 268.2345 | 1.0785 |  | 107.8% |  | 107.8%
25202901 | PRE (100% Humidity) | 2025-02-04 |  | 9:30 | BW/JKW | 29 | [5 empty] | 17.908 | 15.4276 | 6.4073 | 16.6402 | 0.0023094 | 0.0039178 | 0.0025841 | 0.0091302 | 1007 | 1 | 17.8 | 18.9898 | 18.9898 | -0.02899 | 304.4651 | 0.99883 | 304.1099 | 286.1475 | 1.0561 |  | 105.6% |  | 105.6%
|  | 2026-07-20 |  | 4:30 | BW/BD/JKW | 29 | [5 empty] | 16.9957 | 22.4794 | 5.6835 | 19.7803 | 0.0035606 | 0.015997 | 0.003504 | 0.0071163 | 1014 | 1 | 20.1 | 23.126 | 23.126 | -0.050985 | 285.642 | 0.99825 | 285.1409 | 253.8228 | 1.1242 |  | 112.4% |  | 112.4%

## SN034
Spray Air Calibration Log |  | If new aircal is +/- 1.5% from previous trend line then re-do aircal
Test tank Winkler based gains
[11 empty] | Optode reading (average readings over 1 minute after stable) |  |  |  | Std dev |  |  |  | Barometer Reading |  |  | Calculations
Mission | Pre/Post | Date | Time | Operator | Spray ID | Optode S/N | pH S/N | K2 from Data Sheet | Test Tank Date | Test Tank K0 | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Atmospheric Pressure | Relative Humidity | Temp | pH2O @ 100% hum | pH2O real | T_O2soly | O2_soly (umol/kg) | Density water (kg/L) | O2_soly (uM) | O2raw,air (uM) | Gain | Gain QF | QC'd Gain | Comments | Gain Used | Temperature (CTD) | Salinity | O2(mL/L) | Temperature (optode) | phase | Tscaled | Scorr | O2(uM) Glider Salinity corrected | Winkler1 (uM) | Winkler2 | metadatas | Gain1 | Gain2 | Avg Gain
20703401 | PRE (100% Humidity) | 2020-06-29 | 15:10 | JKW | 34 |  | SBE_10330 | -0.0010296 |  | N/A | 16.7875 |  | 6.1235 | 21.1047 | 0.002062 |  | 0.002062 | 0.003401 | 1015 | 1 | 21.1 | 25.0944 | 25.0944 | -0.060265 | 278.3311 | 0.99797 | 277.765 | 273.4731 | 101.7% | 4 | 104.8% | Calculated based on fit of y = 0.00001408 * x + 0.427921 from the following 4 gain values | 104.8%
20903401 | POST (100% Humidity) | 2020-09-17 | 2033 | JKW | 34 |  | T37 | -0.0010572 | 2020-09-17 |  | 17.4438 |  | 6.1575 | 18.3028 | 0.0032692 |  | 0.0033541 | 0.0034766 | 1016 | 1 | 20.5 | 21.0907 | 21.0907 | -0.040634 | 294.225 | 0.99854 | 293.7942 | 274.9915 | 107.1% | 4
20903401 | POST2 (100% Humidity) | 2020-09-24 | 2146 | JKW | 34 |  | T37 | -0.0010572 |  |  | 16.9543 |  | 6.0219 | 20.9147 | 0.0031037 |  | 0.0025314 | 0.00053299 | 1017 | 1 | 21.1 | 24.8033 | 24.8033 | -0.058933 | 279.3587 | 0.99801 | 278.8021 | 268.9337 | 104.1% | 2 | 104.1% |  | 104.1%
20903401 | POST3 SETRA | 2020-09-25 | 1830 | JKW | 34 |  | T37 | -0.0010572 |  |  | 16.9497 |  | 6.0363 | 20.8618 | 0.0040608 |  | 0.0035743 | 0.00036922 | 1019 | 1 | 21.1 | 24.7227 | 24.7227 | -0.058562 | 279.6462 | 0.99802 | 279.0922 | 269.5781 | 104.1% | 2 | 104.1%
20A03401 | PRE (100% Humidity) | 2020-10-06 | 1751 | JKW | 34 |  | 10330 |  | 2020-10-06 |  | 17.292 |  | 6.3018 | 18.2288 | 0.0036332 |  | 0.0034871 | 0.0045452 | 1017 | 1 | 20.5 | 20.9931 | 20.9931 | -0.040116 | 294.667 | 0.99855 | 294.2396 | 281.4359 | 104.9% | 2 | 104.9% |  | 105.4%
20A03401 | POST (100% Humidity) | 2020-10-22 | 1445 | BDJ | 34 |  |  |  | N/A | N/A | 17.1138 |  | 5.883 | 20.9521 | 0.0044452 |  | 0.0023664 | 0.0078778 | 1011 | 1 | 21.1 | 24.8604 | 24.8604 | -0.059195 | 279.1558 | 0.998 | 278.5974 | 262.7324 | 105.8% | 2 | 105.8%
21303401 | PRE (100% Humidity) | 2021-03-23 | 930 | BDJ | 34 |  | 10330 |  |  |  | 17.0047 |  | 5.9782 | 20.9239 | 0.0030912 |  | 0.0025441 | 0.009507 | 1019 | 1 | 21 | 24.8173 | 24.8173 | -0.058998 | 279.3088 | 0.99801 | 278.7517 | 266.9825 | 105% | 2 | 105% |  | 105%
21503401 | PRE (100% Humidity) | 2021-05-05 | 10 | BDJ | 34 |  | T51 | -0.0010626 | 2021-05-05 | -1.4219 | 17.3254 |  | 6.1722 | 20.3347 | 0.0032 |  | 0.0028566 | 0.011252 | 1014 | 1 | 21.1 | 23.9329 | 23.9329 | -0.054869 | 282.5388 | 0.99813 | 282.0108 | 275.648 | 102.4% | 4 | 105.2% | Calculated based on fit of y = 0.00001408 * x + 0.427921 from the prior 4 gain values | 105.2%
21603401 |  | 2021-06-17 | [28 empty] | 105.3% | Calculated based on fit of y = 0.00001408 * x + 0.427921 from the prior 4 gain values | 105.3%
21903401 |  | 2021-09-21 | [28 empty] | 105.4% | Calculated based on fit of y = 0.00001408 * x + 0.427921 from the prior 4 gain values | 105.4%
22303401 | PRE (100% Humidity) | 2022-03-09 | 1123 | JKW/JSL | 34 | 212 | IDK | IDK | 20220309 | IDK | 17.5207 |  | 6.24 | 18.6836 | 0.0020548 |  | 0.0018856 | 0.00062183 | 1017.5 | 1 | 18.9 | 21.5997 | 21.5997 | -0.043302 | 291.9677 | 0.99846 | 291.519 | 278.6759 | 105.1% | 2 | 105.1% |  | 105.6% | 17.3236 | 33.4668 | 6.424 | 17.3287 | 17.6541 | -0.033812 | 0.81743 | 234.5143 | 215.9 | 220.1 |  | 92.06%
22303401 | POST (100% Humidity) | 2022-04-06 |  | JKW | 34 | 212 |  |  |  |  | 17.5686 |  | 6.1701 | 18.8198 | 0.0020062 |  | 0.0015947 | 0.00070676 | 1019.5 | 1 | 18.65 | 21.7843 | 21.7843 | -0.044256 | 291.1682 | 0.99844 | 290.7131 | 275.5547 | 106.2% | 2 | 106.2% |  | 106.2%
22B03401 | PRE (100% Humidity) | 2022-11-03 |  | JKW | 34 |  |  |  | 20221103 |  | 17.4609 | 218.8333 | 6.1336 | 19.4219 | 0.003618 | 105.8047 | 0.0019079 | 0.011348 | 1023 | 1 | 20.9 | 22.6171 | 22.6171 | -0.048474 | 287.6812 | 0.99832 | 287.1974 | 273.9241 | 105.9% | 2 |  |  | 108.3%
23303401 | PRE (100% Humidity) | 2023-03-02 |  | JKW | 34 |  |  |  | 20230302 |  | 17.7415 | 214.8333 | 6.3813 | 17.1794 | 0.0040559 | 104.5902 | 0.0048177 | 0.036124 | 1020 | 1 | 18.7 | 19.6501 | 19.6501 | -0.032766 | 301.073 | 0.99874 | 300.6938 | 284.9863 | 106.2% | 2 |  |  | 106.2%
23803401 | POST (100% Humidity) | 2023-08-01 |  | JKW | 34 |  |  |  | 20230801 |  | 17.6146 | 229 | 6.1969 | 18.5138 | 0.0023419 | 102.041 | 0.0018998 | 0.002586 | 1018 | 1 | 18.6 | 21.3714 | 21.3714 | -0.042112 | 292.9704 | 0.9985 | 292.5297 | 276.75 | 106.2% | 2 |  |  | 106.2%
24403401 | PRE (100% Humidity) | 2024-04-04 |  | JWK | 34 |  |  |  | 20240404 |  | 17.8915 | 222 | 6.3643 | 16.6803 | 0.0022472 | 108.4452 | 0.0017916 | 0.0067324 | 1011 | 1 | 18.3 | 19.0382 | 19.0382 | -0.029271 | 304.2103 | 0.99883 | 303.8534 | 284.2271 | 106.7% |  |  |  | 106.7%
24A03401 | PRE (100% Humidity) | 2024-10-21 | 10:00 | BW | 34 | [5 empty] | 17.771 | 16.9598 | 6.4055 | 17.2888 | 0.0026247 | 0.0040937 | 0.0025055 | 0.0046474 | 1020 | 1 | 17.4 | 19.7865 | 19.7865 | -0.033532 | 300.3937 | 0.99872 | 300.0095 | 286.0671 | 105.6% |  |  | Did not measure CTD offsets in TT so optode was at room temp not TT temp | 105.6%
[23 empty] | 2020-06-29 | 44011 | 1.0175 | 4 | 1.0476
[23 empty] | 2020-09-17 | 44091 | 1.0713 | 2 | 1.0487
[23 empty] | 2020-09-24 | 44098 | 1.0406 | 2 | 1.0488
[23 empty] | 2020-09-25 | 44099 | 1.0413 | 2 | 1.0488
[23 empty] | 2020-10-06 | 44110 | 1.0494 | 2 | 1.049
[23 empty] | 2020-10-22 | 44126 | 1.058 | 2 | 1.0492
[23 empty] | 2021-03-23 | 44278 | 1.0502 | 2 | 1.0514
[23 empty] | 2021-05-05 | 44321 | 1.0239 | 4 | 1.052
[23 empty] | 2022-03-09 | 44629 | 1.0506 | 2 | 1.0563
[23 empty] | 2022-04-06 | 44657 | 1.0617 | 2 | 1.0567
Note on AQ4: Yui Takeshita: Assumes salinity preset for the sensor = 0

## Data Translator
DO | TEMP | RAWPHASE
Z_1> | [5 empty] | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Barometer Temp | Barometer mb/hpa
-1.1 | 16.7103 | 0.0101 |  |  | Avg | 17.8915 | 222 | 6.3643 | 16.6803 | 18.3 | 1011
| 222 | 21710 | 1010 |  | Stdev | 0.0022472 | 108.4452 | 0.0017916 | 0.0067324
6.366 | 16.6916 | 17.886 | err= | 0 |  |  |  | 0.02815%
| 6366 | 21691 | 22886
Z_1>
| -1.1 | 16.7138 | 0.0101
| 222 | 21713 | 1010
6.363 | 16.6883 | 17.891 | err= | 0
| 6363 | 21688 | 22891
Z_1>
| -1.09 | 16.7167 | 0.0101
| 222 | 21716 | 1010
6.361 | 16.6858 | 17.894 | err= | 0
| 6361 | 21685 | 22894
Z_1>
| -1.1 | 16.7188 | 0.0101
| 222 | 21718 | 1010
6.364 | 16.6834 | 17.891 | err= | 0
| 6364 | 21683 | 22891
Z_1>
| -1.09 | 16.7168 | 0.0101
| 222 | 21716 | 1010
6.362 | 16.6812 | 17.894 | err= | 0
| 6362 | 21681 | 22894
Z_1>
| -1.09 | 16.7234 | 0.0102
| 222 | 21723 | 1010
6.365 | 16.6791 | 17.891 | err= | 0
| 6365 | 21679 | 22891
Z_1>
| -1.09 | 16.7264 | 0.0102
| 222 | 21726 | 1010
6.365 | 16.6769 | 17.892 | err= | 0
| 6365 | 21676 | 22892
Z_1>
| -1.09 | 16.7344 | 0.0102
| 222 | 21734 | 1010
6.366 | 16.6745 | 17.891 | err= | 0
| 6366 | 21674 | 22891
Z_1>
| -1.08 | 16.7368 | 0.0102
| 223 | 21736 | 1010
6.367 | 16.6723 | 17.891 | err= | 0
| 6367 | 21672 | 22891
Z_1>
| -1.11 | 16.7404 | 0.0102
| 222 | 21740 | 1010
6.364 | 16.6698 | 17.894 | err= | 0
| 6364 | 21669 | 22894
Z_1>

## old
Spray Air Calibration Log
Air cals start here
[11 empty] | Optode reading (average readings over 1 minute after stable) |  |  |  | Std dev |  |  |  | Barometer Reading |  |  | Calculations
Mission | Pre/Post | Date | Time | Operator | Spray ID | Optode S/N | pH S/N | K2 from Data Sheet | Test Tank Date | Test Tank K0 | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Raw Phase | Raw Thermistor | O2 (mL/L) | Temp | Atmospheric Pressure | Relative Humidity | Temp | pH2O @ 100% hum | pH2O real | T_O2soly | O2_soly (umol/kg) | Density water (kg/L) | O2_soly (uM) | O2raw,air (uM) | Gain
20703401 | PRE (100% Humidity) | 2020-06-29 | 15:10 | JKW | 34 |  | SBE_10330 | -0.0010296 |  | N/A | 16.7875 |  | 6.1235 | 21.1047 | 0.002062 |  | 0.002062 | 0.003401 | 1015 | 1 | 21.1 | 25.0944 | 25.0944 | -0.060265 | 278.3311 | 0.99797 | 277.765 | 273.4731 | 101.7%
20903401 | POST (100% Humidity) | 2020-09-17 | 2033 | JKW | 34 |  | T37 | -0.0010572 | 2020-09-17 |  | 17.4438 |  | 6.1575 | 18.3028 | 0.0032692 |  | 0.0033541 | 0.0034766 | 1016 | 1 | 20.5 | 21.0907 | 21.0907 | -0.040634 | 294.225 | 0.99854 | 293.7942 | 274.9915 | 107.1%
20903401 | POST2 (100% Humidity) | 2020-09-24 | 2146 | JKW | 34 |  | T37 | -0.0010572 |  |  | 16.9543 |  | 6.0219 | 20.9147 | 0.0031037 |  | 0.0025314 | 0.00053299 | 1017 | 1 | 21.1 | 24.8033 | 24.8033 | -0.058933 | 279.3587 | 0.99801 | 278.8021 | 268.9337 | 104.1%
20903401 | POST3 SETRA | 2020-09-25 | 1830 | JKW | 34 |  | T37 | -0.0010572 |  |  | 16.9497 |  | 6.0363 | 20.8618 | 0.0040608 |  | 0.0035743 | 0.00036922 | 1019 | 1 | 21.1 | 24.7227 | 24.7227 | -0.058562 | 279.6462 | 0.99802 | 279.0922 | 269.5781 | 104.1%
20A03401 | PRE (100% Humidity) | 2020-10-06 | 1751 | JKW | 34 |  | 10330 |  | 2020-10-06 |  | 17.292 |  | 6.3018 | 18.2288 | 0.0036332 |  | 0.0034871 | 0.0045452 | 1017 | 1 | 20.5 | 20.9931 | 20.9931 | -0.040116 | 294.667 | 0.99855 | 294.2396 | 281.4359 | 104.9%
20A03401 | POST (100% Humidity) | 2020-10-22 | 1445 | BDJ | 34 |  |  |  | N/A | N/A | 17.1138 |  | 5.883 | 20.9521 | 0.0044452 |  | 0.0023664 | 0.0078778 | 1011 | 1 | 21.1 | 24.8604 | 24.8604 | -0.059195 | 279.1558 | 0.998 | 278.5974 | 262.7324 | 105.8%
21303401 | PRE (100% Humidity) | 2021-03-23 | 930 | BDJ | 34 |  | 10330 |  |  |  | 17.0047 |  | 5.9782 | 20.9239 | 0.0030912 |  | 0.0025441 | 0.009507 | 1019 | 1 | 21 | 24.8173 | 24.8173 | -0.058998 | 279.3088 | 0.99801 | 278.7517 | 266.9825 | 105%
21503401 | PRE (100% Humidity) | 2021-05-05 | 10 | BDJ | 34 |  | T51 | -0.0010626 | 2021-05-05 | -1.4219 | 17.3254 |  | 6.1722 | 20.3347 | 0.0032 |  | 0.0028566 | 0.011252 | 1014 | 1 | 21.1 | 23.9329 | 23.9329 | -0.054869 | 282.5388 | 0.99813 | 282.0108 | 275.648 | 102.4%
22303401 | PRE (100% Humidity) | 2022-03-09 | 1123 | JKW/JSL | 34 | 212 | IDK | IDK | 20220309 | IDK | 17.5207 |  | 6.24 | 18.6836 | 0.0020548 |  | 0.0018856 | 0.00062183 | 1017.5 | 1 | 18.9 | 21.5997 | 21.5997 | -0.043302 | 291.9677 | 0.99846 | 291.519 | 278.6759 | 105.1%
22303401 | POST (100% Humidity) | 2022-04-06 |  | JKW | 34 | 212 |  |  |  |  | 17.5686 |  | 6.1701 | 18.8198 | 0.0020062 |  | 0.0015947 | 0.00070676 | 1019.5 | 1 | 18.65 | 21.7843 | 21.7843 | -0.044256 | 291.1682 | 0.99844 | 290.7131 | 275.5547 | 106.2%
[24 empty] | Tscaled | scorr |  |  | O2(uM) Glider Salinity corrected | #VALUE!
[11 empty] | 17.6541 |  | 6.424 | 17.3236 | [7 empty] | 19.83 | 21.7843 | -0.033812 | 0.81743 | 0.99871 | 0.81638 | 234.5143 | -0.007634%
I just copied this vlue rom above
[6 empty] | Temperature (CTD) | Salinity | O2(mL/L) | Temperature (optode) | phase | Tscaled | Scorr | O2(uM) Glider Salinity corrected | Winkler1 (uM) | Winkler2
[6 empty] | 17.3236 | 33.4668 | 6.424 | 17.3287 | 17.6541 | -0.033812 | 0.81743 | 234.5143 | 215.9 | 220.1
Note on M21: Yui Takeshita: Assumes salinity preset for the sensor = 0