## CALIBRATION SUMMARY
High-Pressure DuraFET pH sensor Logs
runt.ocean.washington.edu/swift/NoaaFloatAudit
| SN#'s |  | APEX Floats |  |  |  |  | HP Housings |  |  | PT-Chamber Tests (0.01 N HCl) |  |  |  | k2 Temperature coefficient(s) - can be a function of pressure | [8 empty] | Vrs pressure coefficients, f(P) = Vrs - Vc - k2T | [12 empty] | Lab k0 calibrations
| DF# | Ref# | APEX# | CTD# | MSC# | ISUS# | Flowcell | Sensor version | DFL# | Comment | Ptchamber | Ch | PT runs | PT Drift Adj (mV/d) | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0 (HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | Lab Cal | k0 Soln | k0 (pOFF) (V) | SdDev k0 (pOFF) (V) | k0 (pON) (V) | SdDev k0 (pON) (V) | Δk0 (OFF-ON) (µV) | K0 HCl vs SW | Comment
Hawaii
| 31 | 2 | 8374 | 4131 | 0 | none |  |  |  |  | 11/4-1/9/2012 |  | 200P & 4 T | 0 | [6 empty] | -0.0008579 | [12 empty] | Dual Stem | SS CE | 31 | [7 empty] | Hawaii | Dead on deployment; shorted bias battery
HOT | 35 | 35 | 7672 | 4132 | 0 | none |  |  |  |  | 3/8-15/2012 |  | 16 P & 2 T | -0.2 | [6 empty] | -0.00091046 | -1.2968 | 0.99702 |  |  | -3.4967e-22 | 1.2491e-18 | 7.7324e-16 | -8.5316e-12 | 1.5406e-08 | -1.6186e-05 | -1.2961 | 0.99995 | Dual Stem | SS CE | 35 | 6/13/2012 | Tris | -1.3025 | 0.0002 |  |  |  | Hawaii | Recovered (bias short) and redeployed
| 35 | 35 | [8 empty] | (recalc Jan 2016) |  |  | -0.25 | [6 empty] | -0.00094643 | -1.2951 | 0.9982 |  |  | 2.028e-22 | -2.6156e-18 | 1.1308e-14 | -2.2656e-11 | 2.5033e-08 | -1.9233e-05 | -1.2952 | 0.99987 | (@20C only)
HOT | 38a | 38 | 8486 | 4791 | 7038 | 7038 |  |  |  |  | 4/6-5/8/2012 |  | 84P & 6T (38a) | 0 | [6 empty] | -0.00095925 | -1.2794 | 0.99933 |  |  | y = -6.624E-22x6 + 4.882E-18x5 - 1.347E-14x4 + 1.700E-11x3 - 7.002E-09x2 - 6.808E-06x - 1.300E+00 | [6 empty] | 0.9971 | Single Stem w/ISUS | SS CE | 38a | 2/26-27/2013 | Tris | -1.286 | 0.0002 | -1.2866 |  | 607 | Hawaii
HOT Recov and sent to SO | 39 | 39 | 8514 | 4792 | 7039 | 7039 |  |  |  |  | 3/26-4/6/2012 |  | 22 P & 4 T | -0.08 | [6 empty] | -0.000959 | -1.2684 | 0.9995 | [10 empty] | Single Stem w/ISUS | SS CE | 39 | 4/12-13/2013 | Tris | -1.3083 | 0.0002 |  |  | 50
| 39 | 39 | 8514 | 4792 | 7039 | 7039 |  |  |  |  | 3/26-4/6/2012 |  | 22 P & 4 T | -0.08 | [6 empty] | -0.000959 | -1.2684 | 0.9995 | [10 empty] | Single Stem w/ISUS | SS CE | 39 | 6/11-12/2013 | Tris | -1.3077 | 0.0002 |  |  | 300
| 39 | 39 | 8514 | 4792 | 7039 | 7039 |  |  |  |  | 3/26-4/6/2012 |  | 22 P & 4 T | -0.08 | [6 empty] | -0.000959 | -1.2684 | 0.9995 | [10 empty] | Single Stem w/ISUS | SS CE | 39 | 6/25-26/2013 | Tris | -1.3085 | 0.0001 |  |  | 200 | Hawaii | Recovered Oct2013 w/multiplexer issue
| 39 | 39 | 8514 | 4792 | 7039 | 7039 |  |  |  |  | Redeployment: |  | 23 P & 4 T | -0.08 | [6 empty] | -0.000959 | -1.2684 | 0.9995 | [10 empty] | Single Stem w/ISUS | SS CE | 39 | 11/7-11/20013 | Tris | -1.3067 | 0.00028 |  |  | 200 |  | Recal with different logging PCB
| 39 | 39 | 8514 | 4792 | 7039 | 7039 |  |  |  |  | Redeployment: |  | 24 P & 4 T | -0.08 | [6 empty] | -0.000959 | -1.2684 | 0.9995 | [10 empty] | Single Stem w/ISUS | SS CE | 39 | 12/30/13-1/2/14 | Tris | -1.3064 | 0.0001 | -1.3066 |  | 202 |  | Recal with different logging PCB minus
Scripps
Yui | 52 | 52 | [5 empty] | Sent to Yui |  |  | 2/28-3/29/2013 |  | 3 Full PT | 0.2 | [6 empty] | -0.001206 |  |  |  |  | y = 2.289E-22x6 + 4.732E-19x5 - 6.714E-15x4 + 1.514E-11x3 - 9.584E-09x2 - 5.906E-06x - 1.383E+00 | [6 empty] | 0.9466 |  |  | 52 | [7 empty] | Yui | Being used at SIO
Yui | 53 | 53 | [5 empty] | Sent to Yui |  |  | 2/28-3/29/2014 |  | 4 Full PT | 0.15 | [6 empty] | -0.00117 |  |  |  |  | y = 5.939E-23x6 + 6.938E-19x5 - 4.823E-15x4 + 9.856E-12x3 - 5.806E-09x2 - 4.871E-06x - 1.384E+00 | [6 empty] | 0.8825 |  |  | 53 | [7 empty] | Yui | Being used at SIO
X-Prize
XP | 37a | 37 | [5 empty] | 4000 | 4001 | Ti CE | 4/6-5/8/2012 |  | 84P & 6T (37a) | 0 | [6 empty] | -0.00094816 | -1.2795 | 0.99946 |  |  | y = 1.100E-21x6 - 5.760E-18x5 + 1.002E-14x4 - 5.751E-12x3 + 1.434E-09x2 - 5.539E-06x - 1.300E+00 | [6 empty] | 0.9967 |  | g? | 37a | 5/3-20/2014 | Tris | -1.2829
[36 empty] | 37a | 7/24-25/2014 | TT | -1.2731 | (-1.3000 per Xprize spreadsheet) |  |  |  |  | In 4000m housing (periodic Ik leakage to 10nA)
XP | 43 | 43 | [5 empty] | 4000 | 4002 |  | 9/20-11/8/2012 |  | 3 Full PT | -0.025 | [6 empty] | -0.00095452 | -1.2961 | 0.95527 |  |  | y = -9.898E-22x6 + 7.208E-18x5 - 1.993E-14x4 + 2.436E-11x3 - 8.425E-09x2 - 7.208E-06x - 1.293E+00 | [6 empty] | 0.9842 |  | g? | 43 | 5/3-20/2014 | Tris | -1.3096
[36 empty] | 43 | 7/24-25/2014 | TT | -1.3096 | (-1.2930 per Xprize spreadsheet)
XP | 44 | 44 | [5 empty] | 4000 | 4003 |  | 11/8-20/2012 |  | 1 Full PT | -0.109 | [6 empty] | -0.00097885 | -1.3045 | 0.9629 |  |  | y = -1.206E-21x6 + 8.541E-18x5 - 2.303E-14x4 + 2.805E-11x3 - 1.169E-08x2 - 5.192E-06x - 1.302E+00 | [6 empty] | 0.9761 |  | g? | 44 | 5/3-20/2014 | Tris | -1.3027
[36 empty] | 44 | 7/24-25/2014 | TT | -1.302 | (-1.3020 per Xprize spreadsheet)
Xpoffset | 49 | 49 | [5 empty] | 4000 | 4002 |  | 12/21-1/14/2013 |  | 2 Full PT | 0.07 | [6 empty] | -0.0005841 | [12 empty] | HW Offset chips |  | 49 | [8 empty] | This contaminated chip run had Vrs=-1.7V.
SO Dec 2013
SO | 55 | 55 | 9099 | 5378 | 106 | 106 | PEEK w/N |  |  |  | 3/39-4/17/13 | B | 3 FullPT | -0.07 | [6 empty] | -0.0010946 | -1.3131 | 0.97606 |  |  | y = -2.856E-21x6 + 2.011E-17x5 - 5.589E-14x4 + 7.686E-11x3 - 5.131E-08x2 + 9.815E-06x - 1.312E+00 | [6 empty] | 0.9155 | SS, ISUS+pH, 1/6 | ng? | 55 | 12/2-5/2013 | Tris | -1.3027 | 5.1e-05 | -1.3028 | 5.7e-05 | 153 | For SO | Replaced pH board due to 15V supplied and recalibrated | -1.3055
[36 empty] | 55 |  | TT | -1.3044 | 7.3e-05 | -1.305 | 6.6e-05 | 569
SO | 56 | 56 | 9125 | 5387 | 105 | 105 | PEEK w/N | Returned from Hobart. Int. SCR open. |  |  | 4/18-5/14/13 |  | 2 FullPT | 0.09 | [6 empty] | -0.001185 | -1.4256 | 0.9359 |  |  | y = -2.282E-16x4 + 5.116E-13x3 + 2.979E-09x2 - 1.041E-05x - 1.422E+00 | [6 empty] | 0.996 | SS, ISUS+pH, 2/6 | ng? | 56 | 10/22-26/2013 |  | -1.4141 | 0.0002 | 120 |  |  | For SO | Returned from Hobart. Has intermitant open on Src. Installed #50 on float. Rebuilt #56 and it works again. But trust it? | -1.4141
SO | 58 | 58 | 9101 | 5381 | 101 | none | PEEK w/N |  |  |  | 5/23-31/2013 | A | 1 Full PT | 0.25 | [6 empty] | -0.0011179 | -1.3944 | 0.95634 |  |  | y = 2.002E-18x5 - 1.073E-14x4 + 2.008E-11x3 - 1.266E-08x2 - 4.141E-06x - 1.392E+00 | [6 empty] | 0.9913 | SS, pH only | ng | 58 | 7/31-8/1/2013 | Tris | -1.3862 | 0.0001 | -1.3867 |  | 572 | For SO |  | -1.3877
SO | 59 | 59 | 9018 | 5382 | 102 | none | PEEK w/N |  |  |  | 5/23-31/2013 | B | 1 Full PT | 0.55 | [6 empty] | -0.0011453 | -1.3911 | 0.9781 |  |  | y = -2.286E-21x6 + 1.506E-17x5 - 3.713E-14x4 + 3.969E-11x3 - 1.085E-08x2 - 6.204E-06x - 1.391E+00 | [6 empty] | 0.9831 | SS, pH only | ng | 59 | 7/31-8/1/2013 | Tris | -1.3773 | 0.0001 | -1.3778 |  | 452 | For SO |  | -1.3784
SO | 61 | 61 | 9031 | 5386 | 107 | 107 | PEEK w/N |  |  |  | 5/31-6/17/2013 |  | 2 FullPT | 0.11 | [6 empty] | -0.001146 | -1.3811 | 0.9323 |  |  | y = 1.540E-21x6 - 7.901E-18x5 + 1.399E-14x4 - 9.937E-12x3 + 5.169E-09x2 - 9.461E-06x - 1.378E+00 | [6 empty] | 0.9717 | SS, ISUS+pH, 3/6 | ng | 61 | 10/9-18/2013 | Tris | -1.3876 | 0.0002 | -1.3875 |  | -51 | For SO | TBT in calibration | -1.3886
SO | 62 | 62 | 9092 | 5379 | 103 | 103 | PEEK w/N |  |  |  | 6/18-7/9/2013 |  | 2 FullPT | 0.11 | [6 empty] | -0.0011988 | -1.3761 | 0.98855 |  |  | Y = -3.353E-21x6 + 2.386E-17x5 - 6.621E-14x4 + 8.961E-11x3 - 5.839E-08x2 + 1.223E-05x - 1.376E+00 | [6 empty] | 0.9019 | SS, ISUS+pH, 1/2 | ng | 62 | 8/29-9/5/2013 | Tirs | -1.3701 | 0.0001 | 150 |  |  | For SO |  | -1.3701
SO | 63 | 63 | 9091 | 5380 | 104 | 104 | PEEK w/N |  |  |  | 6/18-7/9/2013 |  | 2 FullPT | 0.24 | [6 empty] | -0.0011713 | -1.3775 | 0.95902 |  |  | y = -8.686E-22x6 + 6.779E-18x5 - 2.111E-14x4 + 3.286E-11x3 - 2.412E-08x2 + 2.026E-06x - 1.376E+00 | [6 empty] | 0.9638 | SS, ISUS+pH, 2/2 | ng | 63 | 9/24-26/2013 | Tris | -1.3856 | 0.0002 | 100 |  |  | For SO |  | -1.3856
SO | 65 | 65 | 9096 | 5588 | 110 | 110 | PEEK w/N | Returned from Seattle w/APEX issue |  |  | 7/10-8/5/2013 |  | 2 FullPT | 0.13 | [6 empty] | -0.0011237 | -1.3567 | 0.94545 |  |  | y = 1.335E-18x5 - 6.585E-15x4 + 1.116E-11x3 - 4.639E-09x2 - 7.372E-06x - 1.354E+00 | [6 empty] | 0.9626 | SS, ISUS+pH, 5/6 | ng | 65 | 11/7-11/2013 | Tris & SW | -1.3599 | 0.0001 | 200 |  |  | For SO | Leaked onto CTD controller & analog PCB during calibration. Recal with different analog PCB | -1.3599
SO | 64 | 64 | 9254 | 5385 | 108 | 108 | PEEK w/N |  |  |  | 7/10-8/5/2013 |  | 2 FullPT | 0.38 | [6 empty] | -0.0010634 | -1.3572 | 0.97109 |  |  | Y = -1.860E-21x6 + 1.416E-17x5 - 4.247E-14x4 + 6.264E-11x3 - 4.424E-08x2 + 8.567E-06x - 1.356E+00 | [6 empty] | 0.9838 | SS, ISUS+pH, 4/6 | ng | 64 | 10/11-18/2013 | Tris | -1.3708 | 0.0001 | 150 |  |  | For SO | TBT in calibration | -1.3708
SO | 68 | 68 | 9095 | 5383 | 109 | 109 | PEEK w/N |  |  |  | 10/8-10/18/2013 |  | 1 Full PT | 0.18 | [6 empty] | -0.0010752 | -1.3628 | 0.9427 |  |  | y = -9.245E-22x6 + 7.025E-18x5 - 2.089E-14x4 + 2.977E-11x3 - 1.728E-08x2 - 3.478E-06x - 1.360E+00 | [6 empty] | 0.9959 | SS, ISUS+pH, 6/6 | ng | 68 | 11/2-6/2013 | Tris | -1.3689 | 0.0001 | -1.3692 |  | 267 | For SO |  | -1.3698
SO2 Jul 2014 |  |  | To be deployed Dec 2014
SO2 | 48 | 48 | 9313-CF | 5599 | 205 | 205 | PrAc+P | renstalled and recal'ed for CF hull |  |  | 12/21-1/14/2013 |  | 2 Full PT | 0.09 | [6 empty] | -0.0010635 | -1.3905 | 0.9999 |  |  | -3.7451e-22 | 2.6177e-18 | -6.4357e-15 | 5.7296e-12 | 1.4662e-09 | -6.0551e-06 | -1.3905 | 0.99919 |  | ng? | 48 | 6/26-7/1/2014 | TT | -1.3941 | 6.4e-05 | -1.3944 | 0.000154 | 318 | For SO2
[36 empty] | 48 | 7/23-25/2014 | TT | -1.3938 | 6.8e-05 | -1.3943 | 0.000179 | 497 |  | k0 called with XP units in-line
SO2 | 50a | 50 | 9125 | 5387 | 105 | 105 | PrAc+P | Recal'ed and shipped (Hobart) |  |  | 2/1-25/2013 |  | 2 Full PT | 0.13 | [6 empty] | -0.0011604 | -1.3704 | 0.9936 |  |  | -5.9883e-22 | 6.084e-18 | -2.2201e-14 | 3.714e-11 | -2.5701e-08 | -1.2337e-06 | -1.3703 | 0.9999 |  | g? | 50a | 6/21-25/2014 | TT | -1.3698 | 9.9e-05 | -1.3701 | 0.000125 | 385 | For SO2 | Installed on float returned from Hobart (that had intermitant #56
SO | 65 | 65 | 9094 | 5603 | 110 | 110 | PrAc+P | 1/15/16 High Ik and noisy |  |  | 7/10-8/5/2013 |  | 2 FullPT | 0.13 | [6 empty] | -0.0011092 | -1.356 | 0.997 |  |  | -3.2186e-22 | 3.4027e-18 | -1.1841e-14 | 1.7862e-11 | -9.0966e-09 | -5.9346e-06 | -1.3524 | 0.99994 | SS, ISUS+pH, 5/6 | ng | 65 | 11/7-11/2013 | Tris | -1.3592 |  | -1.3594 |  | 208 | For SO2 | Returned from SO1, and Recal'ed and shipped for June ballasting
[36 empty] | 65 | 6/10-13/2014 | TT | -1.3579 |  | -1.3581 |  | 203
SO2 | 67a | 57 | 9096 | 5592 | 202 | 202 | PrAc+P | shipped |  |  | 4/11-22/2014 | A | 3T & 1 full PT | 0.22 | [6 empty] | -0.0011441 | -1.3349 | 0.9991 |  |  | -3.2317e-21 | 2.2315e-17 | -6.0013e-14 | 7.9018e-11 | -5.2033e-08 | 1.4359e-05 | -1.3352 | 0.98634 |  | ng | 67a | 7/7-11/2014 | TT | -1.3384 | 5.7e-05 | -1.3387 | 6.2e-05 | 244 | For SO2
SO2 | 69a | 76 | 9260 | 5583 | 201 | 201 | PrAc+P | shipped |  |  | 4/25-5/5/2014 | A | 3T & 1 full PT | 0.18 | [6 empty] | -0.0010744 | -1.3337 | 0.9999 |  |  | -2.4436e-21 | 1.6929e-17 | -4.521e-14 | 5.7532e-11 | -3.3383e-08 | 4.1344e-06 | -1.3337 | 0.99931 |  | ng | 69a | 7/7-9/2014 | TT | -1.3371 | 5.3e-05 | -1.3373 | 6e-05 | 157 | For SO2 | DF#69 had leak, rebuilt as DF#69a and teted fine
SO2 | 76 | 70 | 9275 | 5582 | 204 | 204 | PrAc+P | shipped |  |  | 5/12-27/2014 | B | 3T & 1 full PT | 0.04 | [6 empty] | -0.0010757 | -1.3652 | 0.9996 |  |  | -1.3429e-21 | 9.8481e-18 | -2.6819e-14 | 3.2857e-11 | -1.3633e-08 | -6.7177e-06 | -1.3649 | 0.99962 |  | ng | 76 | 7/12-15/2014 | TT | -1.3663 | 5.2e-05 | -1.3665 | 5.2e-05 | 191 | For SO2
SO2 | 72 | 72 | 9325-CF | 5588 | 203 | 203 | PrAc+P | on RV Investigator |  |  | 3/24-4/11/2014 | B | 3T & 2 full PT | 0.02 | [6 empty] | -0.0010714 | -1.3825 | 0.9999 |  |  | 1.7322e-21 | -1.0643e-17 | 2.4938e-14 | -2.9666e-11 | 2.4496e-08 | -2.0554e-05 | -1.382 | 0.99979 |  | ng | 72 | 7/23-25/2014 | TT | -1.3844 | 0.000138 | -1.3846 | 0.000146 | 214 | For SO2
[36 empty] | 72 | 10/23-24/2014 | TT | -1.3842 | 6.6e-05 | -1.3843 | 5.9e-05 | 110
SO2 | 80 | 80 | 8514 | 4792 | 7039 | 7039 | PrAc+P | Installed on failed Hawaii float |  |  | 6/27-7/7/2014 | A | 3T & full PT | 0.67 | [6 empty] | -0.0010419 | -1.3918 | 0.9994 |  |  | -1.7243e-21 | 1.2225e-17 | -3.3781e-14 | 4.4838e-11 | -2.5848e-08 | -1.0988e-06 | -1.3916 | 0.99972 |  | ng | 80 | 10/14-22/2014 | TT | -1.3986 | 5.2e-05 | -1.3987 | 5e-05 | 111 | For SO2 | Installed on failed Hawaii float: new pH on same APEX#
|  |  | CF = Carbon Fiber Hull | [42 empty] | Cal'ed by Seabird @ MBARI
Seabird 2014 | [46 empty] | MBARI | SBE | Δk0 MBARI-SBE
SB-HI | 66 | 66 | F0412 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 9/16-10/6/2013 | A | 2 FullPT | 0.5 | [6 empty] | -0.0010698 | -1.3987 | 0.9998 |  |  | -1.8626e-21 | 1.3806e-17 | -3.9537e-14 | 5.3836e-11 | -3.192e-08 | 2.4063e-07 | -1.3986 | 0.99984 |  | ng | 66 | 1/29-31/2014 | Tris | -1.4037 | 7.2e-05 | (NA) |  |  | For SO2 | Developed high Ik's | 2014-10-07 | -1.4 | -1.404 | 0.000282
SB-SO | 75 | 75 | F0511 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 5/12-27/2014 | A | 3T & full PT | 0 | [6 empty] | -0.0010186 | -1.3804 | 0.9999 |  |  | -1.4646e-21 | 1.0132e-17 | -2.6553e-14 | 3.1541e-11 | -1.2986e-08 | -6.1752e-06 | -1.3748 | 0.99992 |  | ng | 75 | 7/11-14/2014 | TT | -1.3768 | 3.9e-05 | -1.377 | 3.9e-05 | 263 | For SO2 |  | 2014-09-03 | -1.3761 | -1.3734 | -0.003352
SB-SO | 78 | 71 | F0506 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 6/2-13/2014 | A | 3T & full PT | 0.2 | [6 empty] | -0.0010585 | -1.3549 | 0.9999 |  |  | -1.031e-21 | 7.9731e-18 | -2.3624e-14 | 3.2053e-11 | -1.471e-08 | -4.7121e-06 | -1.3547 | 0.99914 |  | ng | 78 | 7/15-17/2014 | TT | -1.3603 | 5.9e-05 | -1.3604 | 5.3e-05 | 163 | For SO2 |  | 2014-09-03 | -1.3596 | -1.3569 | -0.003379
SB-SO | 73 | 73 | F0509 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 6/2-13/2014 | B | 3T & full PT | 0.35 | [6 empty] | -0.0010619 | -1.3766 | 0.9998 |  |  | 1.1911e-21 | -8.2415e-18 | 2.2911e-14 | -3.3695e-11 | 3.0227e-08 | -1.8469e-05 | -1.3764 | 0.99962 |  | ng | 73 | 7/19-24/2014 | TT | -1.3838 | 6.4e-05 | -1.384 | 5.6e-05 | 148 | For SO2 |  | 2014-09-03 | -1.3838 | -1.3808 | -0.003023
SB-SO | 74 | 74 | F0507 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 4/11-22/2014 | B | 3T & full PT | 0.06 | [6 empty] | -0.0010769 | -1.3558 | 0.93218 |  |  | -5.4259e-22 | 4.4819e-18 | -1.3927e-14 | 1.9149e-11 | -7.4715e-09 | -8.3184e-06 | -1.3558 | 0.99983 |  | ng | 74 | 7/19-22/2014 | TT | -1.361 | 7.7e-05 | -1.3611 | 6.7e-05 | 135 | For SO2 |  | 2014-09-03 | -1.3605 | -1.3575 | -0.003461
SB-SO | 82 | 82 | F0508 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 7/11-8/4/2014 | A | 3T & 2 full PT | -0.15 | [6 empty] | -0.0010071 | -1.3136 | 0.9991 |  |  | 2.091e-22 | -6.7718e-19 | -7.7313e-16 | 4.3959e-12 | -1.5494e-09 | -7.923e-06 | -1.3135 | 0.99978 |  | ng | 82 | 8/17-21/2014 | TT | -1.3228 | 4.8e-05 | -1.3231 | 5e-05 | 241 | For SO2 | Damaged pH sensort during deployment? | 2014-09-03 | -1.3223 | -1.3197 | -0.003141
SB-SO | 83 | 83 | F0510 |  | SBE Float |  | SB PE | k0 cal'ed by SB @ MBARI |  |  | 7/11-8/4/2015 | B | 3T & 2 full PT | -0.04 | [6 empty] | -0.0010273 | -1.3762 | 0.9998 |  |  | -5.0653e-22 | 4.013e-18 | -1.2696e-14 | 2.0228e-11 | -1.6686e-08 | 5.8761e-06 | -1.3762 | 0.99264 |  | ng | 83 | 8/17-21/2014 | TT | -1.3833 | 5.3e-05 | -1.3834 | 5.1e-05 | 106 | For SO2 |  | 2014-09-03 | -1.3827 | -1.3805 | -0.002815
SO 2015 |  |  |  |  | Beginning here, the Source and Drain are connected the same for PT and float | [30 empty] | DF# |  |  | k0 (pOFF) (V) | SdDev k0 (pOFF) (V) | k0 (pON) (V) | SdDev k0 (pON) (V) | Δk0 (OFF-ON) (µV)
| 81 | 81 | 8501 | 5590 | 702 | New | epoxyABS | Bad S, being returned |  |  | 6/27-7/7/2014 | B | 3T & full PT | 0.14 | [6 empty] | -0.0010402 | -1.3514 | 0.99984 |  |  | -9.3278e-22 | 7.4402e-18 | -2.3062e-14 | 3.3868e-11 | -2.0498e-08 | -2.7952e-06 | -1.3512 | 0.99992 |  | '13 | 81 | 4/2-6/2015 SW |  | -1.3559 | 6.1e-05 | -1.356 | 4.1e-05 | 115 | 4.47 | equil w/SW 3/16-31, cal (SW is seawater in the seawater lab, no ozone) | Bad S in Seattle
[36 empty] | 81 | 7/6-13/2015 SW |  | -1.3563 | 3.9e-05 | -1.3564 | 4e-05 | 92 |  | Rebuild SW
[36 empty] | 81 | 7/14-24/2015 TT |  | -1.3555 | 5.2e-05 | -1.3556 | 4.8e-05 | 63 |  | Rebuild TT
[36 empty] | 81 | 9/21/22/2015 TT |  | -1.3534 | 3e-05 |  |  |  |  | Assy Dana TT
Hawaii | 90 | 90 | 9274 | 5585 | 703 | New | pnPolyPro | 2 |  |  | 10/23-11/3/2014 | A | 3T & full PT | 0 | [6 empty] | -0.0010546 | -1.3762 | 0.9998 |  |  | -3.5379e-21 | 2.5712e-17 | -7.3835e-14 | 1.063e-10 | -8.0006e-08 | 2.8648e-05 | -1.3767 | 0.99062 |  | '14 | 90 | 4/2-6/2015 |  | -1.3769 | 4.8e-05 | -1.377 | 4.6e-05 | 115 | 0.659 | equil w/TT 3/16-31, cal (TT is seawater in the test tank)
| 84 | 84 | 9666 | 6726 | 701 | 701 | bPolyPro | 3 | Dec '15 Palmer |  | 9/12-25/2014 | A | 3T & full PT | 0.035 | [6 empty] | -0.0010186 | -1.3601 | 0.9999 |  |  | -5.3032e-21 | 3.7876e-17 | -1.0634e-13 | 1.4769e-10 | -1.0255e-07 | 2.8005e-05 | -1.3603 | 0.99788 |  | '13 | 84 | 4/25-28/2015 |  | -1.3653 | 4.9e-05 | -1.3652 | 5e-05 | -12 | 5.119 | eqil w/TT 2/25-3/31/2015, cal
[6 empty] | k0 Recal | [29 empty] | 84 | 4/30-5/5/2015 |  | -1.3651 | 5.7e-05 | -1.3651 | 6.8e-05 | 8 |  | Deployed in TestTank right after lab cal
| 85 | 85 | 9662 | 6721 | 704 | 704 | bPolyPro | 4 |  |  | 9/12-25/2014 | B | 3T & full PT | 0.1 | [6 empty] | -0.0010722 | -1.3652 | 0.9989 |  |  | 2.7291e-21 | -1.7917e-17 | 4.4358e-14 | -4.9524e-11 | 2.1081e-08 | 7.9501e-07 | -1.3651 | 0.97923 |  | '13 | 85 | 3/2-5/2015 |  | -1.3705 | 5.6e-05 | -1.3706 | 5.2e-05 | 105 | 5.342 | No Br Equilibration
[6 empty] | k0 Recal | [29 empty] | 85 | 3/16-18/2015 |  | -1.3703 | 4e-05 | -1.3704 | 3.6e-05 | 79 |  | After stem equil 11 days in TT
[6 empty] | k0 Recal | [29 empty] | 85 |  |  | -1.3704 | 5.3e-05 | -1.3706 | 6.1e-05 | 211 |  | With printed flowcell (acrylic)
[6 empty] | k0 Recal | [29 empty] | 85 |  |  | -1.3703 | 3.8e-05 | -1.3704 | 3.7e-05 | 102 |  | Back to PEEK flowcell
[6 empty] | k0 Recal | [29 empty] | 85 |  |  | -1.3715 | 4.1e-05 | -1.3715 | 4.1e-05 | 12 |  | Dep in Test Tank, Ib railed due to unshielded can (ABS flowcell)
[6 empty] | k0 Recal | [29 empty] | 85 |  |  | -1.3716 | 3.4e-05 | -1.3716 | 3.5e-05 | 10 |  | Dep in Test Tank, shielded can, good data (ABS flowcell)
[6 empty] | k0 Recal | [29 empty] | 85 |  |  | -1.3721 | 4.6e-05 | -1.3721 | 4.5e-05 | 6 |  | Same; 4 days later
[6 empty] | k0 Recal | [29 empty] | 85 | 4/24-28/2015 |  | -1.3719 | 5.2e-05 | -1.372 | 4.6e-05 | 107 |  | Back in Lab with TT water
| 86 | 86 | 9652 | 6728 | 705 | 705 | epoxyABS | 5 |  |  | 9/25-10/7/2014 | A | 3T & full PT | -0.02 | [6 empty] | -0.0010226 | -1.3759 | 0.9984 |  |  | -8.8482e-22 | 7.2167e-18 | -2.3329e-14 | 3.8402e-11 | -3.3913e-08 | 1.4461e-05 | -1.3759 | 0.99648 |  | '13 | 86 | 3/2-5/2016 |  | -1.3792 | 0.000111 | -1.3793 | 9.8e-05 | 87 | 3.294 | No Br Equilibration
[6 empty] | k0 Recal | [29 empty] | 86 | 3/16-18/2015 |  | -1.3791 | 6.9e-05 | -1.3791 | 6.4e-05 | 15 |  | Then eqil 11 days w/SW, cal
| 87 | 87 | Stem | 6722 | 706 | 706 | pnPolyPro | Leaked in Seattle |  |  | 9/25-10/7/2014 | B | 3T & full PT | 0.08 | [6 empty] | -0.001024 | -1.3513 |  |  |  | -1.0991e-21 | 8.5475e-18 | -2.6756e-14 | 4.3334e-11 | -3.7787e-08 | 1.4779e-05 | -1.3513 | 0.9951 |  | '13 | 87 | 3/6-10/2015 |  | -1.3545 | 8.1e-05 | -1.3548 | 8.5e-05 | 331 | 3.169 | eqil 2/25-3/5/2015 (8d) w/TT, cal (with current sensing chip) | Leaked in Seattle
[6 empty] | k0 Recal | [29 empty] | 87 | 3/26-31/2015 |  | -1.3545 | 5e-05 | -1.3548 | 3.9e-05 | 269 |  | eqil 3/11-31/2015 (19d) w/TT, cal
'14 Die | 89 | 89 | 9657 | 6724 | 707 | 707 | pnPolyPro | 6 |  |  | 10/7-16/2014 | B | 3T & full PT | 0 | [6 empty] | -0.0010477 | -1.3626 | 0.999 |  |  | 5.6968e-23 | 7.4674e-19 | -6.0455e-15 | 1.6159e-11 | -1.9869e-08 | 1.0533e-05 | -1.3629 | 0.99695 | Start with 2014 Die | '14 | 89 | 3/12-16-2015 |  | -1.3652 | 5e-05 | -1.3653 | 4.5e-05 | 50 | 2.624 | eqil 2/25-3/12/2015 (15d) w/ TT, cal
| 91 | 91 | 9655 | 6723 | 708 | 708 | pAcrylic | 7 |  |  | 10/23-11/3/2014 | B | 3T & full PT | 0.12 | [6 empty] | -0.001091 | -1.3957 |  |  |  | -4.7304e-21 | 3.3086e-17 | -8.9621e-14 | 1.1702e-10 | -7.3275e-08 | 1.6705e-05 | -1.3956 | 0.99114 |  | '14 | 91 | 3/6-10/2015 |  | -1.3955 | 7.2e-05 | -1.3956 | 7.3e-05 | 88 | -0.185 | eqil 2/25-3/5/2015 (8d) w/SW, cal (with current sensing chip)
[6 empty] | k0 Recal | [29 empty] | 91 | 3/26-31/2015 |  | -1.3953 | 3.9e-05 | -1.3954 | 3.7e-05 | 72 |  | eqil 14d w/SW, 5d@2km, cal
| 92 | 92 | 9668 | 6725 | 709 | 709 | bPolyPro | 8 | Dec '15 Palmer |  | 11/3-13/2014 | A | 3T & full PT | 0.1 | [6 empty] | -0.0010575 | -1.3684 | 0.9999 |  |  | 2.4516e-22 | -5.2307e-19 | -3.0911e-15 | 1.4446e-11 | -2.1292e-08 | 7.0736e-06 | -1.3684 | 0.99995 |  | '14 | 92 | 3/12-16-2015 |  | -1.3657 | 4.4e-05 | -1.3659 | 4.2e-05 | 209 | -2.682 | eqil 2/25-3/12/201 (15d) w/SW, cal
| 93 | 93 | 9646 | 6727 | 710 | 710 | bPolyPro | 9 | Dec '15 Palmer |  | 11/3-13/2014 | B | 3T & full PT | 0.07 | [6 empty] | -0.0010729 | -1.3628 | 0.9999 |  |  | -1.3942e-22 | 1.8103e-18 | -7.9551e-15 | 1.6502e-11 | -1.7512e-08 | 8.4696e-06 | -1.3629 | 0.99583 |  | '14 | 93 | 4/10-23/2015 |  | -1.3609 | 4e-05 | -1.3609 | 4.2e-05 | 24 | -1.959 | eqil 2/25-3/31/2015 (34d) w/SW, cal
| 95 | 95 | 9637 | 6729 | 712 | 10 | epoxyABS | 10 |  |  | 11/13-28/2014 | B | 3T & full PT | 0.1 | [6 empty] | -0.0010983 | -1.3582 | 0.9998 |  |  | -5.8504e-22 | 4.1397e-18 | -1.159e-14 | 1.6444e-11 | -1.226e-08 | 4.0828e-06 | -1.3582 | 0.83057 |  | '14 | 95 | 5/1-5/2015 |  | -1.3578 | 6.2e-05 | -1.3582 | 5.1e-05 | 396 | -0.399 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 96 | 96 | 9650 | 6730 | 711 | 11 | epoxyABS | 11 |  |  | 12/8-18/2014 | A | 3T & full PT | 0 | [6 empty] | -0.0010768 | -1.3758 | 0.9999 |  |  | -1.6615e-21 | 1.2547e-17 | -3.6649e-14 | 5.0922e-11 | -3.1355e-08 | 2.0916e-06 | -1.376 | 0.99985 |  | '14 | 96 | 5/1-5/2015 |  | -1.3705 | 5.2e-05 | -1.3707 | 4.3e-05 | 144 | -5.279 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 97 | 97 | 12552 | 6737 | 713 | 12 | bPolyPro | 12 |  |  | 12/19-30/2014 | A | 3T & full PT | 0.24 | [6 empty] | -0.0011404 | -1.3969 | 0.9998 |  |  | 3.2519e-22 | -1.1243e-18 | -5.2261e-17 | 2.6352e-12 | 1.8297e-09 | -9.5541e-06 | -1.397 | 0.99964 |  | '14 | 97 | 5/7-12/2015 |  | -1.3903 | 4.2e-05 | -1.3906 | 3.8e-05 | 265 | -6.649 | eqil 4/15-5/4/2015 (19d) in TT, cal | Unusually light sensitive, spikes removed .
| 98 | 98 | 9632 | 6740 | 714 | 13 | bPolyPro | 13 |  |  | 12/19-30/2014 | B | 3T & full PT | 0.32 | [6 empty] | -0.0011374 | -1.3774 | 0.9998 |  |  | 2.6064e-21 | -1.7474e-17 | 4.5571e-14 | -5.9487e-11 | 4.3879e-08 | -2.1415e-05 | -1.3777 | 0.99883 |  | '14 | 98 | 5/7-12/2015 |  | -1.3715 | 4.4e-05 | -1.3719 | 6.8e-05 | 333 | -5.865 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 99 | 99 | 9630 | 6738 | 715 | 14 | pAcylic | 14 |  |  | 1/13-29/2015 | A | 2@3T & full PT | 0.13 | [6 empty] | -0.0010962 | -1.3729 | 0.9998 |  |  | -1.0101e-22 | 2.2062e-18 | -1.059e-14 | 1.9993e-11 | -1.3464e-08 | -4.1001e-06 | -1.3732 | 0.99974 |  | '14 | 99 | 5/12-15/2015 |  | -1.3782 | 6.3e-05 | -1.3784 | 4.3e-05 | 181 | 5.271 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 100 | 100 | 9631 | 6739 | 716 | 15 | pAcylic | 15 |  |  | 1/13-29/2015 | B | 2@3T & full PT | 0.26 | [6 empty] | -0.0010923 | -1.3539 | 0.9999 |  |  | -1.0498e-21 | 8.3496e-18 | -2.5679e-14 | 3.7135e-11 | -2.2124e-08 | -1.6688e-06 | -1.3537 | 0.99987 |  | '14 | 100 | 5/12-15/2015 |  | -1.3579 | 5.3e-05 | -1.3579 | 4.6e-05 | 47 | 3.958 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 101 | 101 | 12382 | 7183 | 717 | 16 | bPolyPro | 16 |  |  | 2/26-3/2/2015 | A | 1@SO | 0.21 | [6 empty] | -0.0011234 | -1.3791 | 0.9996 |  |  | 7.7316e-23 | 1.0712e-19 | -2.4407e-15 | 6.15e-12 | -3.1218e-09 | -6.2479e-06 | -1.3793 | 0.99991 |  | '14 | 101 | 5/15-19/2015 |  | -1.3802 | 5.3e-05 | -1.3804 | 5.9e-05 | 215 | 1.139 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 102 | 102 | 9750b | 7184 | 718 | 17 | bPolyPro | 17 |  |  | 3/6-14/2015 | A | 2@SO | 0.07 | [6 empty] | -0.0011358 | -1.3515 | 0.9996 |  |  | 3.449e-22 | -1.5413e-18 | 1.4334e-15 | 1.6807e-12 | -5.6651e-10 | -6.9858e-06 | -1.3516 | 0.99985 |  | '14 | 102 | 5/15-19/2016 |  | -1.3487 | 6.1e-05 | -1.3489 | 9.6e-05 | 255 | -2.85 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 87 | 87 | 9749 | 7185 | 719 | 18 | bPolyPro | 18 | Shipped 6/26/15 |  | 9/25-10/7/2014 | B | 3T & full PT | 0.08 | [6 empty] | -0.001024 | -1.3513 |  |  |  | -1.0991e-21 | 8.5475e-18 | -2.6756e-14 | 4.3334e-11 | -3.7787e-08 | 1.4779e-05 | -1.3513 | 0.9951 |  | '13 | 87 | 5/19-22/15 |  | -1.3552 | 8.8e-05 | -1.3553 | 7.9e-05 | 70 | 3.931 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 103 | 103 | 8204 | 6731 | 720 | 19 | bPolyPro | 19 | Shipped 6/26/15 |  | 3/6-14/2015 | B | 2@SO | 0.07 | [6 empty] | -0.0011102 | -1.3991 | 0.9999 |  |  | -4.2489e-22 | 3.8144e-18 | -1.3066e-14 | 2.088e-11 | -1.336e-08 | -2.5853e-06 | -1.3991 | 0.9998 |  | '14 | 103 | 5/22-27/15 |  | -1.3956 | 5.6e-05 | -1.3957 | 6.1e-05 | 106 | -3.52 | previously used in #706
| 104 | 104 | 9600 | 6734 | 721 | 20 | bPolyPro | 20 | Shipped 6/26/15 |  | 3/14-21/2015 | A | 2@SO | 0 | [6 empty] | -0.001108 | -1.3949 | 0.9999 |  |  | -1.3966e-21 | 1.011e-17 | -2.898e-14 | 4.1425e-11 | -2.8942e-08 | 5.1592e-06 | -1.3949 | 0.99892 |  | '14 | 104 | 5/22-27/15 |  | -1.393 | 9.1e-05 | -1.393 | 8.1e-05 | -33 | -1.871 | equil in PT and cal only
'15 die | 105 | 105 | 9645 | 6735 | 722 | 21 | bPPshield | 21 | Dep w/shield only |  | 3/14-21/2015 | B | 2@SO | 0 | [6 empty] | -0.0010829 | -1.3738 | 0.9999 |  |  | 1.5085e-21 | -9.5325e-18 | 2.2511e-14 | -2.51e-11 | 1.6189e-08 | -1.0857e-05 | -1.3739 | 0.99962 |  | '15 | 105 | 5/28-6/4/2015 |  | -1.3718 | 4.7e-05 | -1.3719 | 4.2e-05 | 102 | -2.075 | equil in PT and cal only
| 106 | 106 | 9642 | 6736 | 723 | 22 | bPPshield | 22 | Dep w/shield only |  | 3/21-30/2015 | A | 2@SO | 0 | [6 empty] | -0.0010932 | -1.3775 | 0.9998 |  |  | -1.1265e-22 | 9.6611e-19 | -3.0046e-15 | 3.5867e-12 | 8.7202e-10 | -6.3085e-06 | -1.3774 | 0.99785 | Start with 2015 Die | '15 | 106 | 5/28-6/4/2015 |  | -1.3813 | 5.5e-05 | -1.3815 | 5.8e-05 | 249 | 3.821 | equil in PT and cal only
[36 empty] | 106 | 9/2-9/2015 |  | -1.381 | 8.9e-05 | -1.3813 | 7.1e-05 | 265 |  | Mini-recal after ISUS issue
| 107 | 107 | 9762 | 5601 | 724 | 23 | bPolyPro | 23 | CF Hull |  | 3/21-30/2015 | B | 2@SO | 0.05 | [6 empty] | -0.0010511 | -1.3906 | 0.9999 |  |  | -7.8624e-22 | 6.5671e-18 | -2.1285e-14 | 3.2999e-11 | -2.272e-08 | 1.489e-06 | -1.3905 | 0.99983 |  | '15 | 107 | 6/4-9/2015 |  | -1.3938 | 5.6e-05 | -1.3939 | 5e-05 | 53 | 3.247 | equil in PT and cal only
| 108 | 108 | 9660 | 7186 | 725 | 24 | bPPshield | 24 | Dep w/shield only |  | 3/30-4/6/2015 | A | 2@SO | 0.04 | [6 empty] | -0.0010295 | -1.3584 | 1 |  |  | -1.805e-21 | 1.3291e-17 | -3.8226e-14 | 5.289e-11 | -3.25e-08 | 1.6855e-06 | -1.3583 | 0.99973 |  | '15 | 108 | 6/4-9/2015 |  | -1.3633 | 4.9e-05 | -1.3635 | 5.1e-05 | 198 | 4.888 | equil in PT and cal only
| 109 | 109 | 9659 | 6732 | 726 | 25 | bPPshield | 25 | Dep w/shield only |  | 4/6-15/2015 | A | 2@SO | 0.08 | [6 empty] | -0.0010554 | -1.3468 | 0.9999 |  |  | -2.8805e-21 | 2.1088e-17 | -6.0075e-14 | 8.2407e-11 | -5.1826e-08 | 6.5011e-06 | -1.3467 | 0.99956 |  | '15 | 109 | 6/9-15/2015 |  | -1.3507 | 4.9e-05 | -1.3509 | 5.1e-05 | 246 | 3.9 | equil in PT and cal only
| 110 | 110 | 9602 | 6733 | 727 | 26 | bPPshield | 26 | Dep w/shield only |  | 4/6-15/2015 | B | 2@SO | 0.03 | [6 empty] | -0.0010663 | -1.3896 | 0.9999 |  |  | -3.0619e-22 | 2.4041e-18 | -7.5825e-15 | 1.1488e-11 | -5.699e-09 | -6.0045e-06 | -1.3898 | 0.99951 |  | '15 | 110 | 6/9-15/2015 |  | -1.3934 | 4.4e-05 | -1.3935 | 4.6e-05 | 48 | 3.861 | equil in PT and cal only
| 111 | 111 | 9744 | 7175 | 728 | 27 | bPolyPro | 27 | CF Hull |  | 4/15-23/2015 | A | 2@SO | 0.1 | [6 empty] | -0.0010497 | -1.3545 | 0.9999 |  |  | -4.7147e-21 | 3.2165e-17 | -8.4675e-14 | 1.0743e-10 | -6.4905e-08 | 1.1315e-05 | -1.3544 | 0.99145 |  | '15 | 111 | 6/15-19/2015 |  | -1.3574 | 4.7e-05 | -1.3576 | 4.5e-05 | 248 | 2.876 | equi 6/11-15/2015 (6d) w/SW, cal
| 112 | 112 | 9757 | 7180 | 729 | 28 | bPolyPro | 28 | CF Hull |  | 4/15-23/2015 | B | 2@SO | 0.12 | [6 empty] | -0.0010403 | -1.3667 | 0.9999 |  |  | -8.5558e-22 | 6.7659e-18 | -2.1142e-14 | 3.2651e-11 | -2.367e-08 | 2.4053e-06 | -1.3668 | 0.9995 |  | '15 | 112 | 6/15-19/2015 |  | -1.3675 | 5.7e-05 | -1.3677 | 5.1e-05 | 187 | 0.818 | equi 6/11-15/2015 (6d) w/SW, cal
| 113 | 113 | 9766 | 7181 | 730 | 29 | bPolyPro | 29 | CF Hull |  | 4/23-30/2015 | A | 2@SO | 0.06 | [6 empty] | -0.0010708 | -1.379 | 0.9999 |  |  | 1.4822e-21 | -8.9012e-18 | 1.9178e-14 | -1.7484e-11 | 7.5067e-09 | -6.428e-06 | -1.379 | 0.99787 |  | '15 | 113 | 6/19-25/2015/sw |  | -1.3791 | 5.5e-05 | -1.3796 | 7.7e-05 | 486 | 0.123 | equi 6/11-19/2015 (8d) w/SW, cal
| 115 | 115 | 9761 | 7182 | 731 | 30 | bPolyPro | 30 | CF Hull |  | 4/30-5/7/2015 | A | 2@SO | 0.08 | [6 empty] | -0.0011117 | -1.3625 | 0.99989 |  |  | 2.0889e-21 | -1.2984e-17 | 2.9726e-14 | -3.0598e-11 | 1.5613e-08 | -8.7962e-06 | -1.3625 | 0.99754 |  | '15 | 115 | 6/19-25/2015/sw |  | -1.366 | 6.6e-05 | -1.3662 | 6.3e-05 | 166 | 3.552 | equi 6/11-19/2015 (8d) w/SW, cal
[36 empty] | 115 | 6/27-7/1/2015/TT |  | -1.3651 | 5.3e-05 | -1.3652 | 5.1e-05 | 145 |  | Rerun in TT water
[36 empty] | 115 | 8/7-27/2015/sw |  | -1.3654 | 4e-05 | -1.3658 | 4.9e-05 | 456 |  | Rerun in SW w/XP sensors
| 117 | 117 | 9265 | 5584 | 732 | [5 empty] | 5/7-14/2015 | A | 2@SO | 0.1 | [6 empty] | -0.0010398 | -1.3656 | 1 |  |  | -2.4915e-21 | 1.762e-17 | -4.9353e-14 | 6.9643e-11 | -4.9915e-08 | 1.1404e-05 | -1.3655 | 0.99875 |  | '15 | 117 | 9/2-11/2015/sw |  | -1.3685 | 0.000131 | -1.3688 | 0.000153 | 376 | 2.826 | equil 8/9-26/2015 (16d) w/SW
[37 empty] | 2/1-3/2016/sw |  | -1.3695 | 0.00013378 | -1.3697 | 0.00012026 | 241.8872 |  | recal after returned from Seattle w/NO3 issue
SBE 2015
| 94 | 94 | Stem |  |  |  | SB15- 1 | Sent to SB 3/27/15 |  |  | 11/13-28/2014 | A | 3T & full PT | 0.03 | [6 empty] | -0.0010716 | -1.3799 | 0.9987 |  |  | 1.0768e-21 | -6.8087e-18 | 1.668e-14 | -2.0831e-11 | 1.6455e-08 | -1.1125e-05 | -1.3802 | 0.99849 |  | '14 | 94 | No k0 cal at MBARI | [7 empty] | equil in PT and cal only
| 114 | 114 | Stem |  |  |  | SB15- 2 | Sent to SBE again |  |  | 4/23-30/2015 | B | 2@SO | 0.09 | [6 empty] | -0.0010796 | -1.374 | 0.9999 |  |  | 2.8078e-21 | -1.8647e-17 | 4.7631e-14 | -5.9751e-11 | 4.1434e-08 | -2.0161e-05 | -1.374 | 0.99738 |  | '15 | 114 | No k0 cal at MBARI | [7 empty] | equil in PT and cal only
| 116 | 116 | 0510 |  |  |  | SB15- 3 | Deliverd to SB 6/30/15 |  |  | 4/30-5/7/2015 | B | 2@SO | 0.13 | [6 empty] | -0.0010859 | -1.3616 | 0.9997 |  |  | -2.7692e-21 | 2.0308e-17 | -5.8233e-14 | 8.1859e-11 | -5.6275e-08 | 1.2661e-05 | -1.3618 | 0.99937 |  | '15 | 116 | No k0 cal at MBARI | [7 empty] | equil in PT and cal only
| 118 | 118 | 0564 |  |  |  | SB15- 4 | Deliverd to SB 8/2/15 |  |  | 5/7-14/2015 | B | 2@SO | -0.02 | [6 empty] | -0.0010589 | -1.3478 | 0.9999 |  |  | 8.3734e-22 | -4.9797e-18 | 1.0856e-14 | -1.1536e-11 | 1.0069e-08 | -1.1384e-05 | -1.3479 | 0.99965 |  | '15 | 118 | No k0 cal at MBARI | [7 empty] | equil ~1w + 8/26-xx/2015 (xxd) w/SW
| 119 | 119 | Stem |  |  |  | SB15- 5 | Deliverd to SB 8/2/15 |  |  | 5/14-20/2015 | A | 2@SO | 0.02 | [6 empty] | -0.0010931 | -1.3832 | 1 |  |  | 3.7991e-21 | -2.5009e-17 | 6.1854e-14 | -7.0183e-11 | 3.6516e-08 | -1.2063e-05 | -1.3831 | 0.99879 |  | '15 | 119 | No k0 cal at MBARI | [7 empty] | equil 8/9-26/2015 (16d) w/SW
| 120 | 120 | 0565 |  |  |  | SB15- 6 | Deliverd to SB 8/2/15 |  |  | 5/14-20/2015 | B | 2@SO | 0.03 | [6 empty] | -0.0010593 | -1.3643 | 1 |  |  | -1.587e-21 | 1.0876e-17 | -2.9087e-14 | 3.7881e-11 | -2.2431e-08 | 4.3109e-07 | -1.3643 | 0.99951 |  | '15 | 120 | No k0 cal at MBARI | [7 empty] | equil 8/9-26/2015 (16d) w/SW
| 121 | 121 | 0568 |  |  |  | SB15- 7 | Deliverd to SB 9/23/15 |  |  | 5/20-27/2015 | A | 2@SO | 0.08 | [6 empty] | -0.00099398 | -1.3757 | 0.9998 |  |  | -1.2103e-21 | 9.0166e-18 | -2.699e-14 | 4.1295e-11 | -3.1966e-08 | 6.2071e-06 | -1.3759 | 0.99961 |  | '15 | 121 | No k0 cal at MBARI | [7 empty] | equil 8/9-xx/2015 (xxd) w/SW
| 122 | 122 | 0567 |  |  |  | SB15- 8 | Deliverd to SB 9/23/15 |  |  | 5/20-27/2015 | B | 2@SO | -0.2 | [6 empty] | -0.0010483 | -1.3961 | 0.9998 |  |  | -2.9514e-21 | 1.7997e-17 | -4.1188e-14 | 4.3564e-11 | -1.7359e-08 | -5.13e-06 | -1.3963 | 0.99952 |  | '15 | 122 | No k0 cal at MBARI | [7 empty] | equil 8/9-xx/2015 (xxd) w/SW
| 123 | 123 | Stem |  |  |  | SB15- 9 | Deliverd to SB 9/23/15 |  |  | 6/10-17/2015 | A | 2@SO | 0.11 | [6 empty] | -0.001052 | -1.3798 | 0.9998 |  |  | -3.0102e-21 | 2.0492e-17 | -5.4218e-14 | 7.0323e-11 | -4.4725e-08 | 8.3117e-06 | -1.3799 | 0.99435 |  | '15 | 123 | No k0 cal at MBARI | [7 empty] | equil 8/26-xx/2015 (xxd) w/SW
| 124 | 124 | 0570 |  |  |  | SB15- 10 | Deliverd to SB 10/15 |  |  | 5/27-6/2/2015 | B | 2@SO | 0.08 | [6 empty] | -0.0010663 | -1.3631 | 0.9999 |  |  | 1.4861e-21 | -9.5659e-18 | 2.3634e-14 | -2.8834e-11 | 2.0632e-08 | -1.3212e-05 | -1.3632 | 0.99922 |  | '15 | 124 | No k0 cal at MBARI | [7 empty] | equil 8/26-xx/2015 (xxd) w/SW
| SBE 001 |  | 0571 | [7 empty] | Cal data from Charlie 11/23/16 | [9 empty] | -0.001019 | [6 empty] | -1.4746e-16 | -3.7267e-13 | 5.8315e-09 | -1.372e-05 | -1.3596 |  | Updated f(P) values
| SBE 103 |  | 0566 | [7 empty] | Cal data from Charlie 11/23/16 | [9 empty] | -0.0010387 | [6 empty] | 1.9394e-16 | -2.2056e-12 | 9.4529e-09 | -1.6767e-05 | -1.3845 |  | Updated f(P) values
Spare
| 125 | 125 | k2 cal | [7 empty] | 6/10-17/2015 | B | 2@SO | 0.02 | [6 empty] | -0.0010592 | -1.3669 | 1 |  |  | -1.8388e-21 | 1.3998e-17 | -4.1801e-14 | 6.1007e-11 | -4.2215e-08 | 7.2418e-06 | -1.367 | 0.99994 |  | '15 | 125
| 126 | 126 | k2 cal | [7 empty] | 6/17-7/6/2015 | A | 6@SO | 0.015 | [6 empty] | -0.0010547 | -1.3805 | 0.9999 |  |  | -1.1906e-22 | 2.7799e-18 | -1.4806e-14 | 3.2776e-11 | -3.1568e-08 | 6.9674e-06 | -1.3806 | 0.99816 |  | '15 | 126
| 127 | 127 | k2 cal | [7 empty] | 6/17-7/6/2015 | B | 6@SO | 0.045 | [6 empty] | -0.001093 | -1.3944 | 0.9995 |  |  | -2.8636e-21 | 2.111e-17 | -6.0642e-14 | 8.451e-11 | -5.5971e-08 | 1.0696e-05 | -1.3944 | 0.99769 |  | '15 | 127
SO 2016 |  |  |  |  | Beginning with #128, the sensors were assembled with the thicker covers.
| 97 | 97 | 12552 | 6737 | 713 |  | bPolyPro |  |  |  | 12/19-30/2014 | A | 3T & full PT | 0.24 | [6 empty] | -0.0011404 | -1.3969 | 0.9998 |  |  | 3.2519e-22 | -1.1243e-18 | -5.2261e-17 | 2.6352e-12 | 1.8297e-09 | -9.5541e-06 | -1.397 | 0.99964 | Test new electronics | '14 | 97 | 5/7-12/2015 |  | -1.3903 | 4.2e-05 | -1.3906 | 3.8e-05 | 265 |  | eqil 4/15-5/4/2015 (19d) in TT, cal | Unusually light sensitive, spikes removed .
| 101 | 101 | 12382 | 7183 | 717 |  | bPolyPro |  |  |  | 2/26-3/2/2015 | A | 1@SO | 0.21 | [6 empty] | -0.0011234 | -1.3791 | 0.9996 |  |  | 7.7316e-23 | 1.0712e-19 | -2.4407e-15 | 6.15e-12 | -3.1218e-09 | -6.2479e-06 | -1.3793 | 0.99991 | Test new electronics/returned | '14 | 101 | 5/31-6/2/2016 |  | -1.383 | 0.0001231 | -1.3832 | 0.00012627 | 191.3557 | 3.9751 | eqil 4/15-5/4/2015 (19d) in TT, cal
| 102 | 102 | 12545 | 7973 | 718 |  | bPolyPro |  |  |  | 3/6-14/2015 | A | 2@SO | 0.07 | [6 empty] | -0.0011358 | -1.3515 | 0.9996 |  |  | 3.449e-22 | -1.5413e-18 | 1.4334e-15 | 1.6807e-12 | -5.6651e-10 | -6.9858e-06 | -1.3516 | 0.99985 | '15 DF frm bad S ctd | '14 | 102 | 3/27-30/2016 |  | -1.348 | 4.1694e-05 | -1.3484 | 4.0133e-05 | 447.099 | -3.5623 | recycled 2015 DF from CTD w/bad S ctd
| 106 | 106 | 9642 | 6736 | 723 |  | bPolyPro | Now dep w/reg flowcell |  |  | 3/21-30/2015 | A | 2@SO | 0 | [6 empty] | -0.0010932 | -1.3775 | 0.9998 |  |  | -1.1265e-22 | 9.6611e-19 | -3.0046e-15 | 3.5867e-12 | 8.7202e-10 | -6.3085e-06 | -1.3774 | 0.99785 | '15 return | '15 | 106 | 4/11-14/2016 |  | -1.3804 | 6.7827e-05 | -1.3811 | 5.8981e-05 | 699.843 | 2.944 | 2015 return; new pH board | Now dep w/reg flowcell
| 128 | 128 | 9634 | 7913 | 733 |  | bPolyPro | First w/thicker cover |  |  | 12/23-31/2015 | A | 1@SO | 0 | [6 empty] | -0.0010451 | -1.3584 | 0.99991 |  |  | 3.6488e-21 | -2.5804e-17 | 7.1472e-14 | -9.887e-11 | 7.5474e-08 | -3.6401e-05 | -1.3585 | 0.99935 | '16>'15 swap | '15 | 128 | 1/15-19/2016 |  | -1.3619 | 5.6485e-05 | -1.362 | 5.0806e-05 | 116.8669 | 3.5027
| 129 | 129 | 9752 | 7945 | 734 |  | bPolyPro |  |  |  | 12/23-31/2015 | B | 1@SO | 0.06 | [6 empty] | -0.00099705 | -1.3663 | 0.99996 |  |  | 2.3738e-21 | -1.4854e-17 | 3.4519e-14 | -3.733e-11 | 2.3499e-08 | -1.7017e-05 | -1.3665 | 0.99962 | '16>'15 swap | '15 | 129 | 1/15-19/2016 |  | -1.3714 | 6.6139e-05 | -1.3715 | 8.2236e-05 | 77.0277 | 5.0319
| 130 | 130 | 12381 | 7915 | 735 |  | bPolyPro |  |  |  | 12/31-1/6/2016 | A | 1@SO | 0 | [6 empty] | -0.0010063 | -1.3691 | 0.99984 |  |  | 4.2633e-21 | -2.9203e-17 | 7.7318e-14 | -1.009e-10 | 7.238e-08 | -3.4799e-05 | -1.3692 | 0.99967 |  | '15 | 130 | 2/1-4/2016 |  | -1.3728 | 9.2174e-05 | -1.3728 | 9.9681e-05 | -10.611 | 3.7459
| 132 | 132 | 12551 | 7948 | 737 |  | bPolyPro |  |  |  | 1/7-13/2016 | A | 1@SO | 0.02 | [6 empty] | -0.0010786 | -1.3572 | 0.99997 |  |  | 9.704e-22 | -6.2716e-18 | 1.6076e-14 | -2.2336e-11 | 2.1904e-08 | -1.7949e-05 | -1.3573 | 0.99981 |  | '15 | 132 | 2/29-3/3/2016 |  | -1.3613 | 4.1727e-05 | -1.3616 | 4.5365e-05 | 272.4758 | 4.0782
| 133 | 133 | 12540 | 7982 | 746 |  | bPolyPro |  |  |  | 1/7-13/2016 | B | 1@SO | 0 | [6 empty] | -0.0010462 | -1.3732 | 0.99996 |  |  | 4.6356e-21 | -3.1735e-17 | 8.385e-14 | -1.0877e-10 | 7.6845e-08 | -3.5523e-05 | -1.3733 | 0.99928 |  | '15 | 133 | 3/11-16/2016 |  | -1.3766 | 4.3442e-05 | -1.3767 | 3.9696e-05 | 35.7486 | 3.4642
[36 empty] | 133 | 7/29-8/3/2016 |  | -1.3777 | 4.9391e-05 | -1.3777 | 4.6011e-05 | 51.4763
| 134 | 134 | 12541 | 7974 | 740 |  | bPolyPro |  |  |  | 1/13-19/2016 | A | 1@SO | 0 | [6 empty] | -0.0010604 | -1.3906 | 0.99998 |  |  | 1.7058e-21 | -1.1597e-17 | 3.0447e-14 | -4.0064e-11 | 3.2677e-08 | -2.1942e-05 | -1.3907 | 0.99951 |  | '15 | 134 | 3/11-16/2016 |  | -1.3932 | 4.2763e-05 | -1.3934 | 4.3868e-05 | 111.5975 | 2.6011
| 135 | 135 | 12543 | 7977 | 741 |  | bPolyPro |  |  |  | 1/13-19/2016 | B | 1@SO | 0.03 | [6 empty] | -0.0010776 | -1.3482 | 0.99971 |  |  | 3.9412e-21 | -2.7841e-17 | 7.7183e-14 | -1.0725e-10 | 8.0892e-08 | -3.5903e-05 | -1.3483 | 0.99985 |  | '15 | 135 | 3/30-4/6/2016 |  | -1.3511 | 4.2138e-05 | -1.3514 | 4.2048e-05 | 249.0404 | 2.9319
| 136 | 136 | 12390 | 7983 | 742 |  | bPolyPro |  |  |  | 1/19-28/2016 | A | 1@SO | 0.22 | [6 empty] | -0.0010236 | -1.3724 | 0.99997 |  |  | 1.2085e-21 | -7.8018e-18 | 2.0035e-14 | -2.7526e-11 | 2.544e-08 | -1.9871e-05 | -1.3725 | 0.99938 |  | '15 | 136 | 3/30-4/6/2016 |  | -1.3836 | 4.9774e-05 | -1.3837 | 5.6602e-05 | 105.4985 | 11.1784
| 137 | 137 | 12573 | 7984 | 743 |  | bPolyPro |  |  |  | 1/19-28/2016 | B | 1@SO | 0 | [6 empty] | -0.0010677 | -1.3935 | 0.99997 |  |  | 1.2903e-21 | -9.2548e-18 | 2.7146e-14 | -4.311e-11 | 4.3108e-08 | -3.0019e-05 | -1.3936 | 0.99988 |  | '15 | 137 | 4/8-14/2016 |  | -1.3978 | 7.4552e-05 | -1.3979 | 7.2336e-05 | 112.5234 | 4.3164
| 138 | 138 | 12575 | 8064 | 744 |  | bPolyPro |  |  |  | 1/28-2/5/2016 | A | 1@SO | 0 | [6 empty] | -0.0010292 | -1.3696 | 0.99997 |  |  | 3.4955e-21 | -2.3033e-17 | 5.7802e-14 | -7.0061e-11 | 4.6755e-08 | -2.3936e-05 | -1.3697 | 0.99961 |  | '15 | 138 | 4/8-14/2016 |  | -1.3735 | 4.2025e-05 | -1.3736 | 4.1081e-05 | 108.6086 | 3.9451
| 139 | 139 | 12559 | 8065 | 745 |  | bPolyPro |  |  |  | 1/28-2/5/2016 | B | 1@SO | 0.02 | [6 empty] | -0.0010486 | -1.3746 | 0.99998 |  |  | 7.5059e-22 | -4.3683e-18 | 9.0872e-15 | -8.1606e-12 | 4.1553e-09 | -4.2201e-06 | -1.3747 | 0.99881 |  | '15 | 139 | 4/8-14/2016 |  | -1.3789 | 4.2346e-05 | -1.3791 | 4.3136e-05 | 193.8901 | 4.2353
| 140 | 140 | Bad CTD, thus reasigned as for Cal Current
| 141 | 141 | 12542 | 8059 | 748 |  | bPolyPro |  |  |  | 2/5-12/2016 | B | 1@SO | -0.04 | [6 empty] | -0.001012 | -1.3622 | 0.99993 |  |  | 5.1902e-21 | -3.4932e-17 | 8.9158e-14 | -1.0731e-10 | 6.5033e-08 | -2.5942e-05 | -1.3623 | 0.99887 |  | '15 | 141 | 4/19-22/2016 |  | -1.3665 | 4.723e-05 | -1.3665 | 4.927e-05 | 57.5312 | 4.3041
| 142 | 142 | 12549 | 8062 | 749 |  | bPolyPro |  |  |  | 2/12-18/2016 | A | 1@SO | 0 | [6 empty] | -0.0010243 | -1.3834 | 0.9999 |  |  | 7.642e-22 | -4.9678e-18 | 1.228e-14 | -1.5621e-11 | 1.5462e-08 | -1.6409e-05 | -1.3835 | 0.99996 |  | '15 | 142 | 4/19-22/2016 |  | -1.386 | 5.094e-05 | -1.3861 | 5.1799e-05 | 97.9978 | 2.5805
'16 die | 143 | 143 | 8501 | 5590 | 702 |  | bPolyPro |  |  |  | 2/12-18/2016 | B | 1@SO | 0.05 | [6 empty] | -0.001006 | -1.3813 | 0.99996 |  |  | 6.837e-22 | -5.0119e-18 | 1.4447e-14 | -2.1333e-11 | 2.006e-08 | -1.5366e-05 | -1.3813 | 0.99992 | Start with 2016 Die | '16 | 143 | 4/19-22/2016 |  | -1.3849 | 5.4159e-05 | -1.385 | 5.4951e-05 | 35.6654 | 3.6843 | First '16 die; Was MB-8501
| 144 | 144 | 12558 | 7946 | 750 |  | bPolyPro |  |  |  | 2/23-3/1/2016 | A | 1@SO | 0.05 | [6 empty] | -0.0009331 | -1.3798 | 0.99995 |  |  | 1.9365e-24 | 6.9262e-19 | -4.096e-15 | 7.8682e-12 | -3.2047e-09 | -6.2384e-06 | -1.3799 | 0.99976 |  | '16 | 144 | 4/25-28/2016 |  | -1.383 | 4.3102e-05 | -1.3832 | 4.4902e-05 | 203.5626 | 3.2157
| 145 | 145 | 12537 | 7975 | 751 |  | bPolyPro |  |  |  | 2/23-3/1/2016 | B | 1@SO | 0.05 | [6 empty] | -0.0010117 | -1.395 | 0.9999 |  |  | 3.0728e-21 | -2.0865e-17 | 5.383e-14 | -6.6596e-11 | 4.4741e-08 | -2.2481e-05 | -1.3951 | 0.99963 |  | '16 | 145 | 4/25-5/2/2016 |  | -1.3988 | 4.1128e-05 | -1.3989 | 4.1318e-05 | 101.6199 | 3.8022
| 146 | 146 | 12379 | 7949 | 747 |  | bPolyPro |  |  |  | 3/1-10/2016 | A | 1@SO | 0.06 | [6 empty] | -0.00096272 | -1.383 | 0.99995 |  |  | 1.7468e-21 | -1.2692e-17 | 3.5337e-14 | -4.7655e-11 | 3.6073e-08 | -2.1916e-05 | -1.3831 | 0.9998 |  | '16 | 146 | 4/25-5/2/2016 |  | -1.3856 | 4.3095e-05 | -1.3864 | 4.6132e-05 | 811.133 | 2.5538
| 147 | 147 | 12388 | 8068 | 753 |  | bPolyPro |  |  |  | 3/1-10/2016 | B | 1@SO | 0.06 | [6 empty] | -0.00096864 | -1.3911 | 0.99996 |  |  | -3.2229e-21 | 2.1871e-17 | -5.6456e-14 | 6.748e-11 | -3.3524e-08 | -2.0375e-07 | -1.3912 | 0.99862 |  | '16 | 147 | 5/4-10/2016 |  | -1.3947 | 4.3323e-05 | -1.3949 | 3.9974e-05 | 129.759 | 3.5906
| 148 | 148 | 12386 | 8063 | 754 |  | bPolyPro |  |  |  | 3/16-23/2016 | A | 1@SO | 0.03 | [6 empty] | -0.000935 | -1.3798 | 0.99993 |  |  | 5.2351e-21 | -3.5812e-17 | 9.4369e-14 | -1.2018e-10 | 7.8862e-08 | -3.1264e-05 | -1.3799 | 0.99898 |  | '16 | 148 | 5/12-17/2016 |  | -1.3825 | 5.1365e-05 | -1.3827 | 6.6822e-05 | 143.6265 | 2.7005
| 149 | 149 | 12366 | 8066 | 755 |  | bPolyPro |  |  |  | 3/16-23/2016 | B | 1@SO | 0.02 | [6 empty] | -0.00097806 | -1.3631 | 0.99994 |  |  | 5.7822e-21 | -3.9517e-17 | 1.0347e-13 | -1.2951e-10 | 8.14e-08 | -3.0938e-05 | -1.3632 | 0.99924 |  | '16 | 149 | 5/12-17/2016 |  | -1.3663 | 6.5181e-05 | -1.3668 | 6.0959e-05 | 424.9276 | 3.2779
| 150 | 150 | 12370 | 8067 | 756 |  | bPolyPro |  |  |  | 3/23-30/2016 | A | 1@SO | 0.08 | [6 empty] | -0.00092445 | -1.3858 | 0.99981 |  |  | 9.1462e-22 | -6.605e-18 | 1.8888e-14 | -2.8077e-11 | 2.598e-08 | -1.84e-05 | -1.3859 | 0.99995 |  | '16 | 150 | 5/12-17/2016 |  | -1.39 | 5.7597e-05 | -1.3902 | 5.8258e-05 | 170.8243 | 4.2167
| 151 | 151 | 12378 | 8069 | 757 |  | bPolyPro |  |  |  | 3/23-30/2016 | B | 1@SO | 0 | [6 empty] | -0.00094419 | -1.3537 | 0.99975 |  |  | -3.2933e-21 | 2.3046e-17 | -6.2761e-14 | 8.2944e-11 | -5.1676e-08 | 7.9607e-06 | -1.3538 | 0.99844 |  | '16 | 151 | 5/12-17/2016 |  | -1.3579 | 6.1073e-05 | -1.358 | 5.2352e-05 | 111.3545 | 4.1753
| 152 | 152 | 9750 | 7184 | 758 |  | bPolyPro |  |  |  | 3/31-46/2016 | A | 1@SO | 0 | [6 empty] | -0.00099822 | -1.3767 | 0.99989 |  |  | -3.9798e-21 | 2.6865e-17 | -6.9506e-14 | 8.5884e-11 | -5.1006e-08 | 1.2186e-05 | -1.3767 | 0.88197 |  | '16 | 152 | 5/18-23/2016 |  | -1.3814 | 5.0796e-05 | -1.3815 | 4.8861e-05 | 165.0922 | 4.6801
| 153 | 153 | 12361 | 7877 | 759 |  | bPolyPro |  |  |  | 3/31-4/6/2016 | B | 1@SO | 0 | [6 empty] | -0.00099606 | -1.3658 | 0.99994 |  |  | 3.6665e-21 | -2.5146e-17 | 6.5686e-14 | -8.2944e-11 | 5.7335e-08 | -2.4943e-05 | -1.3658 | 0.99885 |  | '16 | 153 | 5/19-26/2016 |  | -1.3757 | 5.1902e-05 | -1.3762 | 4.8076e-05 | 436.9525 | 9.9663
| 154 | 154 | 12369 | 7943 | 760 |  | bPolyPro |  |  |  | 4/8-25/2016 | A | 1@SO | 0.06 | [6 empty] | -0.00093287 | -1.3713 | 0.99991 |  |  | 3.4524e-21 | -2.3274e-17 | 5.9606e-14 | -7.3036e-11 | 4.7797e-08 | -2.2665e-05 | -1.3714 | 0.99948 |  | '16 | 154 | 5/19-26/2016 |  | -1.3755 | 6.278e-05 | -1.3758 | 6.4277e-05 | 269.784 | 4.212
| 155 | 155 | 12363 | 7944 | 761 |  | bPolyPro |  |  |  | 4/8-15/2016 | B | 1@SO | 0.06 | [6 empty] | -0.0010027 | -1.3604 | 0.99996 |  |  | 2.9617e-21 | -2.0186e-17 | 5.3023e-14 | -6.8723e-11 | 5.0017e-08 | -2.5492e-05 | -1.3605 | 0.99956 |  | '16 | 155 | 5/19-26/2016 |  | -1.3643 | 4.0836e-05 | -1.3644 | 4.1529e-05 | 91.3874 | 3.965
| 156 | 156 | 12380 | 7947 | 762 |  | bPolyPro |  |  |  | 4/25-5/2/2016 | A | 1@SO | 0.05 | [6 empty] | -0.00098636 | -1.3997 | 0.9999 |  |  | 1.9238e-21 | -1.3172e-17 | 3.5366e-14 | -4.8402e-11 | 3.9139e-08 | -2.2918e-05 | -1.3998 | 0.99961 |  | '16 | 156 | 6/16-21/2016 |  | -1.4035 | 4.7243e-05 | -1.4038 | 4.8688e-05 | 325.0377 | 3.7754
| 157 | 157 | 12398 | 8060 | 763 |  | bPolyPro |  |  |  | 4/25-5/2/2016 | B | 1@SO | 0.085 | [6 empty] | -0.0009432 | -1.3844 | 0.99995 |  |  | -1.0147e-21 | 7.7166e-18 | -2.3579e-14 | 3.6629e-11 | -2.8293e-08 | 5.6083e-06 | -1.3844 | 0.99945 |  | '16 | 157 | 6/16-21/2016 |  | -1.3883 | 4.4262e-05 | -1.3886 | 4.8874e-05 | 255.7324 | 3.9825
| 158 | 158 | 12396 | 8061 | 764 |  | bPolyPro |  |  |  | 5/4-11/2016 | A | 1@SO | 0.11 | [6 empty] | -0.00097026 | -1.368 | 0.99987 |  |  | 4.1544e-21 | -2.811e-17 | 7.3903e-14 | -9.585e-11 | 6.6398e-08 | -2.8379e-05 | -1.3681 | 0.99946 |  | '16 | 158 | 6/16-21/2016 |  | -1.3715 | 4.7666e-05 | -1.3716 | 5.1119e-05 | 64.1366 | 3.4811
| 159 | 159 | 12372 | 7976 | 765 |  | bPolyPro |  |  |  | 5/4-11/2016 | B | 1@SO | 0.09 | [6 empty] | -0.0010478 | -1.3832 | 0.99988 |  |  | 4.7858e-21 | -3.1858e-17 | 8.0279e-14 | -9.5731e-11 | 5.9302e-08 | -2.6571e-05 | -1.3833 | 0.99945 |  | '16 | 159 | 6/23-29/2016 |  | -1.387 | 3.982e-05 | -1.3871 | 4.7424e-05 | 90.1538 | 3.8225
| 160 | 160 | 12371 | 7978 | 766 |  | bPolyPro |  |  |  | 5/11-18/2016 | A | 1@SO | 0.07 | [6 empty] | -0.0010002 | -1.3747 | 0.99994 |  |  | 2.9516e-21 | -1.936e-17 | 4.8159e-14 | -5.7371e-11 | 3.7166e-08 | -1.8782e-05 | -1.3748 | 0.99955 |  | '16 | 160 | 6/23-29/2016 |  | -1.3792 | 5.234e-05 | -1.3794 | 0.00012352 | 217.9239 | 4.497 | Was Cal Current, now SOCCOM
[44 empty] | 3.4915 | Average delta k0
[44 empty] | 1.4309 | StdDev
MBARI (long-term testing)
LabTest | 131 | 131 |  | 5589 | 736 | [5 empty] | 12/31-1/6/2016 | B | 1@SO | -0.04 | [6 empty] | -0.0010271 | -1.3547 | 0.99987 |  |  | 6.9688e-21 | -4.7843e-17 | 1.2638e-13 | -1.614e-10 | 1.063e-07 | -4.0907e-05 | -1.3548 | 0.99863 | MBARI LT testing | '15 | 131 | 2/29-3/3/2016 |  | -1.3585 | 8.277e-05 | -1.3585 | 5.0577e-05 | 69.884 | 3.7991
Cal Curr |  |  | 12544 | 9206 |  | (CTD & engine in SEA) |  |  |  |  | 2/5-12/2016 | A | 1@SO | 0.04 | [6 empty] | -0.0010134 | -1.3533 | 0.99999 |  |  | -1.0855e-20 | 7.57e-17 | -2.0427e-13 | 2.6614e-10 | -1.6468e-07 | 3.1455e-05 | -1.3534 | 0.99926 |  | '15 | 139 | 5/4-10/2016 |  | -1.3568 | 7.1699e-05 | -1.357 | 4.8798e-05 | 199.9599 | 3.5028
Yui | 161 | 161 |  | (5589) | (736) | [5 empty] | 5/11-18/2016 | B | 1@SO | 0.03 | [6 empty] | -0.0010481 | -1.4061 | 0.99993 |  |  | 2.1459e-21 | -1.4909e-17 | 4.0316e-14 | -5.395e-11 | 4.0883e-08 | -2.4234e-05 | -1.4062 | 0.9997 | Assigned to CPF |  | 161 | 6/5-9/2017 |  | -1.4145 | 4.9987e-05 | -1.4145 | 4.6433e-05 | 33.8619 | 8.3866
| 162 | 162 | 12544 | 7914 | CTD being be recal'ed ?? |  |  | Partially grease filled | [27 empty] | 162
| 163 | 163 | Grease filled - drifty |  |  |  |  | Grease filled (DO NOT USE) linear f(P) but drifting | [27 empty] | 163
SO 2017 |  |  | (#164 on all with Pt wire) |  |  |  |  | (Ver1 = Std PEEK design, Ver2 = double O-ring groove, Ver3 = smaller wire holes for Pt)
| 164 | 164 | [5 empty] | Ver1(2016), Pt |  |  | 8/4-11/2016 | A | 1@SO | 0.08 | [6 empty] | -0.00094215 | -1.405 | 0.99999 |  |  | -5.4292e-21 | 3.7353e-17 | -9.9455e-14 | 1.2882e-10 | -8.2278e-08 | 1.9244e-05 | -1.405 | 0.99565 | These were run repeatedly | '16 | 164 | NA
| 164 | 164 | 12728 | 9019 | 767 | [5 empty] | 10/26-11/1/2016 | A | 1@SO | -0.02 | [6 empty] | -0.0010492 | -1.4061 | 0.99987 |  |  | -4.0357e-21 | 2.8213e-17 | -7.7198e-14 | 1.0509e-10 | -7.3718e-08 | 2.1625e-05 | -1.4061 | 0.99543 | UW broke pH PCB | '16 | 164 | 1/17-25/2017 |  | -1.4076 | 6.8408e-05 | -1.4077 | 6.807e-05 | 131.8385 | 1.4313
TPOS | [21 empty] | -1.4061 | [11 empty] | New pH board | '16 | 164 | 3/25-30/2017 |  | -1.4075 | 5.4639e-05 | -1.4076 | 5.5601e-05 | 168.6496 | 1.3273
| 165 | 165 | 12717 | 9205 | 768 | ` |  | Ver1(2016), Pt |  |  | 8/4-11/2016 | B | 1@SO | 0.06 | [6 empty] | -0.0009627 | -1.3712 | 0.99996 |  |  | 1.0346e-21 | -6.7421e-18 | 1.6533e-14 | -1.9234e-11 | 1.2966e-08 | -9.0764e-06 | -1.3713 | 0.99967 | These were run repeatedly | '16 | 165 | NA
TPOS | 165 | 165 |  |  |  | ` |  |  |  |  | 9/21-28/2016 | B | 1@SO | -0.04 | [6 empty] | -0.0010014 | -1.3706 | 0.99994 |  |  | 1.8396e-22 | -9.419e-19 | 1.3756e-15 | -4.5349e-13 | 2.177e-09 | -6.6806e-06 | -1.3706 | 0.99978 |  | '16 | 165 | 1/17-25/2017 |  | -1.3746 | 6.9648e-05 | -1.3749 | 6.5925e-05 | 295.8967 | 4.095
1 | 166 | 166 | 12781 | 9015 | 769 |  |  | Ver1(2016), Pt |  |  | 11/18-25/2016 | A | 1@SO | 0.05 | [6 empty] | -0.00097424 | -1.4006 | 0.99998 |  |  | -1.1107e-21 | 7.4471e-18 | -1.9304e-14 | 2.3382e-11 | -1.0406e-08 | -3.4032e-06 | -1.4006 | 0.99979 |  | '16 | 166 | 1/17-25/2017 |  | -1.4032 | 8.4717e-05 | -1.4033 | 6.9433e-05 | 90.6557 | 2.5861
2 | 167 | 167 | 12779 | 9016 | 770 |  |  | Ver1(2017), Pt |  |  | 11/25-12/1/2016 | A | 1@SO | 0.02 | [6 empty] | -0.0010368 | -1.379 | 0.99992 |  |  | -7.2925e-22 | 4.8698e-18 | -1.2129e-14 | 1.3014e-11 | -3.1556e-09 | -5.1585e-06 | -1.3791 | 0.99991 |  | '16 | 167 | 1/17-25/2017 |  | -1.3811 | 8.9457e-05 | -1.3812 | 8.7952e-05 | 69.1707 | 2.1324
3 | 168 | 168 | 12784 | 9017 | 771 |  |  | Ver1(2017), Pt |  |  | 11/25-12/1/2016 | B | 1@SO | 0.06 | [6 empty] | -0.00098854 | -1.3913 | 0.99985 |  |  | -8.4225e-22 | 6.1864e-18 | -1.7969e-14 | 2.5551e-11 | -1.6239e-08 | 4.5634e-07 | -1.3913 | 0.99964 |  | '16 | 168 | 2/15-24/2017 |  | -1.3943 | 5.2036e-05 | -1.3943 | 5.2778e-05 | 77.2844 | 2.9672
4 | 169 | 169 | 12755 | 9025 | 778 | Equilibrating |  | Ver1(2017), Pt |  |  | 12/1-8/2016 | A | 1@SO | 0.07 | [6 empty] | -0.00099312 | -1.3949 | 0.99997 |  |  | -8.1867e-23 | 5.8151e-19 | -1.488e-15 | 9.5892e-13 | 3.0915e-09 | -6.3293e-06 | -1.395 | 0.99985 | Had open CE; Fixed. | '16 | 169 | 3/9-16/2017 |  | -1.3985 | 4.2565e-05 | -1.3987 | 4.8614e-05 | 136.0419 | 3.6071
5 | 170 | 170 | 12769 | 9018 | 772 |  |  | Ver1(2017), Pt |  |  | 12/1-8/2016 | B | 1@SO | 0.07 | [6 empty] | -0.0010069 | -1.3837 | 0.99996 |  |  | -1.1507e-21 | 8.0357e-18 | -2.169e-14 | 2.7707e-11 | -1.4844e-08 | -1.0924e-06 | -1.3838 | 0.99991 |  | '16 | 170 | 2/15-24/2017 |  | -1.3876 | 4.4503e-05 | -1.3878 | 4.2934e-05 | 201.515 | 3.8571
6 | 171 | 171 | 12782 | 9020 | 773 |  |  | Ver1(2017), Pt |  |  | 12/8-14/2016 | A | 1@SO | 0.03 | [6 empty] | -0.00099614 | -1.3833 | 0.99996 |  |  | 6.5256e-23 | -5.2304e-20 | -1.7387e-15 | 5.7605e-12 | -4.2842e-09 | -4.2049e-06 | -1.3834 | 0.99983 |  | '16 | 171 | 2/15-24/2017 |  | -1.3857 | 4.2747e-05 | -1.3859 | 4.4951e-05 | 219.4075 | 2.3753
| 172 | 172 | k0 offset | 9021 | 774 | (now #176) |  | Ver1(2017), Pt |  |  | 12/8-14/2016 | B | 1@SO | -0.15 | [6 empty] | -0.001024 | -1.4477 | 0.99992 |  |  | 3.4232e-22 | -2.0148e-18 | 3.9741e-15 | -2.8081e-12 | 1.9558e-09 | -5.3771e-06 | -1.4478 | 0.99956 |  | '16 | 172 | 3/2-8/2017 |  | -1.4335 | 7.1155e-05 | -1.4336 | 7.2141e-05 | 129.8883 | -14.2599
7 | 173 | 173 | 12778 | 9023 | 775 | Equilibrating |  | Ver1(2017), Pt |  |  | 12/14-1/3/2017 | A | 1@SO | 0.02 | [6 empty] | -0.00097474 | -1.3725 | 0.99999 |  |  | -3.4432e-22 | 3.4643e-18 | -1.2817e-14 | 2.1225e-11 | -1.326e-08 | -2.5486e-06 | -1.3726 | 0.9996 |  | '16 | 173 | 3/3-8/2017 |  | -1.3762 | 6.5783e-05 | -1.3763 | 6.2096e-05 | 143.3175 | 3.6557
8 | 174 | 174 | 12758 | 9206 | 779 | Eqil 2/27 |  | Ver1(2017), Pt |  |  | 12/14-1/3/2017 | B | 1@SO | 0.04 | [6 empty] | -0.00098781 | -1.4308 | 0.99991 |  |  | -8.0605e-23 | 6.4405e-19 | -2.0707e-15 | 2.5627e-12 | 1.7162e-09 | -6.977e-06 | -1.4308 | 0.99976 | High CE resis; Fixed. | '16 | 174 | 3/9-16/2017 |  | -1.434 | 6.7904e-05 | -1.4343 | 7.2659e-05 | 282.3074 | 3.2605
9 | 175 | 175 | 12757 | 9024 | 776 | Equilibrating |  | Ver2-PEEK, Pt |  |  | 1/3-10/2017 | A | 1@SO | 0.14 | [6 empty] | -0.00096652 | -1.4231 | 0.99995 |  |  | 1.5148e-21 | -1.1026e-17 | 3.1783e-14 | -4.6432e-11 | 3.6174e-08 | -1.3337e-05 | -1.4231 | 0.98766 |  | '16 | 175 | 3/9-16/2017 |  | -1.4275 | 6.8815e-05 | -1.4276 | 8.9271e-05 | 108.7727 | 4.3787
10 | 176 | 176 |  | 9021 | 774 | Equilibrating |  | Ver2-Ultem, Pt |  |  | 1/3-10/2017 | B | 1@SO | 0 | [6 empty] | -0.0010056 | -1.3683 | 0.99988 |  |  | 3.0415e-21 | -2.1091e-17 | 5.606e-14 | -7.2103e-11 | 4.8526e-08 | -1.7367e-05 | -1.3684 | 0.99713 |  | '16 | 176 | 3/9-16/2017 |  | -1.3715 | 4.1404e-05 | -1.3716 | 3.7354e-05 | 53.0798 | 3.1968 | not used - Ultem material
11 | 177 | 177 | 12744 | 9268 | 780 | Eqil 2/27 |  | Ver2-Ultem, Pt |  |  | 1/11-17/2017 | A | 1@SO | 0.1 | [6 empty] | -0.00095239 | -1.4298 | 0.99995 |  |  | 1.4616e-21 | -1.01e-17 | 2.7537e-14 | -3.9591e-11 | 3.5331e-08 | -1.8218e-05 | -1.4298 | 0.99921 | Open CE; Fixed. | '16 | 177 | 3/21-24/2017 |  | -1.4328 | 5.4894e-05 | -1.4329 | 5.5727e-05 | 114.2338 | 3.0452
12 | 178 | 178 | 12729 | 9269 | 781 | Equil 10d |  | Ver2-Ultem, Pt |  |  | 1/11-17/2017 | B | 1@SO | 0.08 | [6 empty] | -0.00099001 | -1.4073 | 0.99989 |  |  | 2.7853e-22 | -1.4732e-18 | 3.0164e-15 | -5.7253e-12 | 1.2516e-08 | -1.1817e-05 | -1.4074 | 0.99729 | High CE Ω; Fixed. | '16 | 178 | 3/21-24/2017 |  | -1.4107 | 5.2064e-05 | -1.4108 | 5.7653e-05 | 143.2667 | 3.3579
| 179 | 179 | DAMAGED During Stem Assy |  |  |  |  | Ver2-PEEK, Pt |  |  | 1/17-24/2017 | A | 1@SO | 0.1 | [6 empty] | -0.00098359 | -1.4106 | 1 |  |  | 1.5794e-21 | -1.2263e-17 | 3.7196e-14 | -5.5869e-11 | 4.3331e-08 | -1.3507e-05 | -1.4106 | 0.99932 | R&W wires pulled out | '16 | 179
| 180 | 180 | DAMAGED During Stem Assy |  |  |  |  | Ver2-PEEK, Pt |  |  | 1/17-24/2017 | B | 1@SO | 0.1 | [6 empty] | -0.0010477 | -1.4024 | 0.99997 |  |  | 5.2715e-22 | -5.0803e-18 | 1.7404e-14 | -2.7278e-11 | 2.1512e-08 | -8.4333e-06 | -1.4024 | 0.9986 | Wire pulled out of Ref during stem assy | '16 | 180
13 | 181 | 181 | 12748 | 9270 | 782 | Equil 10d |  | Ver2-PEEK, Pt |  |  | 1/24-31/2017 | A | 1@SO | 0.09 | [6 empty] | -0.00095276 | -1.4291 | 0.99997 |  |  | 1.294e-21 | -9.1032e-18 | 2.55e-14 | -3.6729e-11 | 2.9786e-08 | -1.4622e-05 | -1.4292 | 0.99954 |  | '16 | 181 | 3/21-24/2017 |  | -1.4341 | 4.4023e-05 | -1.4342 | 4.6999e-05 | 108.5278 | 4.9676
| 182x | 182 | DAMAGED During k2 testing |  |  |  |  | Ver2-PEEK, Pt |  |  | 1/24-31/2017 | B | 1@SO | 0.05 | [6 empty] | -0.00094872 | -1.4141 | 0.99986 |  |  | 2.5687e-22 | -2.3626e-18 | 7.9924e-15 | -1.3125e-11 | 1.2765e-08 | -8.5935e-06 | -1.4141 | 0.99955 | Vrs railed inb 3rd P cycle | '16 | 182x
14 | 182 | 182 | 12745 | 9279 | 783 |  |  | Ver2-PEEK, Pt |  |  | 2/3-10/2017 | A | 1@SO | -0.02 | [6 empty] | -0.0010108 | -1.4172 | 0.99995 |  |  | 1.2412e-21 | -1.0089e-17 | 3.214e-14 | -5.1345e-11 | 4.3348e-08 | -1.528e-05 | -1.4172 | 0.99908 |  | '16 | 182 | 3/25-30/2017 |  | -1.4193 | 5.0206e-05 | -1.4194 | 6.6057e-05 | 115.9037 | 2.0612
15 | 183 | 183 | 12733 | 9280 | 784 |  |  | Ver2-PEEK, Pt |  |  | 2/3-10/2017 | B | 1@SO | 0 | [6 empty] | -0.00099299 | -1.4024 | 0.99993 |  |  | 1.5178e-22 | -2.5223e-18 | 1.0959e-14 | -2.019e-11 | 1.8702e-08 | -8.9002e-06 | -1.4024 | 0.99901 |  | '17 | 183 | 3/25-30/2017 |  | -1.4051 | 8.8129e-05 | -1.4052 | 9.4579e-05 | 109.0257 | 2.7087
| 184 | 184 | Drifting, Offset |  |  |  |  | Ver2-PEEK, Pt |  |  | 2/10-21/2017 | A | 1@SO | [21 empty] | '17 | 184 | Rebuild
16 | 185 | 185 | 12742 | 9282 | 786 |  |  | Ver2-PEEK, Pt |  |  | 2/10-21/2017 | B | 1@SO | 0 | [6 empty] | -0.0010063 | -1.4127 | 0.99998 |  |  | 5.5013e-22 | -3.8886e-18 | 9.6864e-15 | -1.0269e-11 | 6.2042e-09 | -4.1016e-06 | -1.4127 | 0.99878 |  | '17 | 185 | 4/7-13/2017 |  | -1.4153 | 5.5913e-05 | -1.4155 | 5.823e-05 | 115.6455 | 2.6307
17 | 186 | 186 | 12723 | 9281 | 785 |  | PT-4k | Ver2-PEEK, Pt |  |  | 2/21-28/2017 | A | 1@SO | 0.04 | [6 empty] | -0.00098388 | -1.39 | 0.99998 |  |  | 1.1661e-21 | -8.1566e-18 | 2.1395e-14 | -2.6822e-11 | 1.926e-08 | -8.5612e-06 | -1.39 | 0.99872 |  | '17 | 186 | 3/25-30/2017 |  | -1.3926 | 7.7621e-05 | -1.3928 | 8.9498e-05 | 150.8851 | 2.6638
18 | 187 | 187 | 12709 | 9027 | 787 |  | PT-4k | Ver2-PEEK, Pt |  |  | 2/21-28/2017 | B | 1@SO | 0.02 | [6 empty] | -0.0010911 | -1.3685 | 0.9999 |  |  | 4.131e-22 | -3.5895e-18 | 1.2585e-14 | -2.2601e-11 | 2.2532e-08 | -1.2876e-05 | -1.3686 | 0.99961 |  | '17 | 187 | 4/7-13/2017 |  | -1.3717 | 4.9238e-05 | -1.3718 | 4.7408e-05 | 52.2818 | 3.2369
19 | 188 | 188 | 12702 | 9028 | 788 |  | PT-4k | Ver2-PEEK, Pt |  |  | 3/7-14/2017 | A | 1@SO | 0.14 | [6 empty] | -0.00095715 | -1.4176 | 0.99989 |  |  | 3.2389e-21 | -2.2596e-17 | 6.1158e-14 | -8.1293e-11 | 5.5296e-08 | -1.5717e-05 | -1.4176 | 0.99702 |  | '17 | 188 | 4/17-21/2017 |  | -1.4216 | 4.6274e-05 | -1.4218 | 4.6273e-05 | 145.4406 | 4.0906
20 | 189 | 189 | 12754 | 9029 | 789 |  | PT-4k | Ver2-PEEK, Pt |  |  | 3/7-14/2017 | B | 1@SO | 0.06 | [6 empty] | -0.0010577 | -1.3788 | 0.99993 |  |  | 1.2831e-21 | -7.6556e-18 | 1.6098e-14 | -1.4811e-11 | 8.9443e-09 | -4.8626e-06 | -1.3788 | 0.9974 |  | '17 | 189 | 4/17-21/2017 |  | -1.3818 | 5.3265e-05 | -1.382 | 5.6284e-05 | 162.6892 | 3.0468
21 | 190 | 190 | 12734 | 9287 | 792 |  | PT-4k | Ver2-PEEK, Pt |  |  | 3/14-21/2017 | A | 1@SO | 0.15 | [6 empty] | -0.0009736 | -1.4339 | 0.99994 |  |  | 1.4125e-22 | -1.7055e-18 | 7.4911e-15 | -1.5974e-11 | 1.8519e-08 | -1.1655e-05 | -1.4339 | 0.99935 |  | '17 | 190 | 5/8-11/2017 |  | -1.438 | 0.00010028 | -1.4381 | 8.7838e-05 | 168.5569 | 4.0607
22 | 191 | 191 | 12768 | 9030 | 790 |  | PT-4k | Ver2-PEEK, Pt |  |  | 3/14-21/2017 | B | 1@SO | 0.2 | [6 empty] | -0.0010114 | -1.4187 | 0.99993 |  |  | 6.2467e-23 | -8.3084e-19 | 3.7862e-15 | -9.2135e-12 | 1.3254e-08 | -6.9757e-06 | -1.4188 | 0.99768 |  | '17 | 191 | 4/17-21/2017 |  | -1.4218 | 5.2216e-05 | -1.422 | 5.3658e-05 | 166.0379 | 3.0422
23 | 192 | 192 | 12736 | 9031 | 791 |  | Pprog-4k | Ver2-PEEK, Pt |  |  | 3/21-27/2017 | A | 1@SO | 0.04 | [6 empty] | -0.0010622 | -1.3733 | 0.99996 |  |  | 7.0791e-22 | -6.7278e-18 | 2.4663e-14 | -4.5493e-11 | 4.5256e-08 | -2.0313e-05 | -1.3733 | 0.99652 |  | '17 | 192 | 4/17-21/2017 |  | -1.3764 | 4.1992e-05 | -1.3767 | 5.3385e-05 | 263.4494 | 3.1222
24 | 193 | 193 | 12747 | 9288 | 793 |  | Pprog-4k | Ver2-PEEK, Pt |  |  | 3/21-27/2017 | B | 1@SO | 0.075 | [6 empty] | -0.0010919 | -1.4108 | 0.99993 |  |  | -7.4531e-23 | -8.5768e-19 | 7.8597e-15 | -2.2777e-11 | 3.0434e-08 | -1.5696e-05 | -1.4108 | 0.99923 |  | '17 | 193 | 5/8-11/2017 |  | -1.4136 | 5.4549e-05 | -1.4138 | 5.8094e-05 | 221.8222 | 2.8012
| 194 | 194 | Railed Ik |  |  |  | Pprog-4k | Ver2-PEEK, Pt |  |  | 3/27-4/3/2017 | A | 1@SO | 0.13 | [6 empty] | -0.00099085 | -1.387 | 0.99996 |  |  | 2.7993e-22 | -3.1349e-18 | 1.2394e-14 | -2.3056e-11 | 2.3034e-08 | -1.3188e-05 | -1.387 | 0.99945 |  | '17 | 194 | [8 empty] | Railed+ Ik (Changed to DF#196)
25 | 195 | 195 | 12700 | 9293 | 795 |  | Pprog-4k | Ver2-PEEK, Pt |  |  | 3/27-4/3/2017 | B | 1@SO | 0.13 | [6 empty] | -0.0010138 | -1.3509 | 0.99992 |  |  | 1.0686e-22 | -1.9413e-18 | 7.9574e-15 | -1.3181e-11 | 1.1557e-08 | -6.0812e-06 | -1.3509 | 0.99889 |  | '17 | 195 | 5/8-11/2017 |  | -1.3533 | 6.0366e-05 | -1.3536 | 6.8381e-05 | 289.8444 | 2.4295
26 | 196 | 196 | 12701 | 9289 | 794 |  | Pprog-4k | Ver2-PEEK, Pt |  |  | 4/3-10/2017 | A | 1@SO | 0.11 | [6 empty] | -0.00095936 | -1.3799 | 0.99998 |  |  | 3.9839e-22 | -3.4142e-18 | 1.073e-14 | -1.613e-11 | 1.4908e-08 | -1.0855e-05 | -1.3799 | 0.99936 |  | '17 | 196 | 5/8-11/2017 |  | -1.3827 | 7.1852e-05 | -1.3828 | 6.2881e-05 | 103.3002 | 2.8244
27 | 197 | 197 | 12730 | 9032 | 796 |  | Pprog-4k | Ver2-PEEK, Pt |  |  | 4/3-10/2017 | B | 1@SO | 0.05 | [6 empty] | -0.0011014 | -1.3937 | 0.99994 |  |  | -1.9022e-21 | 1.18e-17 | -2.6293e-14 | 2.1837e-11 | 3.0509e-09 | -9.7167e-06 | -1.3937 | 0.99819 |  | '17 | 197 | 6/2-9/2017 |  | -1.3962 | 5.5583e-05 | -1.3965 | 5.547e-05 | 291.6602 | 2.5185
28 | 198 | 198 | 12741 | 9065 | 797 |  | Pprog-4k | Ver3-PEEK, Pt |  |  | 4/10-17/2017 | A | 1@SO | 0.09 | [6 empty] | -0.00097553 | -1.3799 | 0.99992 |  |  | -7.158e-22 | 5.0232e-18 | -1.298e-14 | 1.3594e-11 | -4.4363e-10 | -9.8957e-06 | -1.38 | 0.99975 |  | '17 | 198 | 6/2-9/2017 |  | -1.3826 | 6.7152e-05 | -1.3828 | 6.7156e-05 | 176.5806 | 2.6905
29 | 199 | 199 | 12787 | 7914 | 798 |  | Pprog-4k | Ver3-PEEK, Pt |  |  | 4/10-17/2017 | B | 1@SO | 0.01 | [6 empty] | -0.0010989 | -1.4012 | 0.99991 |  |  | -5.958e-22 | 2.8355e-18 | -2.5952e-15 | -7.3309e-12 | 1.8501e-08 | -1.5245e-05 | -1.4012 | 0.99967 |  | '17 | 199 | 6/2-9/2017 |  | -1.4053 | 5.3848e-05 | -1.4056 | 5.5925e-05 | 275.4604 | 4.164
30 | 200 | 200 |  |  |  |  | Pprog-4k | Ver3-PEEK, Pt |  |  | 4/17-24/2017 | A | 1@SO | 0.03 | [6 empty] | -0.0011028 | -1.4111 | 0.9999 |  |  | -2.8407e-21 | 1.8976e-17 | -4.812e-14 | 5.5897e-11 | -2.4491e-08 | -5.6704e-06 | -1.4112 | 0.9995 |  | '17
| 201 | 201 | 4k ok, FAILED during 2k cal | [31 empty] | '17
| 202 | 202
| 203 | 203
| 204 | 204 | 4k ok, Vrs and Vk weird >2d | [30 empty] | 4k ok, Vrs and Vk weird >2d | [9 empty] | 3.1259 | Average delta k0
[44 empty] | 0.77561 | StdDev
For CPF
| 208 | 208 |  |  |  |  | No 6k test | Ver3-PEEK, Pt |  |  | 6/30-7/7/2017 | D | 1@SO | 0.06 | [6 empty] | -0.00098166 | -1.4014 | 0.99996 |  |  | -7.9662e-22 | 5.6892e-18 | -1.5164e-14 | 1.7152e-11 | -3.5296e-09 | -8.3984e-06 | -1.4015 | 0.99993
SO 2018 | [7 empty] | (Ver1 = Std PEEK design, Ver2 = double O-ring groove, Ver3 = smaller wire holes for Pt, Ver4 = Ti CE using roll pin to connect to green wire, Ver5 = pellet oring changed to 6.5x1.5 Buna70A)
| 209 | 209 |  | Yui |  |  | sbe-3.5ksi | Ver3-PEEK, Pt; Yui 4003 |  |  | 10/16-24/2017 | C | 1@SO | 0.05 | [6 empty] | -0.0010762 | -1.3635 | 0.99994 |  |  | -2.5229e-22 | 4.5093e-19 | 3.8458e-15 | -1.5667e-11 | 2.3176e-08 | -1.5203e-05 | -1.3636 | 0.99948 | Yui 4003 | [10 empty] | gave to Yui without K0 run
TPOS | 210 | 210 | 12783 | 10226 | 800 |  | sbe-3.5ksi | Ver3-PEEK, Pt RUN#3 |  |  | 11/8-14/2017 | C | 1@SO | 0.15 | [6 empty] | -0.00094981 | -1.4256 | 0.99996 |  |  | 2.1134e-21 | -1.6118e-17 | 4.8087e-14 | -7.0467e-11 | 5.1801e-08 | -1.5609e-05 | -1.4256 | 0.99807 |  |  | 210 | 12/26-28/2017 |  | -1.4278 | 5.0922e-05 | -1.428 | 6.9317e-05 | 135.5219 | 2.2566
TPOS | 211 | 211 | 12792 | 10223 | 801 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/13-20/2017 | A | 1@SO | -0.006 | [6 empty] | -0.00097083 | -1.418 | 0.99982 |  |  | 8.8448e-22 | -6.2669e-18 | 1.7544e-14 | -2.5382e-11 | 2.1544e-08 | -1.0637e-05 | -1.4181 | 0.99843 | wape |  | 211 | 12/26-28/2017 |  | -1.4195 | 5.3897e-05 | -1.4197 | 8.571e-05 | 167.1945 | 1.5192
TPOS | 212 | 212 | 12780 | 10228 | 802 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/13-20/2017 | B | 1@SO | 0.015 | [6 empty] | -0.0010104 | -1.3997 | 0.99979 |  |  | 1.361e-21 | -1.081e-17 | 3.3669e-14 | -5.1612e-11 | 3.9796e-08 | -1.2997e-05 | -1.3997 | 0.99832 | wape |  | 212 | 12/26-29/2017 |  | -1.4026 | 4.8077e-05 | -1.4029 | 4.5785e-05 | 309.5407 | 2.8531
| 213 | 213 | 12707 | 10224 | 803 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/15-21/2017 | C | 1@SO | 0.17 | [6 empty] | -0.0010236 | -1.3895 | 0.99998 |  |  | 6.7713e-22 | -5.2514e-18 | 1.6392e-14 | -2.7091e-11 | 2.6662e-08 | -1.5688e-05 | -1.3896 | 0.99966 | jaha |  | 213 | 1/29-31/2018 |  | -1.393 | 6.5048e-05 | -1.3932 | 6.2293e-05 | 208.8609 | 3.5099
| 214 | 214 | 12752 | 10225 | 804 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/20-27/2017 | A | 1@SO | 0.06 | [6 empty] | -0.00097794 | -1.4343 | 0.99996 |  |  | 1.7426e-22 | -2.626e-18 | 1.2761e-14 | -2.916e-11 | 3.4783e-08 | -2.0647e-05 | -1.4343 | 0.99968 | jaha |  | 214 | 1/29-31/2018 |  | -1.4389 | 6.7595e-05 | -1.4392 | 7.5423e-05 | 283.9367 | 4.6301
| 215 | 215 | 12786 | 9206 | 774 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/20-27/2017 | B | 1@SO | 0.01 | [6 empty] | -0.0010058 | -1.4006 | 0.99983 |  |  | 4.4203e-22 | -4.7361e-18 | 1.9519e-14 | -3.9645e-11 | 4.1441e-08 | -2.2068e-05 | -1.4007 | 0.99933 | jaha |  | 215 | 3/5-6/2018 |  | -1.4051 | 6.2779e-05 | -1.4055 | 5.3084e-05 | 379.8583 | 4.4524
EMM | 216 | 216 | 8474 | 10599 | 805 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/27-12/4/2017 | A | 1@SO | 0.07 | [6 empty] | -0.00099348 | -1.4238 | 0.99995 |  |  | 8.6964e-22 | -7.2414e-18 | 2.499e-14 | -4.5249e-11 | 4.4333e-08 | -2.3552e-05 | -1.4239 | 0.99992 | jaha |  | 216 | 3/6-7/2018 |  | -1.4275 | 0.00012869 | -1.4277 | 0.00011603 | 263.1563 | 3.6421
EMM | 217 | 217 | 17810 | 10451 | 806 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 11/27-12/4/2017 | B | 1@SO | 0.08 | [6 empty] | -0.00096363 | -1.4214 | 0.99993 |  |  | 9.7473e-22 | -6.6844e-18 | 1.7364e-14 | -2.0892e-11 | 1.1271e-08 | -1.6287e-06 | -1.4214 | 0.99464 | jaha |  | 217 | 3/6-7/2018 |  | -1.4244 | 9.5695e-05 | -1.4249 | 0.00011633 | 507.7199 | 2.9994
| 218 | 218 | 12696 | 10571 | 807 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/4-11/2017 | A | 1@SO | 0.03 | [6 empty] | -0.00095009 | -1.4208 | 0.99997 |  |  | 1.8523e-21 | -1.201e-17 | 3.0089e-14 | -3.8374e-11 | 3.0392e-08 | -1.7938e-05 | -1.4209 | 0.99875 | jaha |  | 218 | 3/14-15/2018 |  | -1.4248 | 4.3022e-05 | -1.4249 | 4.7002e-05 | 98.1366 | 3.9629
| 219 | 219 | 12708 | 10450 | 810 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/4-11/2017 | B | 1@SO | 0.03 | [6 empty] | -0.00092547 | -1.408 | 0.99996 |  |  | 1.1062e-21 | -7.6709e-18 | 2.0773e-14 | -2.8149e-11 | 2.1336e-08 | -1.0647e-05 | -1.408 | 0.99926 | jaha |  | 219 | 3/26-28/2018 |  | -1.4111 | 6.1367e-05 | -1.4116 | 5.794e-05 | 411.3603 | 3.137
EMM | 220 | 220 | 12788 | 10452 | 809 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/11-18/2017 | A | 1@SO | 0.05 | [6 empty] | -0.00091994 | -1.4015 | 0.99993 |  |  | -7.4306e-22 | 5.0483e-18 | -1.2934e-14 | 1.4657e-11 | -5.1238e-09 | -3.2564e-06 | -1.4015 | 0.99877 | jaha |  | 220 | 3/14-15/2018 |  | -1.4156 | 5.784e-05 | -1.4158 | 0.00012114 | 205.2083 | 14.1203 | backup battery voltage low - K0 data noisy but ok
| 221 | 221 | 12727 | 10227 | 820 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/11-18/2017 | B | 1@SO | -0.03 | [6 empty] | -0.0010319 | -1.4044 | 0.99998 |  |  | -9.4361e-22 | 6.8967e-18 | -2.0088e-14 | 2.923e-11 | -1.9805e-08 | 9.016e-07 | -1.4045 | 0.9999 | jaha |  | 221 | 5/30-6/1/2018 |  | -1.4081 | 0.00010499 | -1.4082 | 7.6011e-05 | 93.8939 | 3.6646
| 222 | 222 | 12879 | 10370 | 811 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/19-26/2017 | A | 1@SO | -0.015 | [6 empty] | -0.0010325 | -1.3991 | 0.99997 |  |  | 4.8386e-22 | -4.0518e-18 | 1.3644e-14 | -2.4151e-11 | 2.5026e-08 | -1.5186e-05 | -1.3991 | 0.99991 | wape |  | 222 | 3/26-28/2018 |  | -1.4021 | 5.0894e-05 | -1.4028 | 4.3443e-05 | 675.7824 | 3.014
| 223 | 223 | 12883 | 10440 | 808 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/19-26/2017 | B | 1@SO | -0.01 | [6 empty] | -0.0010265 | -1.3792 | 0.99998 |  |  | 1.3177e-21 | -9.0164e-18 | 2.3519e-14 | -2.9309e-11 | 1.9717e-08 | -1.1672e-05 | -1.3793 | 0.99975 | jaha |  | 223 | 3/14-15/2018 |  | -1.3825 | 0.00013445 | -1.3826 | 0.00012322 | 81.8309 | 3.3144
| 224 | 224 | 12884 | 10441 | 812 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/26-1/2/2018 | A | 1@SO | 0.1 | [6 empty] | -0.00089253 | -1.4368 | 0.99992 |  |  | 6.0342e-22 | -4.5714e-18 | 1.4154e-14 | -2.3795e-11 | 2.4518e-08 | -1.5231e-05 | -1.4369 | 0.99986 | wape |  | 224 | 3/30-4/5/2018 |  | -1.4413 | 5.3573e-05 | -1.4433 | 6.0097e-05 | 2019.2741 | 4.4772
| 225 | 225 | 12880 | 10444 | 813 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 12/26-1/2/2018 | B | 1@SO | 0.09 | [6 empty] | -0.00099381 | -1.4021 | 0.99993 |  |  | -2.8605e-21 | 1.9236e-17 | -4.8633e-14 | 5.5148e-11 | -2.194e-08 | -7.4648e-06 | -1.4022 | 0.99972 | wape |  | 225 | 3/30-4/5/2018 |  | -1.4073 | 6.4196e-05 | -1.4078 | 0.00011865 | 544.7026 | 5.2114
| 226 | 226 | 12889 | 10445 | 814 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 1/24-31/2018 | A | 1@SO | 0.07 | [6 empty] | -0.0010166 | -1.3609 | 0.99998 |  |  | -5.2959e-22 | 3.9841e-18 | -1.1763e-14 | 1.6396e-11 | -7.6456e-09 | -7.2994e-06 | -1.361 | 0.99994 |  |  | 226 | 4/4-5/2018 |  | -1.366 | 8.3605e-05 | -1.3665 | 7.1633e-05 | 483.6685 | 5.0668
| 227 | 227 | 12881 | 10447 | 815 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 1/9-15/2018 | B | 1@SO | 0.095 | [6 empty] | -0.00094022 | -1.3931 | 0.99991 |  |  | 8.942e-22 | -6.2304e-18 | 1.7676e-14 | -2.7525e-11 | 2.8138e-08 | -2.2382e-05 | -1.3933 | 0.99985 |  |  | 227 | 4/16-17/2018 |  | -1.3989 | 4.5374e-05 | -1.3991 | 4.0661e-05 | 124.4571 | 5.8154
| 228 | 228 | 12886 | 10363 | 816 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 1/11-18/2018 | C | 1@SO | 0.095 | [6 empty] | -0.0010971 | -1.3984 | 0.99995 |  |  | -1.4897e-21 | 1.0022e-17 | -2.5375e-14 | 2.8771e-11 | -1.0867e-08 | -5.5399e-06 | -1.3984 | 0.99994 |  |  | 228 | 4/16-17/2018 |  | -1.4034 | 5.2523e-05 | -1.4038 | 5.8179e-05 | 364.4701 | 5.0587
| 229 | 229 | 12878 | 10364 | 817 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 1/11-18/2018 | D | 1@SO | 0.06 | [6 empty] | -0.0010922 | -1.3966 | 0.99993 |  |  | 1.4445e-21 | -1.0656e-17 | 3.0846e-14 | -4.4242e-11 | 3.3006e-08 | -1.2225e-05 | -1.3966 | 0.99729 |  |  | 229 | 4/16-17/2018 |  | -1.4006 | 5.3601e-05 | -1.4013 | 8.4545e-05 | 730.6177 | 3.9984
| 230 | 230 | 17307 | 10454 | 818 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 1/24-31/2018 | B | 1@SO | 0.125 | [6 empty] | -0.0009618 | -1.4258 | 0.99998 |  |  | -1.5095e-22 | 1.0771e-18 | -2.6162e-15 | 1.2789e-12 | 4.976e-09 | -8.537e-06 | -1.4259 | 0.99985 |  |  | 230 | 5/7-9/2018 |  | -1.4344 | 6.9508e-05 | -1.4346 | 5.6392e-05 | 177.0244 | 8.5544 | stopped using vac grease on o-rings. Dieletctic grease used on DF coverplate orings and Haynes Si grease used on all other orings
| 231 | 231 | 11017 | 10365 | 821 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  | stopped using vac grease on o-rings. Dieletctic grease used on DF coverplate orings and Haynes Si grease used on all other orings | 2/28-3/7/2018 | A | 1@SO | 0.12 | [6 empty] | -0.000997 | -1.4147 | 0.99996 |  |  | 7.449e-22 | -5.5283e-18 | 1.6613e-14 | -2.5885e-11 | 2.2535e-08 | -1.1017e-05 | -1.4147 | 0.99962 |  |  | 231 | 5/30-6/1/2018 |  | -1.4205 | 0.00019677 | -1.4219 | 0.00023339 | 1451.2007 | 5.8068 | dielectric grease used on ISFET orings
| 232 | 232 | 17267 | 10455 | 819 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 1/25-31/2018 | D | 1@SO | 0.105 | [6 empty] | -0.0010742 | -1.3906 | 0.99997 |  |  | -1.4623e-21 | 8.7066e-18 | -1.8029e-14 | 1.2895e-11 | 4.6302e-09 | -1.2296e-05 | -1.3907 | 0.99961 |  |  | 232 | 5/8-9/2018 |  | -1.3943 | 6.4774e-05 | -1.3947 | 7.2299e-05 | 443.1086 | 3.6396
| 233 | 233 |  | Yui |  |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 2/28-3/7/2018 | B | 1@SO | 0.145 | [6 empty] | -0.00097889 | -1.372 | 0.99996 |  |  | -1.9608e-24 | 1.4323e-18 | -8.8895e-15 | 1.9531e-11 | -1.5531e-08 | -1.7102e-06 | -1.372 | 0.99952 | [11 empty] | new chip installed in stem - gave to Yui
| 234 | 234 | 12688 | 10229 | 822 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 4/24-30/2018 | C | 1@SO | 0.155 | [6 empty] | -0.0010971 | -1.4015 | 0.99988 |  |  | 1.0905e-21 | -8.3977e-18 | 2.5908e-14 | -4.0979e-11 | 3.5736e-08 | -1.7756e-05 | -1.4015 | 0.9994 |  |  | 234 | 6/26-28/2018 |  | -1.4065 | 9.508e-05 | -1.4066 | 8.2693e-05 | 69.3675 | 5.0317
| 235 | 235 | 12711 | 10448 | 825 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 4/24-30/2018 | D | 1@SO | 0.085 | [6 empty] | -0.0010642 | -1.3824 | 0.99998 |  |  | -5.6381e-22 | 4.092e-18 | -1.1034e-14 | 1.2061e-11 | 6.9078e-11 | -1.0903e-05 | -1.3825 | 0.99977 |  |  | 235 | 7/3-6/2018 |  | -1.3896 | 7.7122e-05 | -1.3902 | 7.6939e-05 | 590.0339 | 7.1411
| 236 | 236 | 12714 | 10222 | 823 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 4/26-5/3/2018 | A | 1@SO | 0.08 | [6 empty] | -0.001025 | -1.3855 | 0.99998 |  |  | 1.2567e-21 | -9.9625e-18 | 3.1329e-14 | -4.9616e-11 | 4.1816e-08 | -1.817e-05 | -1.3855 | 0.99948 |  |  | 236 | 6/11-13/2018 |  | -1.3907 | 4.954e-05 | -1.3919 | 6.486e-05 | 1176.1541 | 5.2025
| 237 | 237 |  |  |  |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 4/26-5/3/2018 | B | 1@SO | 0.09 | [6 empty] | -0.0010297 | -1.408 | 0.99998 |  |  | -1.7261e-22 | 1.8642e-18 | -7.0142e-15 | 1.0891e-11 | -3.5871e-09 | -8.1702e-06 | -1.4081 | 0.99967 |  |  | 237 | 7/3-6/2018 |  | -1.454 | 9.0588e-05 | -1.4543 | 9.2866e-05 | 331.3303 | 45.9699 | K0 offset do not use for SOCCOM
| 238 | 238 | 12739 | 10446 | 826 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 6/5-11/2018 | C | 1@SO | 0.03 | [6 empty] | -0.0010588 | -1.3924 | 0.99997 |  |  | 3.7464e-22 | -2.2003e-18 | 4.6703e-15 | -5.0946e-12 | 6.6828e-09 | -1.0799e-05 | -1.3925 | 0.99979 |  |  | 238 | 7/23-25/2018 |  | -1.3845 | 0.00011472 | -1.3848 | 7.7038e-05 | 250.7611 | -7.8695
| 239 | 239 | 12749 | 10449 | 824 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 6/5-11/2018 | D | 1@SO | 0.095 | [6 empty] | -0.0010635 | -1.4325 | 0.99996 |  |  | -6.7873e-22 | 4.3528e-18 | -9.8109e-15 | 7.085e-12 | 5.7502e-09 | -1.1419e-05 | -1.4326 | 0.9997 |  |  | 239 | 7/23-25/2018 |  | -1.4222 | 6.6028e-05 | -1.4223 | 6.2342e-05 | 133.8191 | -10.3444
TPOS | 240 | 240 | 17350 | 10442 | 829 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 6/23-29/2018 | A | 1@SO | 0.15 | [6 empty] | -0.00097192 | -1.4255 | 0.99999 |  |  | -2.353e-23 | 9.5705e-20 | 5.7631e-16 | -4.4783e-12 | 1.1588e-08 | -1.4007e-05 | -1.4257 | 0.99985 |  |  | 240 | 8/10-13/2018 |  | -1.4155 | 0.00016121 | -1.4159 | 0.0001376 | 426.9196 | -10.0455
| 241 | 241 | 11090 | 10443 | 827 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/5-11/2018 | A | 1@SO | 0.02 | [6 empty] | -0.00096838 | -1.3951 | 0.99998 |  |  | 6.7163e-22 | -4.5427e-18 | 1.1567e-14 | -1.3964e-11 | 9.991e-09 | -8.1974e-06 | -1.3952 | 0.99985 |  |  | 241 | 8/10-13/2018 |  | -1.3846 | 0.00010559 | -1.385 | 8.8818e-05 | 429.9285 | -10.5736 | dielectric grease used on ISFET orings
| 242 | 242 | 12892 | 10572 | 828 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/5-11/2018 | B | 1@SO | 0.09 | [6 empty] | -0.0010283 | -1.4096 | 0.99998 |  |  | -5.7865e-23 | -2.0404e-19 | 2.9553e-15 | -9.0758e-12 | 1.2738e-08 | -8.8704e-06 | -1.4096 | 0.99939 |  |  | 242 | 8/10-13/2018 |  | -1.4 | 8.7539e-05 | -1.4001 | 8.6198e-05 | 103.0142 | -9.5971
TPOS | 243 | 243 | 17534 | 10366 | 832 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/5-11/2018 | C | 1@SO | 0.12 | [6 empty] | -0.0010336 | -1.4307 | 0.99994 |  |  | -1.479e-21 | 9.8406e-18 | -2.4214e-14 | 2.4848e-11 | -4.1107e-09 | -1.0056e-05 | -1.4307 | 0.9996 |  |  | 243 | 10/8-10/2018 |  | -1.4222 | 5.409e-05 | -1.4225 | 5.7201e-05 | 255.8971 | -8.4388
| 244 | 244 | 12885 | 10368 | 833 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/5-11/2018 | D | 1@SO | 0.145 | [6 empty] | -0.001042 | -1.4278 | 0.9999 |  |  | -6.5062e-22 | 3.8951e-18 | -7.9267e-15 | 4.156e-12 | 7.0714e-09 | -1.1118e-05 | -1.4279 | 0.99969 |  |  | 244 | 10/8-10/2018 |  | -1.4244 | 7.1805e-05 | -1.4248 | 7.5166e-05 | 376.8745 | -3.4024
TPOS | 245 | 245 | 18820 | 11428 | 834 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/23-30/2018 | A | 1@SO | 0.1 | [6 empty] | -0.0009605 | -1.4238 | 0.99997 |  |  | -2.2861e-21 | 1.4201e-17 | -3.1449e-14 | 2.6512e-11 | 1.7663e-09 | -1.5879e-05 | -1.4239 | 0.99963 |  |  | 245 | 1/21-23/2019 |  | -1.4147 | 6.7295e-05 | -1.4154 | 7.1396e-05 | 751.0909 | -9.0747
TPOS | 246 | 246 | 18709 | 11431 | 835 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/23-30/2018 | x | 1@SO | 0.06 | [6 empty] | -0.0010509 | -1.4041 | 0.99997 |  |  | -1.0989e-21 | 7.9711e-18 | -2.2197e-14 | 2.8182e-11 | -1.1217e-08 | -9.9879e-06 | -1.4042 | 0.99977 |  |  | 246 | 1/21-23/2019 |  | -1.3951 | 3.9944e-05 | -1.3954 | 4.1741e-05 | 283.8875 | -8.9862
TPOS | 247 | 247 | 18114 | 11432 | 837 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/24-30/2018 | C | 1@SO | 0.12 | [6 empty] | -0.0010856 | -1.3918 | 0.99978 |  |  | -4.5408e-22 | 1.4913e-18 | 2.3418e-15 | -1.6342e-11 | 2.7287e-08 | -1.9767e-05 | -1.3919 | 0.99956 |  |  | 247 | 2/11-13/2019 |  | -1.3835 | 7.1954e-05 | -1.3839 | 6.5687e-05 | 459.2691 | -8.2846
TPOS | 248 | 248 | 18601 | 11439 | 836 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/24-30/2018 | D | 1@SO | 0.12 | [6 empty] | -0.0010812 | -1.3956 | 0.9999 |  |  | 1.6742e-21 | -1.0902e-17 | 2.7079e-14 | -3.2827e-11 | 2.2908e-08 | -1.2639e-05 | -1.3957 | 0.99745 |  |  | 248 | 2/11-13/2019 |  | -1.3861 | 7.1694e-05 | -1.3864 | 5.313e-05 | 379.7798 | -9.5828
| 249 | 249 | 18796 | 11433 | 839 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  |  | 7/31-8/6/2018 | A | 1@SO | 0.11 | [6 empty] | -0.00092879 | -1.437 | 0.99995 |  |  | -9.2806e-22 | 6.4548e-18 | -1.6777e-14 | 1.8758e-11 | -4.6771e-09 | -7.7486e-06 | -1.4371 | 0.99984 |  |  | 249 | 3/13-17/2019 |  | -1.4307 | 0.00027415 | -1.4313 | 0.0002689 | 605.7648 | -6.3085
Australian Float | 250 | 250 | 12544 | 11434 | 838 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  | Cert 0.01N HCl used in PT chambers | 9/19-25/2018 | A | 1@SO | 0.02 | [6 empty] | -0.0010485 | -1.401 | 0.99998 |  |  | -3.3577e-22 | 3.0707e-18 | -1.0063e-14 | 1.3367e-11 | -1.6583e-09 | -1.2601e-05 | -1.4012 | 0.99994 |  |  | 250 | 3/13-17/2019 |  | -1.4047 | 9.6063e-05 | -1.4054 | 7.2437e-05 | 727.0839 | 3.6311 | PT chamber filled with certified 0.01N HCl soln
SBE DF stems
| C1 | T1 |  |  |  |  | sbe-3.5ksi | FuzzButton | [27 empty] | C1 | [8 empty] | K0 comparison with Navis head
| C2 | T2 |  |  |  |  | sbe-3.5ksi | FuzzButton | [27 empty] | C2 | [8 empty] | K0 comparison with Navis head
| C3 | T3 | 18864 | 11520 | 851 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | Sept 21 2018 |  | at SBE | [7 empty] | -0.0011066 | -1.3779 |  |  |  | -1.4412e-21 | 1.0634e-17 | -3.1846e-14 | 4.9955e-11 | -4.4566e-08 | 2.4666e-05 |  |  |  |  | C3 | 6/13-17/2019 |  | -1.3862 | 0.00015033 | -1.3867 | 0.0001374 | 588.7185 | 8.2607
| C4 | T4 |  |  |  |  | sbe-3.5ksi | FuzzButton |  | Cert 0.01N HCl | 9/19-25/2018 | B | 1@SO | 0.045 | [6 empty] | -0.0011045 | -1.4199 | 0.99998 |  |  | 1.0962e-22 | 8.7918e-19 | -8.7622e-15 | 2.4296e-11 | -3.0684e-08 | 2.1251e-05 | -1.4198 | 0.99963 |  |  | C4 | [8 empty] | PT chamber filled with certified 0.01N HCl soln
| C4 | T4 |  |  |  |  | sbe-3.5ksi | FuzzButton |  | Diluted 1N to 0.01N HCl | 9/6-12/2018 | B | 1@SO | 0.08 | [6 empty] | -0.0011174 | -1.4194 | 0.99985 |  |  | -3.3778e-22 | 3.8866e-18 | -1.6386e-14 | 3.3311e-11 | -3.5602e-08 | 2.2321e-05 | -1.4193 | 0.99951 |  |  | C4 | [8 empty] | PT chamber filled with 0.01N HCl diluted from 1N HCl stock
| C4 | T4 |  |  |  |  | sbe-3.5ksi | FuzzButton |  | on Gene's CPF | Aug 23 2018 |  | at SBE | [7 empty] | -0.0011416 | -1.4131 |  |  |  | -7.3622e-22 | 5.7894e-18 | -1.871e-14 | 3.2171e-11 | -3.1889e-08 | 2.038e-05 |  |  |  |  | C4 | [8 empty] | f(P) and K2 from SBE, K0 done at MBARI
| C7 | T7 | 18013 | 11440 | 846 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | Oct 24 2018 |  | at SBE | [7 empty] | -0.0010786 | -1.3832 |  |  |  | -2.189e-21 | 1.5404e-17 | -4.2856e-14 | 6.0012e-11 | -4.5146e-08 | 2.0909e-05 |  |  |  |  | C7 | 5/6-8/2019 |  | -1.3901 | 4.0803e-05 | -1.3908 | 4.0677e-05 | 679.1563 | 6.9173 | f(P) and K2 from SBE, K0 done at MBARI
| C8 | T8 | 12882 | 10367 | 830 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | Aug 31 2018 |  | at SBE | [7 empty] | -0.0010925 | -1.3703 |  |  |  | -9.5486e-22 | 7.404e-18 | -2.341e-14 | 3.903e-11 | -3.6937e-08 | 2.2023e-05 |  |  |  |  | C8 | 9/21-25/2018 |  | -1.3784 | 8.4757e-05 | -1.3788 | 6.6071e-05 | 373.9169 | 8.139 | f(P) and K2 from SBE, K0 done at MBARI
| F10253 | R10247 | 12888 | 10369 | 831 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | Aug 6 2018 |  | at SBE | [7 empty] | -0.0011281 | -1.3656 |  |  |  | -1.188e-21 | 8.0354e-18 | -2.0951e-14 | 2.7124e-11 | -1.9487e-08 | 1.0728e-05 |  |  |  |  | F10253 | 9/21-25/2018 |  | -1.3733 | 6.3292e-05 | -1.3741 | 5.443e-05 | 762.1704 | 7.741 | f(P) and K2 from SBE, K0 done at MBARI
| C9 | C9 | 18994 | 11430 | 841 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | 2019-02-11 |  | at SBE | [7 empty] | -0.0011011 | -1.3921 |  |  |  | -1.4414e-21 | 1.0489e-17 | -3.0795e-14 | 4.7007e-11 | -4.0562e-08 | 2.2298e-05 |  |  |  |  | C9 | 4/15-17/2019 |  | -1.4055 | 4.9508e-05 | -1.406 | 7.1464e-05 | 488.8575 | 13.412 | f(P) and K2 from SBE, K0 done at MBARI
| C10 | C10 |  |  |  |  | sbe-3.5ksi | FuzzButton |  | to Yui T for glider | 2018-11-30 |  | at SBE | [7 empty] | -0.0010747 | -1.3888 |  |  |  | -1.2138e-21 | 8.7851e-18 | -2.5645e-14 | 3.8874e-11 | -3.301e-08 | 1.8078e-05 |  |  |  |  | C10 | K0 will be run in Yuis lab or in-situ using glider data | [7 empty] | March 2019 spray glider deployement - Yui T
| C12 | C12 | 18739 | 11435 | 844 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | Oct 22 2018 |  | at SBE | [7 empty] | -0.0010809 | -1.4149 |  |  |  | 5.5783e-22 | -3.085e-18 | 4.7791e-15 | 2.197e-12 | -1.2978e-08 | 1.4452e-05 |  |  |  |  | C12 | 5/6-8/2019 |  | -1.4237 | 3.789e-05 | -1.424 | 4.3856e-05 | 302.518 | 8.8251 | f(P) and K2 from SBE, K0 done at MBARI
| C13 | C13 | 18098 | 11519 | 852 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | Oct 22 2018 |  | at SBE | [7 empty] | -0.0010948 | -1.3904 |  |  |  | 2.2435e-22 | -1.177e-18 | 1.4548e-15 | 2.2194e-12 | -7.3499e-09 | 9.2521e-06 |  |  |  |  | C13 | 6/13-17/2019 |  | -1.3965 | 0.00016404 | -1.3973 | 0.00013012 | 729.8727 | 6.1207 | f(P) and K2 from SBE, K0 done at MBARI
| C14 | C14 | 17545 | 13305 | 914 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0011473 | -1.4168 |  |  |  | -1.1525e-21 | 8.6968e-18 | -2.6926e-14 | 4.4212e-11 | -4.1281e-08 | 2.3047e-05 |  |  |  |  | C14 | 2/8-10/2021 |  | -1.4269 | 4.1991e-05 | -1.427 | 5.8407e-05 | 126.2584 | 10.1125 | f(P) and K2 from SBE, K0 done at MBARI
| C29 | C29 | 19875 | 12317 | 878 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010762 | -1.3725 |  |  |  | -1.1912e-21 | 8.8675e-18 | -2.694e-14 | 4.3184e-11 | -3.9692e-08 | 2.2444e-05 |  |  |  |  | C29 | 1/13-15/2020 |  | -1.3827 | 5.0872e-05 | -1.3831 | 5.7267e-05 | 377.8016 | 10.1728 | f(P) and K2 from SBE, K0 done at MBARI
| 10302 | 10302 | 18721 | 11429 | 857 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010708 | -1.4213 |  |  |  | -6.7749e-22 | 5.0682e-18 | -1.5517e-14 | 2.512e-11 | -2.3277e-08 | 1.3916e-05 |  |  |  |  | 10302 | 7/23-25/2019 |  | -1.4296 | 3.6766e-05 | -1.43 | 3.9613e-05 | 383.9913 | 8.2774 | f(P) and K2 from SBE, K0 done at MBARI
| 10303 | 10303 | 18082 | 11767 | 859 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.001089 | -1.3904 |  |  |  | -5.5304e-22 | 4.2487e-18 | -1.3481e-14 | 2.2822e-11 | -2.2034e-08 | 1.3485e-05 |  |  |  |  | 10303 | 7/29-31/2019 |  | -1.3988 | 4.4646e-05 | -1.3988 | 3.8555e-05 | 36.1622 | 8.4081 | f(P) and K2 from SBE, K0 done at MBARI
| 10311 | 10311 | 18852 | 11517 | 862 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010731 | -1.413 |  |  |  | 1.0533e-21 | -6.6445e-18 | 1.5091e-14 | -1.3327e-11 | 5.5161e-11 | 8.0198e-06 |  |  |  |  | 10311 | 8/12-14/2019 |  | -1.4213 | 6.4534e-05 | -1.4214 | 6.3963e-05 | 170.2382 | 8.254 | f(P) and K2 from SBE, K0 done at MBARI
| 10312 | 10312 | 18829 | 11527 | 870 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00107 | -1.4109 |  |  |  | -9.3645e-22 | 6.8781e-18 | -2.0517e-14 | 3.2076e-11 | -2.838e-08 | 1.6222e-05 |  |  |  |  | 10312 | 9/9-11/2019 |  | -1.4207 | 6.7432e-05 | -1.4212 | 5.2367e-05 | 564.4181 | 9.7751 | f(P) and K2 from SBE, K0 done at MBARI
| 10314 | 10314 | 18417 | 11526 | 867 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0011062 | -1.4001 |  |  |  | -1.5833e-21 | 1.1314e-17 | -3.2383e-14 | 4.7661e-11 | -3.8585e-08 | 1.8666e-05 |  |  |  |  | 10314 | 8/24-27/2019 |  | -1.4124 | 0.00019996 | -1.4124 | 0.00021722 | 22.7711 | 12.286 | f(P) and K2 from SBE, K0 done at MBARI
| 10315 | 10315 | 18169 | 11528 | 868 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010909 | -1.4048 |  |  |  | -9.7516e-22 | 7.0386e-18 | -2.0462e-14 | 3.0833e-11 | -2.5752e-08 | 1.3145e-05 |  |  |  |  | 10315 | 9/9-11/2019 |  | -1.4138 | 5.3464e-05 | -1.4144 | 7.9671e-05 | 505.3711 | 9.0465 | f(P) and K2 from SBE, K0 done at MBARI
| 10299 | 10299 | 19327 | 12315 | 888 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.001148 | -1.4055 |  |  |  | -1.3613e-21 | 9.7406e-18 | -2.7847e-14 | 4.0769e-11 | -3.2532e-08 | 1.5334e-05 |  |  |  |  | 10299 | 2/10-12/2020 |  | -1.4129 | 6.5786e-05 | -1.413 | 5.2783e-05 | 137.9869 | 7.3527
| 10301 | 10301 | 19118 | 12314 | 885 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0011149 | -1.3835 |  |  |  | -2.2103e-22 | 2.2715e-18 | -9.5873e-15 | 2.1099e-11 | -2.5167e-08 | 1.6865e-05 |  |  |  |  | 10301 | 2/10-12/2020 |  | -1.3926 | 5.3332e-05 | -1.3928 | 5.3863e-05 | 197.4432 | 9.0561
| 10390 | 10390 | 19050 | 12306 | 882 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010653 | -1.3989 |  |  |  | 1.15e-21 | -7.1967e-18 | 1.6049e-14 | -1.332e-11 | -1.6292e-09 | 1.0002e-05 |  |  |  |  | 10390 | 1/27-29/2020 |  | -1.4079 | 6.4304e-05 | -1.4083 | 4.6108e-05 | 422.5407 | 9.0148
| 10399 | 10399 | 19142 | 12488 | 900 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010827 | -1.4037 |  |  |  | 2.0249e-21 | -1.2971e-17 | 3.0452e-14 | -2.9717e-11 | 5.9926e-09 | 9.5298e-06 |  |  |  |  | 10399 | 7/15-17/2020 |  | -1.4116 | 8.3804e-05 | -1.4117 | 0.00011025 | 155.7271 | 7.8882
| 10402 | 10402 | 19017 | 12310 | 881 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010706 | -1.4027 |  |  |  | 4.4307e-22 | -2.3858e-18 | 3.3531e-15 | 2.9277e-12 | -1.1765e-08 | 1.16e-05 |  |  |  |  | 10402 | 1/27-29/2020 |  | -1.4106 | 6.0831e-05 | -1.4106 | 5.2819e-05 | 53.8736 | 7.8513
| 10403 | 10403 | 19085 | 12463 | 895 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.001049 | -1.399 |  |  |  | 5.0796e-22 | -2.8319e-18 | 4.552e-15 | 1.3397e-12 | -1.0716e-08 | 1.1963e-05 |  |  |  |  | 10403 | 6/8-10/2020 |  | -1.4077 | 3.3398e-05 | -1.4079 | 3.5465e-05 | 138.4397 | 8.7207
| 10410 | 10410 | 19090 | 12305 | 886 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010592 | -1.4048 |  |  |  | 1.09e-21 | -6.6847e-18 | 1.4218e-14 | -9.8105e-12 | -5.1769e-09 | 1.0983e-05 |  |  |  |  | 10410 | 2/10-12/2020 |  | -1.4139 | 4.9713e-05 | -1.4139 | 4.4816e-05 | 47.4879 | 9.0991
| 10414 | 10414 | 19061 | 12444 | 889 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010469 | -1.4279 |  |  |  | -9.7523e-23 | 1.2011e-18 | -5.6819e-15 | 1.3435e-11 | -1.6872e-08 | 1.2188e-05 |  |  |  |  | 10414 | 3/11-13/2020 |  | -1.4317 | 6.21e-05 | -1.4318 | 5.9627e-05 | 124.8995 | 3.788
| 10418 | 10418 | 19412 | 12307 | 887 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010752 | -1.4347 |  |  |  | 2.4914e-21 | -1.5929e-17 | 3.7258e-14 | -3.6005e-11 | 6.6904e-09 | 1.176e-05 |  |  |  |  | 10418 | 2/10-12/2020 |  | -1.4447 | 5.765e-05 | -1.4448 | 4.9288e-05 | 112.7952 | 9.9981
| 10419 | 10419 | 19588 | 12449 | 901 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010752 | -1.3883 |  |  |  | 1.5522e-21 | -9.8189e-18 | 2.2444e-14 | -2.0247e-11 | 8.5049e-10 | 1.0863e-05 |  |  |  |  | 10419 | 7/15-17/2020 |  | -1.397 | 4.7782e-05 | -1.3972 | 4.8044e-05 | 150.275 | 8.7365
| 10590 | 10590 | 19072 | 12451 | 903 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010587 | -1.3666 |  |  |  | -8.0486e-23 | 1.1864e-18 | -6.1138e-15 | 1.514e-11 | -1.9399e-08 | 1.4786e-05 |  |  |  |  | 10590 | 8/10-12/2020 |  | -1.3732 | 3.4899e-05 | -1.3733 | 3.5701e-05 | 96.1992 | 6.6404
| 10592 | 10592 | 19129 | 12453 | 906 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010602 | -1.3997 |  |  |  | -2.2507e-22 | 2.3301e-18 | -9.8794e-15 | 2.1794e-11 | -2.5882e-08 | 1.7488e-05 |  |  |  |  | 10592 | 8/17-19/2020 |  | -1.4079 | 5.0921e-05 | -1.4081 | 5.0483e-05 | 197.5331 | 8.2364
| 10593 | 10593 | 19512 | 12450 | 907 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010705 | -1.4011 |  |  |  | -5.5082e-22 | 4.3107e-18 | -1.4005e-14 | 2.4382e-11 | -2.3923e-08 | 1.4689e-05 |  |  |  |  | 10593 | 8/17-19/2020 |  | -1.409 | 5.6338e-05 | -1.4094 | 9.2326e-05 | 435.9785 | 7.89
| 10599 | 10599 | 19605 | 12458 | 911 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.001068 | -1.434 |  |  |  | -7.0092e-22 | 5.3782e-18 | -1.7038e-14 | 2.8794e-11 | -2.7629e-08 | 1.5875e-05 |  |  |  |  | 10599 | 8/31-9/2/2020 |  | -1.442 | 4.6905e-05 | -1.4422 | 4.4619e-05 | 195.7026 | 8.0055
SOCCOM2 start | 10605 | 10605 | 19045 | 13308 | 915 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010519 | -1.4207 |  |  |  | -9.6412e-22 | 7.0994e-18 | -2.1255e-14 | 3.3396e-11 | -2.9253e-08 | 1.5198e-05 |  |  |  |  | 10605 | 2/8-10/2021 |  | -1.4289 | 6.8357e-05 | -1.4295 | 0.00014012 | 587.9785 | 8.208
| 10619 | 10619 | 19302 | 13330 | 919 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010612 | -1.4019 |  |  |  | 7.3889e-22 | -3.8634e-18 | 4.7536e-15 | 7.2568e-12 | -2.2901e-08 | 2.1959e-05 |  |  |  |  | 10619 | 2/22-24/2021 |  | -1.412 | 5.3385e-05 | -1.4123 | 4.7824e-05 | 236.7024 | 10.1453
| 10646 | 10646 | 19996 | 13332 | 922 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0011211 | -1.4559 |  |  |  | 1.3157e-22 | 1.9393e-20 | -4.3227e-15 | 1.6002e-11 | -2.4188e-08 | 1.8174e-05 |  |  |  |  | 10646 | 3/8-10/2021 |  | -1.4633 | 4.2462e-05 | -1.4639 | 5.139e-05 | 628.0588 | 7.4045
| 10697 | 10697 | 19014 | 13331 | 921 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.001095 | -1.3921 |  |  |  | 6.6543e-22 | -3.6359e-18 | 5.4448e-15 | 3.1716e-12 | -1.584e-08 | 1.5086e-05 |  |  |  |  | 10697 | 3/8-10/2021 |  | -1.398 | 4.9444e-05 | -1.3982 | 4.9331e-05 | 155.9124 | 5.9039
| 10742 | 10742 | 20662 | 13624 | 933 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010816 | -1.4761 |  |  |  | -1.0757e-21 | 8.1657e-18 | -2.5511e-14 | 4.2392e-11 | -4.0282e-08 | 2.2811e-05 |  |  |  |  | 10742 | 6/2-3/2021 |  | -1.4846 | 4.5654e-05 | -1.4847 | 4.5716e-05 | 137.7702 | 8.482
| 10747 | 10747 | 20644 | 13578 | 927 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0011007 | -1.4382 |  |  |  | 1.4577e-21 | -8.8438e-18 | 1.8318e-14 | -1.1152e-11 | -9.6914e-09 | 1.627e-05 |  |  |  |  | 10747 | 4/12-14/2021 |  | -1.4446 | 4.4498e-05 | -1.445 | 6.9093e-05 | 331.4909 | 6.4487
| 10749 | 10749 | 20220 | 13623 | 928 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010377 | -1.407 |  |  |  | -3.2575e-22 | 2.8367e-18 | -1.0386e-14 | 2.0412e-11 | -2.2641e-08 | 1.4562e-05 |  |  |  |  | 10749 | 4/12-14/2021 |  | -1.4136 | 4.096e-05 | -1.414 | 4.3078e-05 | 363.4758 | 6.6251
| 10850 | 10850 | 20043 | 14221 | 935 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010581 | -1.5172 |  |  |  | -7.0323e-22 | 5.0482e-18 | -1.4552e-14 | 2.1648e-11 | -1.7971e-08 | 9.5727e-06 |  |  |  |  | 10850 | 5/17-19/2021 |  | -1.5218 | 4.3506e-05 | -1.5221 | 4.9838e-05 | 257.6667 | 4.6477
| 10852 | 10852 | 19751 | 14218 | 936 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010661 | -1.5037 |  |  |  | 4.9467e-23 | 1.6974e-19 | -2.8115e-15 | 9.39e-12 | -1.3941e-08 | 1.0946e-05 |  |  |  |  | 10852 | 5/17-19/2021 |  | -1.5103 | 4.8866e-05 | -1.5106 | 5.269e-05 | 259.1828 | 6.6304
| 10868 | 10868 | 20103 | 14099 | 938 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.001045 | -1.3927 |  |  |  | -2.6583e-22 | 2.3015e-18 | -8.3763e-15 | 1.6372e-11 | -1.7946e-08 | 1.1722e-05 |  |  |  |  | 10868 | 6/7-9/2021 |  | -1.3998 | 4.2442e-05 | -1.4 | 4.1512e-05 | 176.0779 | 7.1211
| 10894 | 10894 | 20002 | 14098 | 940 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.0010562 | -1.4966 |  |  |  | 4.6708e-23 | 2.444e-19 | -3.2678e-15 | 1.0607e-11 | -1.5215e-08 | 1.1052e-05 |  |  |  |  | 10894 | 6/7-9/2021 |  | -1.5015 | 4.1311e-05 | -1.5017 | 3.6626e-05 | 206.7118 | 4.8868
| 10896 | 10896 | 20060 | 14301 | 942 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010589 | -1.4748 |  |  |  | -7.3232e-22 | 5.4885e-18 | -1.6849e-14 | 2.7379e-11 | -2.5161e-08 | 1.3348e-05 |  |  |  |  | 10896 | 6/28-29/2021 |  | -1.524 | 4.0517e-05 | -1.5243 | 3.9185e-05 | 260.925 | 49.2192
| 10952 | 10952 | 20532 | 14293 | 943 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010356 | -1.3797 |  |  |  | 1.0753e-21 | -6.6225e-18 | 1.4235e-14 | -1.0284e-11 | -4.3975e-09 | 1.014e-05 |  |  |  |  | 10952 | 6/28-29/2021 |  | -1.3867 | 4.2473e-05 | -1.3868 | 3.9204e-05 | 109.1304 | 6.9789
| 10953 | 10953 | 20520 | 14433 | 946 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010393 | -1.3763 |  |  |  | -7.6874e-22 | 5.5133e-18 | -1.5872e-14 | 2.3572e-11 | -1.9495e-08 | 1.0458e-05 |  |  |  |  | 10953 | 7/15-19/2021 |  | -1.3818 | 6.7399e-05 | -1.3821 | 6.3153e-05 | 302.3802 | 5.4979
MATLAB | 10954 | 10954 | 20144 | 14458 | 949 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010283 | -1.3612 |  |  |  | 7.5279e-22 | -4.2278e-18 | 6.9788e-15 | 1.3037e-12 | -1.4526e-08 | 1.4566e-05 |  |  |  |  | 10954 | 07/27-07/29/2021 |  | -1.3685 | 5.8467e-05 | -1.3687 | 4.6368e-05 | 260 | 7.27
| 10958 | 10958 | 20135 | 14461 | 950 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010423 | -1.4067 |  |  |  | -3.8419e-22 | 3.0324e-18 | -9.9589e-15 | 1.7552e-11 | -1.7952e-08 | 1.1539e-05 |  |  |  |  | 10958 | 07/26-07/29/2021 |  | -1.4135 | 6.6183e-05 | -1.4137 | 5.6844e-05 | 170 | 6.8
| 10960 | 10960 | 20150 | 14457 | 952 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010174 | -1.3941 |  |  |  | -2.5499e-22 | 2.1082e-18 | -7.3078e-15 | 1.3635e-11 | -1.4736e-08 | 1.0006e-05 |  |  |  |  | 10960 | 08/02-08/05/2021 |  | -1.4005 | 5.8009e-05 | -1.4005 | 5.9495e-05 | 60 | 6.39
| 10961 | 10961 | 20592 | 14460 | 953 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010155 | -1.3561 |  |  |  | -8.2863e-22 | 6.012e-18 | -1.7619e-14 | 2.6875e-11 | -2.27e-08 | 1.0632e-05 |  |  |  |  | 10961 | 08/02-08/05/2021 |  | -1.3618 | 0.00013767 | -1.3618 | 0.00011684 | 70 | 5.66
| 11018 | 11018 | 20084 | 14459 | 960 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00099759 | -1.3025 |  |  |  | -2.0068e-22 | 1.6506e-18 | -5.6898e-15 | 1.0554e-11 | -1.1084e-08 | 6.9364e-06 |  |  |  |  | 11018 | 08/23-08/25/2021 |  | -1.3073 | 0.00015942 | -1.3074 | 0.00014711 | 60 | 4.8
| 11021 | 11021 | 20134 | 15328 | 961 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00098133 | -1.3053 |  |  |  | -3.9967e-22 | 2.8704e-18 | -8.2999e-15 | 1.2432e-11 | -1.0268e-08 | 5.2143e-06 |  |  |  |  | 11021 | 08/30-09/01/2021 |  | -1.3103 | 5.3989e-05 | -1.3103 | 5.3227e-05 | 20 | 5.02
| 11022 | 11022 | 20138 | 15164 | 963 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00098231 | -1.3165 |  |  |  | 1.6884e-22 | -7.7905e-19 | 3.5314e-16 | 3.6785e-12 | -7.9813e-09 | 7.1434e-06 |  |  |  |  | 11022 | 08/30-09/01/2021 |  | -1.3217 | 3.6602e-05 | -1.3218 | 3.8555e-05 | 100 | 5.22
| 11045 | 11045 | 20492 | 15165 | 962 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00097711 | -1.3531 |  |  |  | 1.6045e-22 | -5.9539e-19 | -7.5152e-16 | 6.6635e-12 | -1.1931e-08 | 9.3985e-06 |  |  |  |  | 11045 | 08/30-09/01/2021 |  | -1.3588 | 6.4816e-05 | -1.3589 | 5.9507e-05 | 60 | 5.73
| 11046 | 11046 | 20620 | 15480 | 971 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00098775 | -1.3402 |  |  |  | -5.6689e-23 | 1.0237e-18 | -5.6845e-15 | 1.4597e-11 | -1.8924e-08 | 1.2886e-05 |  |  |  |  | 11046 | 10/04-10/06/2021 |  | -1.3458 | 6.1013e-05 | -1.3459 | 5.717e-05 | 130 | 5.58
| 11055 | 11055 | 20136 | 15850 | 974 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0010009 | -1.3199 |  |  |  | -6.5375e-22 | 4.6949e-18 | -1.3575e-14 | 2.0336e-11 | -1.6702e-08 | 7.8411e-06 |  |  |  |  | 11055 | 09/27-09/29/2021 |  | -1.3232 | 4.7892e-05 | -1.3234 | 4.308e-05 | 210 | 3.25
| 11057 | 11057 | 20169 | 15863 | 975 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00097955 | -1.3376 |  |  |  | 1.0142e-21 | -6.2767e-18 | 1.3738e-14 | -1.0813e-11 | -2.2907e-09 | 8.1984e-06 |  |  |  |  | 11057 | 10/04-10/06/2021 |  | -1.342 | 6.4127e-05 | -1.3421 | 5.6824e-05 | 70 | 4.39
| 11058 | 11058 | 21291 | 16646 | 976 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00098012 | -1.331 |  |  |  | -3.5443e-22 | 2.7557e-18 | -8.8896e-15 | 1.5358e-11 | -1.5109e-08 | 8.8993e-06 |  |  |  |  | 11058 | 01/31-02/02/2022 |  | -1.3355 | 7.8313e-05 | -1.3359 | 8.0942e-05 | 420 | 4.46
| 11085 | 11085 | 20962 | 15864 | 977 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.0009235 | -1.2563 |  |  |  | -6.343e-23 | 4.772e-19 | -1.4712e-15 | 2.4021e-12 | -2.3613e-09 | 2.6127e-06 |  |  |  |  | 11085 | 11/01-11/03/2021 |  | -1.2604 | 9.087e-05 | -1.2604 | 8.6906e-05 | 20 | 4.1
| 11111 | 11111 | 20380 | 15865 | 978 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00092377 | -1.2646 |  |  |  | -1.7909e-22 | 1.5319e-18 | -5.5065e-15 | 1.0641e-11 | -1.1652e-08 | 7.1966e-06 |  |  |  |  | 11111 | 11/01-11/03/2021 |  | -1.2698 | 4.774e-05 | -1.2698 | 4.0624e-05 | 50 | 5.2
| 11096 | 11096 | 20189 | 16132 | 990 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00093087 | -1.2526 |  |  |  | -7.2318e-22 | 5.2318e-18 | -1.5243e-14 | 2.3013e-11 | -1.9186e-08 | 9.208e-06 |  |  |  |  | 11096 | 12/13-12/15/2021 |  | -1.2583 | 6.851e-05 | -1.2584 | 6.5534e-05 | 110 | 5.68
| 11150 | 11150 | 18340 | 14103 | 991 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00090016 | -1.3289 |  |  |  | -8.9958e-23 | 8.6891e-19 | -3.4887e-15 | 7.3879e-12 | -8.7048e-09 | 6.3928e-06 |  |  |  |  | 11150 | 12/13-12/15/2021 |  | -1.3347 | 0.00011379 | -1.3348 | 8.2369e-05 | 10 | 5.84
| 11232 | 11232 | 20623 | 13782 | 995 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00085054 | -1.3646 |  |  |  | -4.2173e-22 | 3.0352e-18 | -8.774e-15 | 1.3094e-11 | -1.0905e-08 | 6.5828e-06 |  |  |  |  | 11232 | 12/20-12/23/2021 |  | -1.3713 | 5.1684e-05 | -1.3715 | 4.8733e-05 | 170 | 6.72
| 11202 | 11202 | 20057 | 14100 | 996 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00086626 | -1.3253 |  |  |  | -5.0659e-22 | 3.7704e-18 | -1.1452e-14 | 1.8331e-11 | -1.6359e-08 | 8.6167e-06 |  |  |  |  | 11202 | 12/20-12/22/2021 |  | -1.3315 | 6.582e-05 | -1.3316 | 6.6461e-05 | 80 | 6.23
| 11254 | 11254 | 21883 | 17089 | 1004 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00086835 | -1.2793 |  |  |  | -2.2828e-22 | 1.7545e-18 | -5.5784e-15 | 9.4767e-12 | -9.1672e-09 | 6.1672e-06 |  |  |  |  | 11254 | 02/28-03/02/2022 |  | -1.2852 | 6.0349e-05 | -1.2854 | 7.0481e-05 | 250 | 5.87
| 11270 | 11270 | 22873 | 18990 | 1002 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00088895 | -1.2549 |  |  |  | 4.6844e-22 | -2.8632e-18 | 6.0182e-15 | -3.8964e-12 | -2.7396e-09 | 6.1547e-06 |  |  | SBE |  | 11270 | 10/02-10/03/2023 |  | -1.2646 | 5.5734e-05 | -1.2648 | 5.9188e-05 | 250 | 9.68
OCR | 11307 | 11307 | 23599 | 17086 | 1006 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00091104 | -1.2758 |  |  |  | 4.2981e-24 | 1.8085e-19 | -1.4516e-15 | 4.2334e-12 | -6.013e-09 | 5.2903e-06 |  |  |  |  | 11307 | 03/07-03/10/2022 |  | -1.2856 | 6.3684e-05 | -1.2857 | 6.8129e-05 | 60 | 9.85
OCR | 11311 | 11311 | 22826 | 18596 | 1007 |  | sbe-3.5ksi | FuzzButton |  | GO-BGC |  |  | at SBE | [7 empty] | -0.00090368 | -1.283 |  |  |  | -2.3395e-22 | 1.823e-18 | -5.8868e-15 | 1.0172e-11 | -1.0062e-08 | 6.4328e-06 |  |  | SBE |  | 11311 | 10/02-10/03/2023 |  | -1.2889 | 4.6752e-05 | -1.2892 | 5.1616e-05 | 320 | 5.88
| 11296 | 11296 | 21838 | 16637 | 1010 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00091447 | -1.2721 |  |  |  | 1.404e-22 | -6.8928e-19 | 5.6602e-16 | 2.3658e-12 | -5.8109e-09 | 6.0833e-06 |  |  |  |  | 11296 | 03/14-03/16/2022 |  | -1.2787 | 7.6598e-05 | -1.279 | 5.4258e-05 | 280 | 6.62
| 11310 | 11310 | 21250 | 16666 | 1011 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00092409 | -1.2863 |  |  |  | -8.9553e-22 | 6.2236e-18 | -1.7078e-14 | 2.3537e-11 | -1.718e-08 | 8.0143e-06 |  |  |  |  | 11310 | 03/14-03/16/2022 |  | -1.2914 | 5.2655e-05 | -1.2916 | 5.5547e-05 | 150 | 5.12
| 11360 | 11360 | 21871 | 16792 | 1009 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM |  |  | at SBE | [7 empty] | -0.00088757 | -1.3008 |  |  |  | -6.5817e-22 | 4.6705e-18 | -1.3206e-14 | 1.9041e-11 | -1.4754e-08 | 7.4435e-06 |  |  |  |  | 11360 | 03/14-03/16/2022 |  | -1.3108 | 5.4539e-05 | -1.311 | 8.4145e-05 | 190 | 9.97
| 11372 | 11372 | [7 empty] | April 2022 returned to SBE | 03/04-03/07/2022 | A | at MBARI | 0.117 | [6 empty] | -0.00095186 | -1.2856 | 0.9999 | 5.23e-27 | -4.7236e-23 | 1.7667e-19 | -3.5287e-16 | 4.0381e-13 | -2.6201e-10 | 8.7288e-08 | -8.8245e-06 | -1.2857 | 0.9935 | SBE |  | 11372
| 11472 | 11472 | [7 empty] | April 2022 returned to SBE | 03/04-03/07/2022 | B | at MBARI | 0.037 | [6 empty] | -0.00098969 | -1.3071 | 1 |  | -3.2323e-24 | 2.5488e-20 | -8.0084e-17 | 1.2689e-13 | -1.0453e-10 | 3.9538e-08 | -1.3984e-06 | -1.3071 | 0.9967 | SBE |  | 11472
| 11455 | 11455 | [7 empty] | April 2022 returned to SBE | 03/09-03/12/2022 | A | at MBARI | -0.041 | [6 empty] | -0.00091072 | -1.3414 | 0.9997 |  |  | 1.239e-21 | -8.0394e-18 | 1.9669e-14 | -2.1999e-11 | 1.0075e-08 | 1.6333e-06 | -1.3414 | 0.9956 | SBE |  | 11455
| 11987 | 11987 | 21499 | 17687 | 1040 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | 08/26-08/30/2022 |  |  | -0.033 | [6 empty] | -0.00084159 | -1.2859 | 0.9995 |  |  | 1.1592e-21 | -7.2511e-18 | 1.638e-14 | -1.4221e-11 | -7.3795e-10 | 1.0055e-05 | -1.2861 | 0.9955 | SBE |  | 11987 | 09/19-09/22/2022 |  | -1.2888 | 0.00020926 | -1.2889 | 0.00020588 | 90 | 2.85
| 11988 | 11988 | 21105 | 17684 | 1041 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | 08/26-08/30/2022 |  |  | -0.074 | [6 empty] | -0.00084231 | -1.2743 | 0.9998 |  |  | -3.7556e-22 | 3.598e-18 | -1.3974e-14 | 2.8377e-11 | -3.2046e-08 | 2.1254e-05 | -1.2746 | 0.996 | SBE |  | 11988 | 09/16-09/21/2022 |  | -1.2785 | 0.00035045 | -1.2785 | 0.00035069 | 70 | 4.17
| 11991 | 11991 | 21675 | 17685 | 1042 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | 08/27-09/01/2022 |  |  | -0.137 | [6 empty] | -0.0009686 | -1.2809 | 0.9999 |  |  | 1.1532e-21 | -7.096e-18 | 1.5469e-14 | -1.2183e-11 | -2.8318e-09 | 1.0735e-05 | -1.2811 | 0.9976 | SBE |  | 11991 | 09/26-09/28/2022 |  | -1.2871 | 7.0309e-05 | -1.2872 | 6.7929e-05 | 110 | 6.18
| 11992 | 11992 | 21902 | 17686 | 1043 |  | sbe-3.5ksi | FuzzButton |  | SOCCOM | 08/27-09/01/2022 |  |  | -0.168 | [5 empty] | -2.2397e-08 | -0.00088774 | -1.3074 | 0.9987 |  |  | -4.2068e-22 | 3.8784e-18 | -1.4195e-14 | 2.6618e-11 | -2.7119e-08 | 1.6153e-05 | -1.3077 | 0.9901 | SBE |  | 11992 | 09/26-09/28/2022 |  | -1.3116 | 0.00012199 | -1.3117 | 0.00011017 | 90 | 4.2
Raw cal data from SBE used for calcs | 12026 | 12026 | 17777 | 17421 | 1053 |  | sbe-3.5ksi | FuzzButton |  |  | 09/26-10/01/2022 |  |  | -0.024 |  |  |  | -7.7877e-15 | 3.6032e-11 | -5.9074e-08 | -0.00087423 | -1.2994 | 0.9999 | 4.5869e-27 | -4.109e-23 | 1.524e-19 | -3.0084e-16 | 3.3576e-13 | -2.0229e-10 | 4.8234e-08 | 1.0656e-05 | -1.2998 | 0.9993 | SBE |  | 12026 | 10/28-11/01/2022 |  | -1.3017 | 0.00013253 | -1.3018 | 0.00013843 | 180 | 2.21
| 12027 | 12027 | 21467 | 17434 | 1052 |  | sbe-3.5ksi | FuzzButton |  |  | 09/26-10/01/2022 |  |  | -0.033 |  |  |  | -1.5726e-14 | 5.9e-11 | -7.75e-08 | -0.00094825 | -1.2678 | 0.9999 |  |  |  | 1.0668e-18 | -7.4683e-15 | 2.09e-11 | -2.8872e-08 | 2.1263e-05 | -1.2684 | 0.9983 | SBE |  | 12027 | 10/28-11/01/2022 |  | -1.2719 | 0.00015529 | -1.2724 | 0.00019238 | 410 | 4.18
| 12028 | 12028 | 21865 | 17349 | 1069 |  | sbe-3.5ksi | FuzzButton |  |  | 09/25-09/30/2022 |  |  | -0.049 |  |  |  | -1.2912e-14 | 4.7763e-11 | -6.0413e-08 | -0.0010099 | -1.259 | 0.9998 |  |  | 5.9936e-22 | -2.729e-18 | 1.6187e-15 | 1.0649e-11 | -2.3562e-08 | 2.0625e-05 | -1.2596 | 0.9982 | SBE |  | 12028 | 12/19-12/22/2022 |  | -1.2636 | 5.8425e-05 | -1.2637 | 5.7298e-05 | 150 | 4.59
| 12114 | 12114 | 20563 | 16410 | 1054 |  | sbe-3.5ksi | FuzzButton |  |  | 09/25-09/30/2022 |  |  | -0.104 |  |  |  | -8.2538e-15 | 2.9484e-11 | -3.1505e-08 | -0.00094851 | -1.281 | 0.9998 | 3.446e-27 | -3.0608e-23 | 1.1196e-19 | -2.1642e-16 | 2.3437e-13 | -1.3561e-10 | 3.1004e-08 | 6.3933e-06 | -1.2813 | 0.997 | SBE |  | 12114 | 11/07-11/09/2022 |  | -1.2858 | 6.6167e-05 | -1.286 | 4.8947e-05 | 210 | 4.8
| 12115 | 12115 | 21075 | 17590 | 1057 |  | sbe-3.5ksi | FuzzButton |  |  | 10/06-10/11/2022 |  |  | -0.055 | [5 empty] | -1.6988e-08 | -0.00086816 | -1.284 | 0.9994 |  | -1.9253e-24 | 1.4516e-20 | -4.2783e-17 | 6.0369e-14 | -3.6552e-11 | -2.0744e-09 | 1.5266e-05 | -1.2845 | 0.9959 | SBE |  | 12115 | 11/14-11/16/2022 |  | -1.2861 | 8.3074e-05 | -1.2863 | 9.873e-05 | 190 | 2.12
| 12113 | 12113 | 21810 | 17435 | 1058 |  | sbe-3.5ksi | FuzzButton |  |  | 08/31-09/03/2022 |  |  | 0.228 |  |  |  |  | 5.6545e-12 | -2.551e-08 | -0.00090414 | -1.2904 | 0.9996 |  | -4.5851e-24 | 3.6919e-20 | -1.1995e-16 | 2.0005e-13 | -1.779e-10 | 7.5286e-08 | -3.5581e-06 | -1.2904 | 0.9983 | SBE |  | 12113 | 11/14-11/16/2022 |  | -1.2918 | 6.2183e-05 | -1.2919 | 5.7879e-05 | 130 | 1.4
| 12109 | 12109 | 21861 | 17591 | 1059 |  | sbe-3.5ksi | FuzzButton |  |  | 09/02-09/05/2022 | E |  | 0.025 |  |  | -4.2338e-18 | 3.2705e-14 | -6.6647e-11 | 2.4927e-08 | -0.00085187 | -1.2896 | 0.9999 | 3.647e-27 | -2.9943e-23 | 9.9269e-20 | -1.6849e-16 | 1.4937e-13 | -5.3641e-11 | -1.3789e-08 | 2.0517e-05 | -1.2905 | 0.9976 | SBE |  | 12109 | 11/14-11/16/2022 |  | -1.2936 | 6.5894e-05 | -1.2937 | 7.6475e-05 | 140 | 3.97
| 12128 | 12128 | 21818 | 17589 | 1060 |  | sbe-3.5ksi | FuzzButton |  |  | 09/09-09/12/2022 |  |  | -0.14 |  |  |  |  | 1.5345e-11 | -5.0103e-08 | -0.00081315 | -1.3139 | 0.9997 | 6.3553e-27 | -5.3385e-23 | 1.846e-19 | -3.381e-16 | 3.4854e-13 | -1.9158e-10 | 3.7537e-08 | 1.3278e-05 | -1.3145 | 0.9988 | SBE |  | 12128 | 11/14-11/16/2022 |  | -1.3173 | 7.2351e-05 | -1.3174 | 7.7769e-05 | 110 | 3.39
| 12108 | 12108 | 22185 | 18479 | 1072 |  | sbe-3.5ksi | FuzzButton |  |  | 09/02-09/05/2022 |  |  | -0.077 |  |  | 4.1245e-17 | -1.2002e-13 | 6.6865e-11 | 6.0699e-09 | -0.00083141 | -1.2875 | 0.9998 | 4.8761e-27 | -4.081e-23 | 1.3957e-19 | -2.5031e-16 | 2.4968e-13 | -1.3199e-10 | 2.6462e-08 | 7.738e-06 | -1.2879 | 0.9987 | SBE |  | 12108 | 01/23-01/25/2023 |  | -1.2915 | 3.9857e-05 | -1.2915 | 3.7828e-05 | 30 | 3.99
| 12112 | 12112 | 22213 | 17596 | 1064 |  | sbe-3.5ksi | FuzzButton |  |  | 09/02-09/05/2022 | D |  | 0.058 |  |  |  | 2.3964e-14 | -6.3303e-11 | 1.98e-08 | -0.00089107 | -1.2982 | 0.9999 |  |  | -7.1753e-23 | 9.3733e-19 | -4.6521e-15 | 1.1682e-11 | -1.5892e-08 | 1.369e-05 | -1.2987 | 0.9989 | SBE |  | 12112 | 12/05-12/08/2022 |  | -1.3018 | 6.1626e-05 | -1.3021 | 6.5936e-05 | 240 | 3.64
| 12116 | 12116 | 21576 | 17383 | 1065 |  | sbe-3.5ksi | FuzzButton |  |  | 08/31-09/03/2022 | C |  | -0.087 | [5 empty] | -2.5085e-08 | -0.00081168 | -1.3097 | 0.9995 |  |  |  | 1.5098e-18 | -9.3382e-15 | 2.2892e-11 | -2.806e-08 | 1.956e-05 | -1.3104 | 0.9965 | SBE |  | 12116 | 12/12-12/14/2022 |  | -1.3123 | 6.9186e-05 | -1.3123 | 6.035e-05 | 40 | 2.58
| 12117 | 12117 | 21905 | 17386 | 1063 |  | sbe-3.5ksi | FuzzButton |  |  | 08/31-09/03/2022 | A |  | -0.04 |  |  | 4.1498e-17 | -1.5556e-13 | 1.9015e-10 | -1.1111e-07 | -0.00083417 | -1.2883 | 0.9996 | 3.4394e-27 | -3.0258e-23 | 1.0956e-19 | -2.0878e-16 | 2.1931e-13 | -1.1511e-10 | 1.2356e-08 | 1.6822e-05 | -1.289 | 0.9986 | SBE |  | 12117 | 12/05-12/07/2022 |  | -1.2912 | 6.3475e-05 | -1.2917 | 6.7522e-05 | 480 | 2.87
| 12118 | 12118 | 22080 | 18548 | 1071 |  | sbe-3.5ksi | FuzzButton |  |  | 08/31-09/03/2022 | B |  | -0.054 |  |  |  | 2.5467e-14 | -5.6132e-11 | -4.8436e-09 | -0.00084517 | -1.2622 | 0.9997 |  |  | 2.5828e-22 | -3.0349e-19 | -5.1423e-15 | 1.9749e-11 | -2.961e-08 | 2.2456e-05 | -1.2631 | 0.9967 | SBE |  | 12118 | 01/23-01/25/2023 |  | -1.2654 | 3.8722e-05 | -1.2656 | 3.5347e-05 | 280 | 3.2
| 12127 | 12127 | 22693 | 18481 | 1073 |  | sbe-3.5ksi | FuzzButton |  |  | 09/09-09/12/2022 | B |  | -0.039 |  |  |  | 3.4755e-14 | -9.1067e-11 | 3.2476e-08 | -0.00086445 | -1.2847 | 0.9998 |  |  | -4.0389e-23 | 2.2801e-19 | -9.1854e-16 | 3.9783e-12 | -9.9663e-09 | 1.2882e-05 | -1.285 | 0.9992 | SBE |  | 12127 | 01/23-01/25/2023 |  | -1.2878 | 4.4804e-05 | -1.2881 | 3.9501e-05 | 250 | 3.2
| 12024 | 12024 | 22496 | 18546 | 1074 |  | sbe-3.5ksi | FuzzButton |  |  | 09/06-09/10/2022 | D |  | -0.055 |  |  |  | -8.207e-15 | 4.135e-11 | -6.9006e-08 | -0.00088644 | -1.253 | 0.9992 | 4.6781e-27 | -4.2537e-23 | 1.5996e-19 | -3.1952e-16 | 3.5985e-13 | -2.1841e-10 | 5.344e-08 | 8.9489e-06 | -1.2535 | 0.9966 | SBE |  | 12024 | 01/23-01/25/2023 |  | -1.2549 | 3.6067e-05 | -1.255 | 3.8394e-05 | 100 | 1.93
| 12183 | 12183 |  | 18925 | 1119 |  | sbe-3.5ksi | FuzzButton |  |  | 09/30-10/03/2022 | B |  | 0.073 | [5 empty] | -1.3352e-08 | -0.00085412 | -1.3397 | 0.9999 |  |  | 1.2607e-21 | -7.1626e-18 | 1.3552e-14 | -6.5421e-12 | -8.3801e-09 | 1.2042e-05 | -1.3401 | 0.9953 | SBE |  | 12183 | 11/07-11/09/2023 |  | -1.3419 | 8.3298e-05 | -1.342 | 8.3238e-05 | 100 | 2.22
| 12186 | 12186 | 22678 | 18528 | 1081 |  | sbe-3.5ksi | FuzzButton |  |  | 09/30-10/03/2022 | E |  | 0.053 |  |  |  | 2.2859e-14 | -5.4196e-11 | -3.2932e-10 | -0.00088533 | -1.3115 | 0.9999 |  |  | -1.1331e-22 | 1.3993e-18 | -6.2324e-15 | 1.3936e-11 | -1.7651e-08 | 1.4949e-05 | -1.312 | 0.9983 | SBE |  | 12186 | 03/06-03/08/2023 |  | -1.3185 | 5.0049e-05 | -1.3186 | 5.9254e-05 | 110 | 6.97
| 12241 | 12241 | 20116 | 15862 | 1085 |  | sbe-3.5ksi | FuzzButton |  |  | 11/04-11/07/2022 | B |  | 0.12 |  |  |  |  | 8.2776e-12 | -3.1899e-08 | -0.00086879 | -1.3091 | 0.9997 |  |  | 1.23e-21 | -7.1195e-18 | 1.4171e-14 | -8.4319e-12 | -7.1508e-09 | 1.388e-05 | -1.3095 | 0.9977 | SBE |  | 12241 | 04/04-04/06/2023 |  | -1.3114 | 5.3572e-05 | -1.3115 | 5.8163e-05 | 70 | 2.32
| 12219 | 12219 | 22257 | 18547 | 1083 |  | sbe-3.5ksi | FuzzButton |  |  | 12/08-12/11/2022 |  |  | 0.12 |  |  |  | 8.6129e-15 | -2.2948e-11 | -2.8164e-09 | -0.00083065 | -1.3462 | 0.9997 |  |  | -6.1777e-22 | 5.3436e-18 | -1.8528e-14 | 3.3358e-11 | -3.3509e-08 | 2.0584e-05 | -1.3469 | 0.9971 | SBE |  | 12219 | 04/04-04/06/2023 |  | -1.3471 | 7.5937e-05 | -1.3472 | 7.7099e-05 | 70 | 0.87
| 12302 | 12302 | 22508 | 18682 | 1084 |  | sbe-3.5ksi | FuzzButton |  |  | 01/19-01/22/2023 |  |  | 0.022 |  |  |  |  | -2.1755e-12 | -6.4854e-10 | -0.00096036 | -1.2958 | 0.9998 |  |  | 2.279e-21 | -1.5215e-17 | 3.9134e-14 | -4.7425e-11 | 2.4826e-08 | 7.0932e-07 | -1.2959 | 0.9983 | SBE |  | 12302 | 04/04-04/06/2023 |  | -1.2975 | 8.4266e-05 | -1.2976 | 8.5121e-05 | 80 | 1.66
| 11914 | 11914 | 22934 | 18989 | 1108 |  | sbe-3.5ksi | FuzzButton |  |  | 07/21-07/24/2023 | B |  | -0.124 |  | 2.9778e-20 | -1.6019e-16 | 3.0134e-13 | -2.3463e-10 | 6.3268e-08 | -0.00095754 | -1.3243 | 0.9998 |  |  | -8.5065e-22 | 6.3293e-18 | -1.9267e-14 | 3.1551e-11 | -2.9864e-08 | 1.6745e-05 | -1.3249 | 0.9979 | SBE |  | 11914 | 09/18-09/20/2023 |  | -1.3289 | 6.7879e-05 | -1.3294 | 9.1584e-05 | 440 | 4.6
| 11915 | 11915 | 22365 | 18687 | 1109 |  | sbe-3.5ksi | FuzzButton |  |  | 07/21-07/24/2023 | F |  | -0.106 |  |  |  |  | 1.917e-11 | -5.0167e-08 | -0.00098934 | -1.2855 | 0.9999 |  |  | -5.8472e-22 | 5.1805e-18 | -1.8644e-14 | 3.5111e-11 | -3.7031e-08 | 2.296e-05 | -1.2863 | 0.9983 | SBE |  | 11915 | 09/18-09/20/2023 |  | -1.289 | 7.9211e-05 | -1.2892 | 7.1533e-05 | 240 | 3.46
| 11916 | 11916 |  | 18922 | 1110 |  | sbe-3.5ksi | FuzzButton |  |  | 07/21-07/24/2023 | E |  | -0.132 |  |  | -1.002e-17 | 3.5903e-14 | -2.0609e-11 | -3.3741e-08 | -0.00095896 | -1.3047 | 0.9997 |  |  | -7.9968e-22 | 6.5862e-18 | -2.2016e-14 | 3.8818e-11 | -3.896e-08 | 2.3304e-05 | -1.3055 | 0.9981 | SBE |  | 11916 | 10/16-10/18/2023 |  | -1.3075 | 8.4415e-05 | -1.3077 | 9.1024e-05 | 200 | 2.84
| 11989 | 11989 | 22771 | 18988 | 1111 |  | sbe-3.5ksi | FuzzButton |  |  | 07/21-07/24/2023 | A |  | -0.187 |  |  |  | -4.7701e-15 | 2.389e-11 | -4.2587e-08 | -0.00087598 | -1.2659 | 0.9998 |  |  | -4.2375e-22 | 3.8742e-18 | -1.4104e-14 | 2.6474e-11 | -2.7443e-08 | 1.6878e-05 | -1.2665 | 0.9984 | SBE |  | 11989 | 09/29-10/03/2023 |  | -1.2713 | 0.00013107 | -1.2717 | 0.00014373 | 330 | 5.45
| 12543 | 12543 |  | 19038 | 1129 |  | sbe-3.5ksi | FuzzButton |  |  | 12/07-12/10/2023 | A |  | -0.033 |  |  |  | 8.2114e-15 | -2.9454e-11 | 1.8821e-08 | -0.0009514 | -1.5096 | 0.9998 | -5.6866e-28 | 3.5247e-24 | -4.7399e-21 | -1.4641e-17 | 5.7042e-14 | -7.544e-11 | 4.2532e-08 | -3.464e-06 | -1.5096 | 0.9988 | SBE |  | 12543 | 01/08-01/10/2024 |  | -1.5136 | 6.3534e-05 | -1.5137 | 5.1015e-05 | 110 | 4.02
CSIRO Float with new pH sensor (prior deployment with DF 250) | 12581 | 12581 | 12544 | 11434 | 838 | [5 empty] | 12/20-12/23/2023 | C |  | 0.014 | [5 empty] | -1.0912e-08 | -0.00094839 | -1.4385 | 0.9998 |  |  |  |  | -7.5269e-16 | 4.1226e-12 | -8.3539e-09 | 9.7417e-06 | -1.4388 | 0.9984 | SBE |  | 12581 | 04/02-04/03/2024 |  | -1.443 | 4.6999e-05 | -1.4432 | 5.7887e-05 | 210 | 4.48
| 20110 | 20110 |  | 20210 | 1134 | [5 empty] | 03/14-03/17/2024 | B |  | 0.005 | [5 empty] | -9.9409e-09 | -0.0010042 | -1.3555 | 0.9998 |  |  | 2.8325e-21 | -1.8875e-17 | 4.7571e-14 | -5.439e-11 | 2.4023e-08 | 3.5009e-06 | -1.3555 | 0.9973 | SBE |  | 20110 | 05/28-05/30/2024 |  | -1.3566 | 4.9799e-05 | -1.3568 | 4.5129e-05 | 140 | 1.11
| 20073 | 20073 |  | 20211 | 1133 | [5 empty] | 03/14-03/17/2024 | D |  | 0.128 |  |  |  |  | 8.6331e-12 | -2.6034e-08 | -0.0010134 | -1.3619 | 0.9997 | 2.7735e-27 | -2.5088e-23 | 9.4077e-20 | -1.8835e-16 | 2.1469e-13 | -1.3447e-10 | 3.6185e-08 | 5.1483e-06 | -1.3621 | 0.9976 | SBE |  | 20073 | 05/28-05/30/2024 |  | -1.3628 | 6.0898e-05 | -1.3629 | 5.2806e-05 | 120 | 0.9
| 20167 | 20167 |  | 20217 | 1136 | [5 empty] | 03/26-03/29/2024 | F |  | -0.056 | [5 empty] | -1.1889e-08 | -0.00092469 | -1.5641 | 0.9998 | 2.6853e-27 | -2.4451e-23 | 9.2295e-20 | -1.8578e-16 | 2.1265e-13 | -1.3458e-10 | 3.9327e-08 | 1.5326e-06 | -1.5643 | 0.9983 | SBE |  | 20167 | 06/10-06/12/2024 |  | -1.5663 | 4.5671e-05 | -1.5665 | 5.6079e-05 | 270 | 2.12
| 20066 | 20066 |  | 20220 | 1137 | [5 empty] | 01/27-01/30/2024 | F |  | -0.028 | [6 empty] | -0.0010268 | -1.3852 | 0.9999 |  | -2.7749e-24 | 2.1835e-20 | -6.9414e-17 | 1.1369e-13 | -1.0023e-10 | 4.3182e-08 | -2.709e-06 | -1.3852 | 0.9983 | SBE |  | 20066 | 06/10-06/12/2024 |  | -1.3897 | 7.3221e-05 | -1.3903 | 5.5599e-05 | 650 | 4.5
AOML refurb | 20162 | 20162 | 19097 | 12487 | 897 | [5 empty] | 03/26-03/29/2024 | C |  | -0.024 | [6 empty] | -0.0010087 | -1.3794 | 0.9998 |  | -4.1701e-24 | 3.3226e-20 | -1.0688e-16 | 1.7634e-13 | -1.5421e-10 | 6.2735e-08 | -2.7173e-06 | -1.3792 | 0.9962 | SBE |  | 20162 | 06/10-06/12/2024 |  | -1.3822 | 6.1297e-05 | -1.3824 | 5.085e-05 | 140 | 2.84
| 20067 | 20067 |  | 20212 | 1138 | [5 empty] | 01/27-01/30/2024 | G |  | -0.049 | [5 empty] | -1.2069e-08 | -0.001 | -1.3578 | 0.9998 |  | -2.3917e-24 | 1.9185e-20 | -6.1787e-17 | 1.0075e-13 | -8.4385e-11 | 2.9233e-08 | 4.5846e-06 | -1.3581 | 0.9986 | SBE |  | 20067 | 06/10-06/12/2024 |  | -1.3643 | 6.5575e-05 | -1.365 | 4.5082e-05 | 710 | 6.48
| 12537 | 12537 |  | 20219 | 1141 | [5 empty] | 12/01-12/04/2023 | D |  | -0.047 | [5 empty] | -5.8249e-09 | -0.00095104 | -1.4944 | 0.9998 |  |  | 5.3508e-21 | -3.6169e-17 | 9.4136e-14 | -1.1588e-10 | 6.4049e-08 | -6.8759e-06 | -1.4944 | 0.9974 | SBE |  | 12537 | 06/17-06/19/2024 |  | -1.508 | 4.7809e-05 | -1.5086 | 4.0388e-05 | 560 | 13.55
| 20044 | 20044 |  | 20209 | 1140 | [5 empty] | 01/27-01/30/2024 | D |  | 0.047 |  |  |  | 8.9561e-15 | -4.2071e-11 | 3.3629e-08 | -0.0010284 | -1.3735 | 0.9998 |  |  |  | 6.4467e-19 | -4.5867e-15 | 1.2615e-11 | -1.5778e-08 | 1.0404e-05 | -1.3739 | 0.9966 | SBE |  | 20044 | 06/17-06/19/2024 |  | -1.3766 | 5.1017e-05 | -1.3769 | 6.4393e-05 | 250 | 3.14
| 12538 | 12538 |  | 17090 | 1146 | [5 empty] | 12/01-12/04/2023 | J |  | 0.015 |  |  |  |  | -1.1095e-11 | 9.7692e-09 | -0.00091939 | -1.4916 | 0.9997 |  | -1.662e-24 | 1.4549e-20 | -5.239e-17 | 9.8759e-14 | -1.0054e-10 | 4.8266e-08 | -2.0587e-06 | -1.4918 | 0.9985 | SBE |  | 12538 | 07/15-07/17/2024 |  | -1.494 | 6.7709e-05 | -1.4941 | 6.4518e-05 | 130 | 2.32
| 20055 | 20055 |  | 20283 | 1147 | [5 empty] | 03/26-03/29/2024 | H |  | 0.055 |  |  | -9.2262e-18 | 4.9993e-14 | -9.1003e-11 | 3.3146e-08 | -0.00089477 | -1.5888 | 0.9998 |  | -2.4336e-24 | 1.7821e-20 | -5.1116e-17 | 7.1294e-14 | -4.6552e-11 | 7.7529e-09 | 7.4121e-06 | -1.5891 | 0.9985 | SBE |  | 20055 | 07/29-07/31/2024 |  | -1.5902 | 5.1019e-05 | -1.5904 | 5.5469e-05 | 230 | 1.41
| 20173 | 20173 |  | 20265 | 1150 | [5 empty] | 03/26-03/29/2024 | J |  | 0.035 |  | 2.2227e-20 | -1.3006e-16 | 2.7611e-13 | -2.4938e-10 | 7.0452e-08 | -0.00090438 | -1.5826 | 0.9998 |  | -3.6467e-24 | 2.9134e-20 | -9.4483e-17 | 1.5849e-13 | -1.4373e-10 | 6.4952e-08 | -7.6095e-06 | -1.5825 | 0.9975 | SBE |  | 20173 | 07/29-07/31/2024 |  | -1.5836 | 5.0659e-05 | -1.584 | 9.4271e-05 | 410 | 1.05
SO 2019 | [7 empty] | (Ver1 = Std PEEK design, Ver2 = double O-ring groove, Ver3 = smaller wire holes for Pt, Ver4 = Ti CE using roll pin to connect to green wire, Ver5 = pellet oring changed to 6.5x1.5 Buna70A)
| 251 | 251 | 18821 | 11426 | 840 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  | CE silver epoxy to wire | 11/26-12/3/2018 | A | 1@SO | 0.095 | [6 empty] | -0.00091988 | -1.4203 | 0.99998 |  |  | -7.5783e-22 | 5.3911e-18 | -1.4699e-14 | 1.778e-11 | -5.0921e-09 | -8.3684e-06 | -1.4204 | 0.99994 |  |  | 251 | 3/13-17/2019 |  | -1.4236 | 0.00027331 | -1.4237 | 0.00028088 | 116.1435 | 3.2571
| 252 | 252 | 18299 | 11436 | 843 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  | CE silver epoxy to wire | 11/9-15/2018 | A | 1@SO | 0.015 | [6 empty] | -0.00099635 | -1.3758 | 0.99995 |  |  | 2.305e-21 | -1.5847e-17 | 4.2735e-14 | -5.8391e-11 | 4.4584e-08 | -1.942e-05 | -1.3758 | 0.99845 |  |  | 252 | 4/18-19/2019 |  | -1.3932 | 5.2479e-05 | -1.3935 | 6.1982e-05 | 283.007 | 17.4901
| 253 | 253 | 17898 | 11438 | 845 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 2/6-12/2019 | A | 1@SO | 0.11 | [6 empty] | -0.00092937 | -1.4219 | 0.99997 |  |  | -2.6126e-21 | 1.7317e-17 | -4.2691e-14 | 4.6028e-11 | -1.5406e-08 | -8.7298e-06 | -1.4218 | 0.99972 |  |  | 253 | 5/6-8/2019 |  | -1.4272 | 9.1714e-05 | -1.429 | 7.7664e-05 | 1820.7393 | 5.3267
| 254 | 254 | 18081 | 11427 | 842 |  | sbe-3.5ksi | Ver3-PEEK, Pt |  | CE silver epoxy to wire | 11/26-12/3/2018 | B | 1@SO | 0.055 | [6 empty] | -0.0010282 | -1.4131 | 0.99998 |  |  | -1.468e-21 | 9.5178e-18 | -2.3485e-14 | 2.6376e-11 | -8.3019e-09 | -1.1065e-05 | -1.4133 | 0.99993 |  |  | 254 | 3/25-27/2019 |  | -1.4169 | 6.4524e-05 | -1.4172 | 4.6001e-05 | 291.5964 | 3.7648
| 255 | 255 | 12887 | 11437 | 847 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 2/6-12/2019 | B | 1@SO | -0.07 | [6 empty] | -0.0010425 | -1.4045 | 0.99995 |  |  | 3.3509e-22 | -2.4364e-18 | 7.4957e-15 | -1.3058e-11 | 1.4751e-08 | -1.1473e-05 | -1.4046 | 0.99975 |  |  | 255 | 5/28-30/2019 |  | -1.4087 | 4.5593e-05 | -1.4091 | 4.4342e-05 | 352.0789 | 4.2087
| 256 | 256 | 18608 | 11516 | 849 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 2/20-26/2019 | A | 1@SO | 0.04 | [6 empty] | -0.00092459 | -1.4051 | 0.99992 |  |  | -6.6292e-22 | 4.1275e-18 | -8.6861e-15 | 4.9091e-12 | 7.1284e-09 | -1.2529e-05 | -1.4052 | 0.99992 |  |  | 256 | 5/28-30/2019 |  | -1.4091 | 5.0753e-05 | -1.4094 | 5.4215e-05 | 324.3131 | 3.9803
| 257 | 257 | 17862 | 11441 | 848 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 2/20-26/2019 | B | 1@SO | 0.07 | [6 empty] | -0.000924 | -1.3861 | 0.99995 |  |  | -1.8667e-22 | 1.3626e-18 | -3.9176e-15 | 4.7163e-12 | 1.5609e-09 | -1.0608e-05 | -1.3862 | 0.99992 |  |  | 257 | 5/28-30/2019 |  | -1.3918 | 4.7969e-05 | -1.3922 | 5.5983e-05 | 436.4622 | 5.6576
| 258 | 258 | 18862 | 11518 | 864 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 5/27-6/3/2019 | A | 1@SO | 0.14 | [6 empty] | -0.00094401 | -1.3843 | 0.99997 |  |  | -1.8567e-21 | 1.2625e-17 | -3.2442e-14 | 3.6914e-11 | -1.169e-08 | -1.1417e-05 | -1.3844 | 0.99997 |  |  | 258 | 8/12-14/2019 |  | -1.3877 | 6.5255e-05 | -1.3888 | 4.9019e-05 | 1026.7837 | 3.4173
| 260 | 260 | 18242 | 11764 | 853 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 3/25-31/2019 | A | 1@SO | 0.08 | [6 empty] | -0.001003 | -1.4129 | 0.99997 |  |  | -1.5448e-21 | 9.7478e-18 | -2.1958e-14 | 1.865e-11 | 2.3801e-09 | -1.4123e-05 | -1.413 | 0.99992 |  |  | 260 | 6/30-7/2/2019 |  | -1.4174 | 0.0001119 | -1.4178 | 0.00011861 | 412.2406 | 4.5201
| 261 | 261 | 18771 | 11760 | 855 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 3/25-31/2019 | B | 1@SO | 0.02 | [6 empty] | -0.0010196 | -1.3693 | 0.99998 |  |  | -5.9564e-22 | 4.6095e-18 | -1.3625e-14 | 1.7689e-11 | -4.8711e-09 | -1.1591e-05 | -1.3694 | 0.99992 |  |  | 261 | 6/30-7/2/2019 |  | -1.3756 | 0.00010108 | -1.3757 | 0.00010073 | 165.7765 | 6.3057
| 262 | 262 | 18320 | 11765 | 858 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 3/30-4/6/2019 | C | 1@SO | 0.05 | [6 empty] | -0.0010934 | -1.3875 | 0.99998 |  |  | -1.3917e-21 | 9.4269e-18 | -2.3673e-14 | 2.4974e-11 | -4.4187e-09 | -1.1537e-05 | -1.3876 | 0.99991 |  |  | 262 | 7/15-17/2019 |  | -1.3911 | 5.0789e-05 | -1.3912 | 5.0789e-05 | 128.5423 | 3.6106
| 263 | 263 | 18677 | 11769 | 874 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 3/31-4/8/2019 | A | 1@SO | 0.04 | [6 empty] | -0.001022 | -1.4045 | 0.99999 |  |  | -6.5659e-22 | 3.8336e-18 | -7.6674e-15 | 5.0123e-12 | 2.6069e-09 | -6.1782e-06 | -1.4046 | 0.99961 |  |  | 263 | 9/30-10/2/2019 |  | -1.4088 | 4.1304e-05 | -1.4089 | 4.2735e-05 | 52.5455 | 4.3409
| 264 | 264 | 18545 | 11524 | 861 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 3/31-4/6/2019 | B | 1@SO | 0.12 | [6 empty] | -0.00092236 | -1.4384 | 0.99993 |  |  | -9.5725e-22 | 5.4875e-18 | -1.0425e-14 | 5.0372e-12 | 7.1969e-09 | -1.1941e-05 | -1.4385 | 0.99975 |  |  | 264 | 7/29-31/2019 |  | -1.4448 | 8.6602e-05 | -1.4449 | 9.8717e-05 | 127.5262 | 6.3894
| 265 | 265 | 18110 | 11515 | 860 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 4/17-23/2019 | A | 1@SO | 0.1 | [6 empty] | -0.00093649 | -1.4053 | 0.99996 |  |  | -1.3122e-21 | 9.3294e-18 | -2.54e-14 | 3.1787e-11 | -1.4488e-08 | -5.3698e-06 | -1.4054 | 0.99993 |  |  | 265 | 7/29-31/2019 |  | -1.4088 | 4.242e-05 | -1.4088 | 4.4006e-05 | 81.1802 | 3.4298
| 266 | 266 | 18643 | 11521 | 863 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 4/17-23/2019 | B | 1@SO | 0.102 | [6 empty] | -0.00090551 | -1.3967 | 0.99997 |  |  | -6.0423e-22 | 3.8896e-18 | -8.9067e-15 | 7.0125e-12 | 4.3764e-09 | -1.189e-05 | -1.3968 | 0.9999 |  |  | 266 | 8/12-14/2019 |  | -1.4009 | 4.7356e-05 | -1.4014 | 3.9904e-05 | 531.1725 | 4.1664
| 267 | 267 | 18097 | 11529 | 869 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 4/10-16/2019 | B | 1@SO | -0.03 | [6 empty] | -0.0010417 | -1.391 | 0.99999 |  |  | 4.8154e-22 | -2.8532e-18 | 6.4633e-15 | -7.8078e-12 | 7.7657e-09 | -8.5771e-06 | -1.3911 | 0.99959 |  |  | 267 | 9/21-23/2019 |  | -1.397 | 8.5817e-05 | -1.3974 | 7.3352e-05 | 325.2925 | 6.033
| 268 | 268 | 18861 | 11522 | 866 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 5/27-6/3/2019 | B | 1@SO | 0.12 | [6 empty] | -0.00093018 | -1.4124 | 0.99992 |  |  | -2.969e-22 | 3.1793e-18 | -1.2007e-14 | 1.9827e-11 | -1.1463e-08 | -5.3844e-06 | -1.4125 | 0.99987 |  |  | 268 | 8/24-27/2019 |  | -1.4151 | 0.00010432 | -1.4156 | 8.7688e-05 | 524.8762 | 2.7241
| 270 | 270 | 18161 | 11759 | 872 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 6/10-17/2019 | B | 1@SO | 0.04 | [6 empty] | -0.0010286 | -1.3967 | 0.99997 |  |  | -6.177e-22 | 4.3136e-18 | -1.1013e-14 | 1.1704e-11 | -1.9408e-09 | -7.2974e-06 | -1.3968 | 0.99993 |  |  | 270 | 9/21-23/2019 |  | -1.4018 | 4.5416e-05 | -1.4018 | 4.2712e-05 | 25.0265 | 5.0273
| 272 | 272 | 18450 | 11761 | 871 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 7/23-29/2019 | B | 1@SO | 0.18 | [6 empty] | -0.00095397 | -1.4161 | 0.99997 |  |  | -1.3921e-21 | 1.0308e-17 | -2.9516e-14 | 4.0358e-11 | -2.5169e-08 | 2.9245e-06 | -1.4161 | 0.99926 |  |  | 272 | 9/21-23/2019 |  | -1.4195 | 9.3592e-05 | -1.4198 | 9.0948e-05 | 239.0352 | 3.4543
| 273 | 273 | 18321 | 11768 | 875 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 7/30-8/8/2019 | A | 1@SO | 0.04 | [6 empty] | -0.0010278 | -1.412 | 0.99996 |  |  | -2.4286e-22 | 1.9726e-18 | -5.8613e-15 | 6.6427e-12 | 1.8462e-09 | -1.1035e-05 | -1.4121 | 0.99996 |  |  | 273 | 9/30-10/2/2019 |  | -1.4174 | 5.9626e-05 | -1.4177 | 5.0841e-05 | 279.3453 | 5.3771
| 275 | 275 | 18431 | 11763 | 873 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 8/26-9/1/2019 | A | 1@SO | 0.002 | [6 empty] | -0.0010626 | -1.3783 | 0.99993 |  |  | 1.9887e-21 | -1.4927e-17 | 4.4553e-14 | -6.693e-11 | 5.2821e-08 | -2.0714e-05 | -1.3784 | 0.99908 |  |  | 275 | 9/30-10/2/2019 |  | -1.3817 | 6.4171e-05 | -1.3818 | 6.2326e-05 | 95.4187 | 3.3498
| 276 | 276 | 18764 | 11762 | 876 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire | 8/26-9/1/2019 | B | 1@SO | 0.04 | [6 empty] | -0.0009616 | -1.4289 | 0.99996 |  |  | -6.1361e-22 | 2.7766e-18 | -1.3418e-15 | -1.1608e-11 | 2.4616e-08 | -2.0266e-05 | -1.429 | 0.99984 |  |  | 276 | 9/30-10/2/2019 |  | -1.431 | 5.3939e-05 | -1.4312 | 5.301e-05 | 189.0039 | 2.07
SO 2020 | [7 empty] | (Ver1 = Std PEEK design, Ver2 = double O-ring groove, Ver3 = smaller wire holes for Pt, Ver4 = Ti CE using roll pin to connect to green wire, Ver5 = pellet oring changed to 6.5x1.5 Buna70A, Ver6 = Ti CE roll pin crimp on 0.025 bus wire)
| 259 | 259 | 19139 | 12316 | 879 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  | CE roll pin to grn wire on all | 3/4-9/2019 | B | 1@SO | 0.18 | [6 empty] | -0.0008785 | -1.4356 | 0.99999 |  |  | -6.3834e-22 | 4.2409e-18 | -1.0344e-14 | 1.0328e-11 | -7.34e-10 | -6.9565e-06 | -1.4356 | 0.99991 |  |  | 259 | 1/13-15/2020 |  | -1.4398 | 6.2856e-05 | -1.44 | 6.7607e-05 | 194.5881 | 4.2137
| 269 | 269 | 18069 | 11525 | 877 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  |  | 6/10-17/2019 | A | 1@SO | 0.155 | [6 empty] | -0.00091848 | -1.4129 | 0.99999 |  |  | -1.5532e-22 | 1.2702e-18 | -3.6984e-15 | 4.0273e-12 | 1.2625e-09 | -7.2271e-06 | -1.4129 | 0.99995 |  |  | 269 | 1/13-15/2020 |  | -1.4384 | 8.3937e-05 | -1.4388 | 9.0827e-05 | 443.3587 | 25.5481
| 271 | 271 |  |  |  |  | sbe-3.5ksi | Ver4-PEEK, Pt |  |  | 8/10-16/2019 | A | 1@SO | 0.085 | [6 empty] | -0.00091585 | -1.4407 | 0.99993 |  |  | 2.0663e-22 | -2.0596e-18 | 8.2962e-15 | -1.7834e-11 | 2.2522e-08 | -1.6197e-05 | -1.4407 | 0.99987 |  |  | 271
| 274 | 274 | 19063 | 12308 | 880 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  |  | 8/10-16/2019 | C | 1@SO | 0.02 | [6 empty] | -0.0010398 | -1.4017 | 0.99995 |  |  | -8.0643e-22 | 5.3484e-18 | -1.3115e-14 | 1.317e-11 | -3.3572e-10 | -1.0274e-05 | -1.4018 | 0.99993 |  |  | 274 | 1/13-15/2020 |  | -1.4056 | 4.6828e-05 | -1.4062 | 5.0403e-05 | 582.9596 | 3.9081
| 277 | 277 | 19211 | 12311 | 883 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  |  | 8/26-9/1/2019 | C | 1@SO | 0.001 | [6 empty] | -0.0010989 | -1.3746 | 0.99997 |  |  | 8.0199e-23 | -1.9523e-18 | 1.0393e-14 | -2.3517e-11 | 2.6666e-08 | -1.6713e-05 | -1.3747 | 0.99987 |  |  | 277 | 1/27-29/2020 |  | -1.3783 | 4.9966e-05 | -1.3787 | 7.8362e-05 | 404.9364 | 3.6849
| 278 | 278 | 19644 | 12309 | 884 |  | sbe-3.5ksi | Ver4-PEEK, Pt |  |  | 8/26-9/1/2019 | D | 1@SO | 0.02 | [6 empty] | -0.0011273 | -1.3761 | 0.99985 |  |  | -1.5223e-21 | 8.7951e-18 | -1.612e-14 | 3.8421e-12 | 2.0511e-08 | -2.3067e-05 | -1.3762 | 0.99947 |  |  | 278 | 1/27-29/2020 |  | -1.3792 | 5.7273e-05 | -1.3794 | 4.8025e-05 | 262.7978 | 3.0793
| 279 | 279 | 18808 | 11523 | 890 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  | Pellet oring changed to 6.5 x 1.5 Buna 70A | 9/10-16/2019 | A | 1@SO | -0.045 | [6 empty] | -0.0010334 | -1.4035 | 0.99985 |  |  | 2.8891e-21 | -2.2618e-17 | 7.055e-14 | -1.1131e-10 | 9.2986e-08 | -3.8709e-05 | -1.4036 | 0.9997 |  |  | 279 | 3/11-13/2020 |  | -1.4067 | 6.7963e-05 | -1.4068 | 8.0681e-05 | 100.1482 | 3.2142 | Pellet oring changed to 6.5 x 1.5 Buna 70A
| 280 | 280 | 20627 | 12486 | 891 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 9/10-16/2019 | B | 1@SO | -0.08 | [6 empty] | -0.00097696 | -1.4009 | 0.99987 |  |  | -1.7619e-21 | 9.784e-18 | -1.6881e-14 | 3.1745e-12 | 1.9207e-08 | -1.883e-05 | -1.4009 | 0.99854 |  |  | 280 | 4/26-28/2021 |  | -1.4087 | 4.7698e-05 | -1.4091 | 6.5746e-05 | 400.7422 | 7.8116
| 281 | 281 | 17328 | 12483 | 892 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 9/18-23/2019 | C | 1@SO | 0.05 | [6 empty] | -0.0011099 | -1.4201 | 0.99996 |  |  | -7.3585e-23 | 3.2922e-19 | 2.2773e-16 | -4.3653e-12 | 1.2178e-08 | -1.5774e-05 | -1.4202 | 0.99992 |  |  | 281 | 3/11-13/2020 |  | -1.4249 | 6.2735e-05 | -1.4249 | 8.2286e-05 | 56.1704 | 4.8414
| 282 | 282 | 19881 | 12480 | 894 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 9/18-23/2019 | D | 1@SO | 0.11 | [6 empty] | -0.00095841 | -1.4181 | 0.99994 |  |  | -2.4232e-21 | 1.4703e-17 | -3.1247e-14 | 2.3599e-11 | 5.4923e-09 | -1.6286e-05 | -1.4182 | 0.99952 |  |  | 282 | 6/8-10/2020 |  | -1.4212 | 7.3256e-05 | -1.4214 | 8.5519e-05 | 249.6981 | 3.0853
| 283 | 283 | 19067 | 12479 | 896 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 1/13-18/2020 | C | 1@SO | 0.02 | [6 empty] | -0.0011788 | -1.373 | 0.99993 |  |  | 3.8702e-22 | -4.6027e-18 | 1.9461e-14 | -3.9204e-11 | 4.0891e-08 | -2.1261e-05 | -1.3731 | 0.99949 |  |  | 283 | 6/8-10/2020 |  | -1.379 | 4.5466e-05 | -1.3796 | 5.2801e-05 | 589.4386 | 6.0224
| 284 | 284 |  |  |  |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 1/13-18/2020 | D | 1@SO | 0.02 | [6 empty] | -0.0010603 | -1.3841 | 0.99984 |  |  | 1.1277e-21 | -9.2471e-18 | 3.0158e-14 | -5.0081e-11 | 4.5488e-08 | -2.1929e-05 | -1.3842 | 0.99946 |  |  | 284
| 285 | 285 | 19364 | 13306 | 893 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 1/20-26/2020 | A | 1@SO | 0.04 | [6 empty] | -0.00109 | -1.3967 | 0.99998 |  |  | 1.5654e-22 | -2.9235e-18 | 1.4471e-14 | -3.14e-11 | 3.3982e-08 | -1.7421e-05 | -1.3967 | 0.99817 |  |  | 285 | 6/22-24/2020 |  | -1.4004 | 0.00011912 | -1.4011 | 0.00014894 | 737.5844 | 3.71
AOML refurb was Apfid 19097 | 286 | 286 |  | 12487 | 897 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 1/20-26/2020 | B | 1@SO | 0.08 | [6 empty] | -0.00094874 | -1.4456 | 0.99996 |  |  | 1.0551e-21 | -8.7098e-18 | 2.869e-14 | -4.8437e-11 | 4.5028e-08 | -2.2068e-05 | -1.4456 | 0.99945 |  |  | 286 | 6/22-24/2020 |  | -1.4494 | 7.6995e-05 | -1.4494 | 8.8601e-05 | 59.6036 | 3.8297
| 287 | 287 | 19107 | 12457 | 899 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 2/5-11/2020 | D | 1@SO | 0.06 | [6 empty] | -0.0011331 | -1.3822 | 0.99984 |  |  | 1.0342e-21 | -9.0417e-18 | 3.137e-14 | -5.5085e-11 | 5.1966e-08 | -2.4931e-05 | -1.3823 | 0.9992 |  |  | 287 | 6/22-24/2020 |  | -1.3863 | 9.6912e-05 | -1.3867 | 0.0001395 | 408.7585 | 4.1536
| 288 | 288 | 19018 | 12446 | 898 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 3/12-18/2020 | A | 1@SO | 0.05 | [6 empty] | -0.0010253 | -1.4117 | 0.99997 |  |  | 2.1063e-22 | -3.1585e-18 | 1.4841e-14 | -3.1546e-11 | 3.3369e-08 | -1.7235e-05 | -1.4118 | 0.9994 |  |  | 288 | 8/10-12/2020 |  | -1.4179 | 6.6804e-05 | -1.4209 | 8.7043e-05 | 2980.4907 | 6.2274
| 289 | 289 | 19970 | 12452 | 902 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  | Ver 6 bus wire in CE crimp | 3/12-18/2020 | B | 1@SO | 0.045 | [6 empty] | -0.001125 | -1.4177 | 0.99998 |  |  | 1.4215e-21 | -1.0579e-17 | 3.0968e-14 | -4.526e-11 | 3.4813e-08 | -1.2969e-05 | -1.4178 | 0.99835 |  |  | 289 | 8/10-12/2020 |  | -1.4222 | 7.3296e-05 | -1.4231 | 8.5994e-05 | 868.0762 | 4.5012
| 290 | 290 | 19093 | 12459 | 904 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/12-18/2020 | C | 1@SO | 0.2 | [6 empty] | -0.0011714 | -1.3849 | 0.99976 |  |  | 1.2861e-21 | -1.0601e-17 | 3.4675e-14 | -5.7236e-11 | 5.0401e-08 | -2.2692e-05 | -1.385 | 0.99964 |  |  | 290 | 2/22-24/2021 |  | -1.39 | 4.4427e-05 | -1.3906 | 4.4586e-05 | 592.2956 | 5.0262
| 291 | 291 | 19531 | 12482 | 905 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 6/12-18/2020 | D | 1@SO | 0.08 | [6 empty] | -0.0011348 | -1.4206 | 0.99998 |  |  | 5.5136e-22 | -4.9838e-18 | 1.7205e-14 | -2.8873e-11 | 2.5178e-08 | -1.0986e-05 | -1.4206 | 0.99954 |  |  | 291 | 8/17-19/2020 |  | -1.4248 | 4.2682e-05 | -1.4252 | 3.9135e-05 | 426.2168 | 4.1548
| 292 | 292 | 19951 | 12455 | 909 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/1-7/2020 | C | 1@SO | 0.06 | [6 empty] | -0.0011156 | -1.4031 | 0.99995 |  |  | -1.2983e-21 | 8.0312e-18 | -1.7326e-14 | 1.2616e-11 | 5.8051e-09 | -1.3473e-05 | -1.4033 | 0.99971 |  |  | 292 | 10/26-28/2020 |  | -1.4097 | 9.713e-05 | -1.4107 | 9.305e-05 | 955.2133 | 6.5839
| 293 | 293 | 19443 | 12456 | 908 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/1-7/2020 | D | 1@SO | 0.065 | [6 empty] | -0.0010303 | -1.4019 | 0.99998 |  |  | -1.4211e-22 | -9.7896e-20 | 4.5894e-15 | -1.5804e-11 | 2.4328e-08 | -1.7525e-05 | -1.402 | 0.99963 |  |  | 293 | 10/26-28/2020 |  | -1.4066 | 0.00011011 | -1.4065 | 9.6519e-05 | -135.2001 | 4.6862
| 294 | 294 | 19319 | 12462 | 912 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/16-21/2020 | A | 1@SO | 0.065 | [6 empty] | -0.0010856 | -1.3994 | 0.9999 |  |  | 2.254e-23 | -8.495e-20 | 5.2613e-16 | -3.7729e-12 | 1.1849e-08 | -1.6083e-05 | -1.3995 | 0.99995 |  |  | 294 | 8/31-9/2/2020 |  | -1.4033 | 6.4304e-05 | -1.4042 | 4.839e-05 | 856.2315 | 3.8971
| 295 | 295 | 19073 | 12460 | 910 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/16-21/2020 | B | 1@SO | 0.02 | [6 empty] | -0.0010941 | -1.3882 | 0.99998 |  |  | 1.7138e-21 | -1.3236e-17 | 4.0373e-14 | -6.1657e-11 | 4.9441e-08 | -1.9069e-05 | -1.3883 | 0.99861 |  |  | 295 | 8/31-9/2/2020 |  | -1.3929 | 5.9435e-05 | -1.3937 | 5.3962e-05 | 756.8093 | 4.675
| 296 | 296 | 19356 | 13329 | 917 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 7/13-18/2020 | C | 1@SO | -0.05 | [6 empty] | -0.0010918 | -1.4265 | 0.99998 |  |  | 5.2312e-22 | -3.7789e-18 | 1.1089e-14 | -1.842e-11 | 2.2329e-08 | -2.0571e-05 | -1.4267 | 0.99979 |  |  | 296 | 2/22-24/2021 |  | -1.4329 | 4.1078e-05 | -1.4332 | 6.9592e-05 | 331.1642 | 6.4144
| 297 | 297 | 19378 | 13307 | 913 |  | sbe-3.5ksi | Ver5-PEEK, Pt |  |  | 9/28-10/4/2020 | A | 1@SO | 0.05 | [6 empty] | -0.0010596 | -1.3986 | 0.99997 |  |  | -2.3537e-22 | 1.4206e-18 | -2.9822e-15 | 9.6529e-13 | 7.4475e-09 | -1.3769e-05 | -1.3987 | 0.99987 |  |  | 297 | 2/8-10/2021 |  | -1.4051 | 4.4127e-05 | -1.4058 | 6.5745e-05 | 770.996 | 6.4249
| 298 | 298 | 19598 | 13328 | 916 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/29-7/5/2020 | A | 1@SO | 0.015 | [6 empty] | -0.0010043 | -1.4005 | 0.99994 |  |  | 2.1027e-21 | -1.5149e-17 | 4.2773e-14 | -6.0017e-11 | 4.4562e-08 | -1.7882e-05 | -1.4006 | 0.99901 |  |  | 298 | 2/16-18/2021 |  | -1.4055 | 7.5215e-05 | -1.4063 | 9.4234e-05 | 771.0211 | 5.0051
| 299 | 299 | 19180 | 13309 | 918 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 7/8-14/2020 | B | 1@SO | 0.02 | [6 empty] | -0.0010132 | -1.4512 | 0.99997 |  |  | 3.7655e-22 | -4.9729e-18 | 2.225e-14 | -4.6705e-11 | 5.073e-08 | -2.7756e-05 | -1.4513 | 0.99917 |  |  | 299 | 2/16-18/2021 |  | -1.4544 | 4.3108e-05 | -1.455 | 4.6067e-05 | 600.794 | 3.1488
| 300 | 300 | 19445 | 13580 | 920 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 7/15-21/2020 | A | 1@SO | 0.02 | [6 empty] | -0.0010421 | -1.3965 | 0.99998 |  |  | 4.7333e-22 | -3.5001e-18 | 1.0649e-14 | -1.8237e-11 | 2.1766e-08 | -1.9677e-05 | -1.3967 | 0.99991 |  |  | 300 | 3/8-10/2021 |  | -1.4015 | 5.8867e-05 | -1.4024 | 5.5747e-05 | 852.4527 | 5.0063
| 301 | 301 | 19821 | 12461 | 923 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/08-10/12/2020 | B |  | -0.034 |  |  |  |  | -1.9621e-11 | 4.7023e-08 | -0.0010953 | -1.4395 | 0.9998 |  |  | 9.6618e-22 | -8.0732e-18 | 2.7013e-14 | -4.6304e-11 | 4.3325e-08 | -2.0882e-05 | -1.439 | 0.9907 |  |  | 301 | 02/14-02/17/2023 |  | -1.4439 | 0.00097591 | -1.4439 | 0.00097822 | 0 | 4.39
| 302 | 302 | 19976 | 13622 | 924 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/6-12/2020 | A | 1@SO | 0.01 | [6 empty] | -0.001037 | -1.4397 | 0.99998 |  |  | 5.1659e-22 | -5.0125e-18 | 1.8403e-14 | -3.2881e-11 | 3.0206e-08 | -1.2327e-05 | -1.4397 | 0.99947 |  |  | 302 | 3/22-24/2021 |  | -1.4445 | 4.3484e-05 | -1.445 | 4.4516e-05 | 439.3749 | 4.8303
| 303 | 303 | 19441 | 13581 | 925 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/22-28/2020 | A | 1@SO | 0.05 | [6 empty] | -0.001064 | -1.4197 | 0.99997 |  |  | 1.8745e-21 | -1.5121e-17 | 4.8184e-14 | -7.7191e-11 | 6.5625e-08 | -2.7379e-05 | -1.4197 | 0.99876 |  |  | 303 | 3/22-24/2021 |  | -1.4259 | 6.3404e-05 | -1.4267 | 7.5075e-05 | 791.7205 | 6.2644
| 304 | 304 | 19389 | 13621 | 926 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/22-28/2020 | B | 1@SO | 0.055 | [6 empty] | -0.0010528 | -1.4049 | 0.99998 |  |  | 7.6934e-22 | -6.951e-18 | 2.4599e-14 | -4.3384e-11 | 3.9829e-08 | -1.6494e-05 | -1.4049 | 0.99852 |  |  | 304 | 3/22-24/2021 |  | -1.4094 | 5.3231e-05 | -1.4103 | 6.3776e-05 | 860.8015 | 4.5113
| 305 | 305 | 20932 | 13579 | 929 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/4-10/2020 | A | 1@SO | 0.095 | [6 empty] | -0.0010208 | -1.4443 | 0.99998 |  |  | 1.2279e-22 | -2.8414e-18 | 1.4912e-14 | -3.343e-11 | 3.6815e-08 | -1.8527e-05 | -1.4443 | 0.99974 |  |  | 305 | 4/12-14/2021 |  | -1.4479 | 4.9129e-05 | -1.4483 | 5.1549e-05 | 406.233 | 3.6205
| 306 | 306 | 19806 | 13310 | 930 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/4-10/2020 | B | 1@SO | 0.075 | [6 empty] | -0.0010587 | -1.4167 | 0.99998 |  |  | -2.9478e-22 | -4.335e-19 | 1.0154e-14 | -3.0271e-11 | 3.7759e-08 | -2.0784e-05 | -1.4168 | 0.99707 |  |  | 306 | 4/26-28/2021 |  | -1.4207 | 6.5906e-05 | -1.4214 | 6.0042e-05 | 739.4619 | 3.978
SO 2021
| 307 | 307 | 20108 | 13781 | 932 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 2/9-15/2021 | A | 1@SO | 0.01 | [6 empty] | -0.0010188 | -1.4409 | 0.99992 |  |  | -1.3877e-22 | -1.373e-18 | 1.2346e-14 | -3.3214e-11 | 4.0972e-08 | -2.355e-05 | -1.4409 | 0.99913 |  |  | 307 | 5/10-12/2021 |  | -1.4449 | 3.6233e-05 | -1.4451 | 3.472e-05 | 217.988 | 3.9895
| 308 | 308 | 20926 | 13625 | 931 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 2/9-15/2021 | B | 1@SO | 0.01 | [6 empty] | -0.0010776 | -1.4351 | 0.99997 |  |  | -2.7402e-21 | 1.6601e-17 | -3.5596e-14 | 2.826e-11 | 3.5562e-09 | -1.5212e-05 | -1.4351 | 0.99868 |  |  | 308 | 5/10-12/2021 |  | -1.4401 | 3.9486e-05 | -1.4407 | 5.0218e-05 | 568.6141 | 5.0656
| 309 | 309 | 20704 | 14209 | 934 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 2/18-24/2021 | A | 1@SO | 0 | [6 empty] | -0.0010361 | -1.45 | 0.99993 |  |  | 2.8895e-22 | -3.6349e-18 | 1.6305e-14 | -3.5396e-11 | 4.092e-08 | -2.5073e-05 | -1.4501 | 0.99969 |  |  | 309 | 5/17-19/2021 |  | -1.455 | 7.93e-05 | -1.4557 | 0.00011473 | 680.2354 | 4.9332
| 310 | 310 | 20328 | 14220 | 937 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 2/18-24/2021 | B | 1@SO | 0.03 | [6 empty] | -0.0010486 | -1.431 | 0.99994 |  |  | 3.9711e-21 | -2.7452e-17 | 7.3895e-14 | -9.8949e-11 | 7.3178e-08 | -3.5556e-05 | -1.4312 | 0.99934 |  |  | 310 | 5/17-19/2021 |  | -1.4433 | 5.0272e-05 | -1.4441 | 7.7555e-05 | 741.3226 | 12.3041
| 311 | 311 | 20804 | 14219 | 939 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 3/8-15/2021 | A | 1@SO | 0.02 | [6 empty] | -0.0010379 | -1.4347 | 0.99994 |  |  | -7.4225e-22 | 4.5717e-18 | -9.9977e-15 | 7.6791e-12 | 4.6974e-09 | -1.3549e-05 | -1.4348 | 0.99983 |  |  | 311 | 6/2-3/2021 |  | -1.4395 | 3.9215e-05 | -1.4397 | 7.9576e-05 | 240.6426 | 4.8166
| 312 | 312 | 20970 | 14292 | 941 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  | Go-BGC | 3/2-7/2021 | B | 1@SO | 0.01 | [6 empty] | -0.0010534 | -1.4216 | 0.99985 |  |  | 3.0972e-21 | -2.1507e-17 | 5.8074e-14 | -7.8105e-11 | 5.9809e-08 | -3.3705e-05 | -1.4218 | 0.99957 |  |  | 312 | 6/7-9/2021 |  | -1.4275 | 3.9432e-05 | -1.4288 | 4.7765e-05 | 1378.6824 | 5.8357
| 313 | 313 | 20132 | 14302 | 944 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 3/8-15/2021 | B | 1@SO | 0 | [6 empty] | -0.0010021 | -1.4436 | 0.99986 |  |  | 5.9817e-22 | -4.5194e-18 | 1.3979e-14 | -2.3834e-11 | 2.6756e-08 | -2.0776e-05 | -1.4437 | 0.99956 |  |  | 313 | 7/15-19/2021 |  | -1.4483 | 8.4207e-05 | -1.4487 | 8.8513e-05 | 413.3786 | 4.6841
| 314 | 314 | 20175 | 14434 | 945 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 5/7-13/2021 | A | 1@SO | 0.02 | [6 empty] | -0.0010467 | -1.4317 | 0.99996 |  |  | -2.885e-22 | 2.0315e-18 | -5.3971e-15 | 5.1627e-12 | 4.1682e-09 | -1.2364e-05 | -1.4318 | 0.99956 |  |  | 314 | 7/15-19/2021 |  | -1.4347 | 7.6918e-05 | -1.4351 | 7.4759e-05 | 367.4626 | 2.9866
| 315 | 315 | 20912 | 14437 | 947 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 5/7-13/2021 | B | 1@SO | 0.02 | [6 empty] | -0.0010579 | -1.4399 | 0.99994 |  |  | -8.9996e-22 | 6.0579e-18 | -1.5631e-14 | 1.8241e-11 | -5.3154e-09 | -8.5134e-06 | -1.44 | 0.99981 |  |  | 315 | 7/26-28/2021 |  | -1.4436 | 6.5992e-05 | -1.4441 | 6.9397e-05 | 529.3863 | 3.6588
| 316 | 316 | 20137 | 14436 | 948 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 3/25-4/1/2021 | A | 1@SO | 0.02 | [6 empty] | -0.0010546 | -1.415 | 0.99996 |  |  | 1.1128e-21 | -7.2309e-18 | 1.7791e-14 | -2.1427e-11 | 1.7341e-08 | -1.5981e-05 | -1.4152 | 0.99954 |  |  | 316 | 7/26-28/2021 |  | -1.4181 | 5.8349e-05 | -1.4183 | 5.682e-05 | 233.6764 | 3.0478
| 317 | 317 | 20162 | 14435 | 951 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 3/25-4/1/2021 | B | 1@SO | 0 | [6 empty] | -0.0010749 | -1.4374 | 0.99986 |  |  | 1.3983e-21 | -9.1305e-18 | 2.3231e-14 | -3.1003e-11 | 2.8542e-08 | -2.2707e-05 | -1.4375 | 0.99938 |  | MATLAB | 317 | 08/03-08/05/2021 |  | -1.4407 | 8.3154e-05 | -1.4411 | 8.3775e-05 | 370 | 3.3054
| 318 | 318 | 20913 | 14303 | 956 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 5/14-20/2021 | A | 1@SO | 0.05 | [6 empty] | -0.0010058 | -1.397 | 0.99997 |  |  | 1.2205e-21 | -9.7436e-18 | 3.0476e-14 | -4.7452e-11 | 3.866e-08 | -1.4487e-05 | -1.397 | 0.999 |  |  | 318 | 08/16-08/18/2021 |  | -1.3997 | 5.6319e-05 | -1.3998 | 6.1757e-05 | 100 | 2.7399
| 319 | 319 | 20358 | 14603 | 955 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 5/14-20/2021 | B | 1@SO | 0.02 | [6 empty] | -0.0010999 | -1.4242 | 0.99997 |  |  | -1.1211e-22 | -6.9778e-19 | 7.3417e-15 | -2.0507e-11 | 2.5881e-08 | -1.3965e-05 | -1.4243 | 0.99755 |  |  | 319 | 08/16-08/18/2021 |  | -1.4291 | 6.8992e-05 | -1.4291 | 7.8451e-05 | 60 | 4.831
| 320 | 320 | 20632 | 14602 | 954 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 5/21-27/2021 | A | 1@SO | 0.04 | [6 empty] | -0.0010397 | -1.4379 | 0.99998 |  |  | 1.7257e-21 | -1.1963e-17 | 3.1577e-14 | -3.9367e-11 | 2.3223e-08 | -4.2485e-06 | -1.4379 | 0.99884 |  |  | 320 | 08/16-08/18/2021 |  | -1.4417 | 6.5554e-05 | -1.442 | 7.5169e-05 | 270 | 3.7868
| 321 | 321 | 21996 | 16648 | 957 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 05/23-05/27/2021 | B |  | -0.064 | [6 empty] | -0.0010407 | -1.4264 | 0.9998 |  |  | 8.9186e-22 | -6.6115e-18 | 1.9113e-14 | -2.6771e-11 | 1.8145e-08 | -3.8078e-06 | -1.4267 | 0.9657 |  |  | 321 | 01/31-02/02/2022 |  | -1.4316 | 7.9544e-05 | -1.4317 | 6.3117e-05 | 140 | 5.15
| 322 | 322 | 21900 | 16647 | 958 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/04-06/08/2021 | A |  | -0.027 | [6 empty] | -0.0010758 | -1.4259 | 0.9995 |  |  | 2.4887e-21 | -1.7109e-17 | 4.5601e-14 | -5.9049e-11 | 3.812e-08 | -1.0304e-05 | -1.4261 | 0.8319 |  |  | 322 | 01/31-02/02/2022 |  | -1.4303 | 0.00010193 | -1.4307 | 7.9552e-05 | 390 | 4.37
| 323 | 323 | 20001 | 14605 | 959 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 6/2-8/2021 | B | 1@SO | 0.02 | [6 empty] | -0.001072 | -1.3959 | 0.99999 |  |  | 1.2746e-21 | -1.0433e-17 | 3.3422e-14 | -5.3312e-11 | 4.4554e-08 | -1.7156e-05 | -1.3959 | 0.99534 |  |  | 323 | 08/23-08/25/2021 |  | -1.4013 | 0.00011387 | -1.4016 | 9.4013e-05 | 240 | 5.4263
MATLAB
| 324 | 324 | 17182 | 15143 | 966 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/23-07/27/2021 | A |  | -0.127 | [6 empty] | -0.0010656 | -1.4256 | 0.9993 |  |  | 7.891e-22 | -6.2179e-18 | 1.9753e-14 | -3.2479e-11 | 2.9592e-08 | -1.3544e-05 | -1.4256 | 0.932 |  |  | 324 | 09/20-09/21/2021 |  | -1.4342 | 5.1912e-05 | -1.4347 | 4.8892e-05 | 500
| 325 | 325 | 20549 | 15166 | 965 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/23-07/27/2021 | B |  | -0.077 | [6 empty] | -0.0010796 | -1.3887 | 0.9997 |  |  | 2.4465e-21 | -1.6817e-17 | 4.513e-14 | -5.9865e-11 | 4.0624e-08 | -1.2877e-05 | -1.3888 | 0.9215 |  |  | 325 | 09/08-09/10/2021 |  | -1.3935 | 0.00010239 | -1.3935 | 7.2393e-05 | 30 | 4.76
| 326 | 326 | 20842 | 15329 | 964 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/05-08/09/2021 | A |  | -0.014 | [6 empty] | -0.0010908 | -1.4025 | 0.9989 |  |  | -2.1573e-22 | 1.4827e-18 | -2.887e-15 | -1.5108e-12 | 1.2393e-08 | -1.5826e-05 | -1.4027 | 0.9918 |  |  | 326 | 09/08-09/10/2021 |  | -1.4068 | 7.9948e-05 | -1.4069 | 7.4978e-05 | 160 | 4.27
| 327 | 327 | 20603 | 15370 | 969 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/13-08/18/2021 | A |  | -0.001 | [6 empty] | -0.0010983 | -1.4127 | 0.9994 |  |  | 1.1458e-21 | -8.6154e-18 | 2.5771e-14 | -3.8923e-11 | 3.1154e-08 | -1.1962e-05 | -1.4128 | 0.8634 |  |  | 327 | 09/20-09/21/2021 |  | -1.4176 | 0.00010951 | -1.4182 | 0.00012887 | 650 | 4.9
| 328 | 328 | 20075 | 15366 | 967 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/30-08/03/2021 | A |  | -0.033 | [6 empty] | -0.0010739 | -1.3925 | 0.9996 |  |  | -7.4521e-22 | 4.2237e-18 | -7.2955e-15 | 2.2068e-13 | 1.1978e-08 | -1.1039e-05 | -1.3925 | 0.9746 |  |  | 328 | 09/08-09/10/2021 |  | -1.3964 | 9.5761e-05 | -1.3966 | 7.7765e-05 | 230 | 3.91
| 329 | 329 | 20547 | 15368 | 968 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/30-08/03/2021 | B |  | -0.105 | [6 empty] | -0.0010773 | -1.4051 | 0.9999 |  |  | 2.0495e-21 | -1.4307e-17 | 3.896e-14 | -5.2079e-11 | 3.5358e-08 | -1.0137e-05 | -1.4052 | 0.9693 |  |  | 329 | 09/20-09/21/2021 |  | -1.4093 | 5.6751e-05 | -1.4097 | 9.6765e-05 | 400 | 4.21
| 330 | 330 | 20528 | 15367 | 970 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/06-08/10/2021 | A |  | -0.058 | [6 empty] | -0.0011158 | -1.3928 | 0.9997 |  |  | -9.4643e-22 | 6.3493e-18 | -1.5771e-14 | 1.6296e-11 | -2.2553e-09 | -7.5488e-06 | -1.3928 | 0.9879 |  |  | 330 | 11/01-11/03/2021 |  | -1.3984 | 7.7281e-05 | -1.3993 | 7.1139e-05 | 830 | 5.61
| 331 | 331 | 20618 | 15369 | 972 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/06-08/10/2021 | B |  | -0.063 | [6 empty] | -0.0010489 | -1.4208 | 0.9994 |  |  | 3.5119e-22 | -3.6738e-18 | 1.4624e-14 | -2.8406e-11 | 2.853e-08 | -1.3363e-05 | -1.4209 | 0.9191 |  |  | 331 | 09/27-09/29/2021 |  | -1.4255 | 4.892e-05 | -1.4258 | 4.0705e-05 | 340 | 4.66
| 332 | 332 | 20496 | 15479 | 973 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/13-08/18/2021 | B |  | 0.014 | [6 empty] | -0.0010832 | -1.4138 | 0.9995 |  |  | 1.9291e-21 | -1.381e-17 | 3.8903e-14 | -5.4697e-11 | 4.0778e-08 | -1.5278e-05 | -1.4139 | 0.9542 |  |  | 332 | 10/04-10/06/2021 |  | -1.4173 | 6.1498e-05 | -1.4181 | 6.1077e-05 | 720 | 3.56
| 333 | 333 | 20343 | 16078 | 983 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/21-08/25/2021 | A |  | -0.045 | [6 empty] | -0.0010617 | -1.3674 | 0.9997 |  |  | 5.2039e-22 | -3.7364e-18 | 1.0867e-14 | -1.6521e-11 | 1.4377e-08 | -6.5691e-06 | -1.3674 | 0.9119 |  |  | 333 | 11/15-11/17/2021 |  | -1.3761 | 9.967e-05 | -1.3766 | 9.5825e-05 | 500 | 8.71
| 334 | 334 | 19993 | 12445 | 985 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/21-08/25/2021 | B |  | -0.029 | [6 empty] | -0.001105 | -1.3657 | 0.9998 |  |  | -6.6119e-22 | 4.3603e-18 | -1.0339e-14 | 9.3323e-12 | 1.094e-09 | -6.7578e-06 | -1.3656 | 0.9885 |  |  | 334 | 12/06-12/08/2021 |  | -1.3722 | 5.2865e-05 | -1.3736 | 5.6009e-05 | 1410 | 6.46
| 335 | 335 | 20329 | 16079 | 982 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/28-09/01/2021 | B |  | -0.016 | [6 empty] | -0.0010516 | -1.414 | 0.9996 |  |  | 3.1788e-21 | -2.19e-17 | 5.8807e-14 | -7.7279e-11 | 5.0926e-08 | -1.4024e-05 | -1.4141 | 0.9491 | DF |  | 335 | 11/15-11/17/2021 |  | -1.4182 | 4.952e-05 | -1.4185 | 5.1061e-05 | 270 | 4.16
| 336 | 336 | 20589 | 16081 | 981 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/28-09/01/2021 | A |  | 0.01 | [6 empty] | -0.0010387 | -1.4429 | 0.9999 |  |  | -5.7617e-22 | 3.6697e-18 | -8.1096e-15 | 5.4218e-12 | 5.1575e-09 | -8.3983e-06 | -1.443 | 0.9896 | DF |  | 336 | 11/15-11/17/2021 |  | -1.4469 | 4.8318e-05 | -1.4473 | 5.4461e-05 | 380 | 3.97
| 337 | 337 | 20832 | 16082 | 988 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/10-09/15/2021 | A |  | -0.022 | [6 empty] | -0.0010965 | -1.3744 | 0.9998 |  |  | 2.0257e-23 | 7.9582e-19 | -4.7793e-15 | 8.3836e-12 | -1.4547e-09 | -8.9844e-06 | -1.3743 | 0.9985 |  |  | 337 | 12/06-12/08/2021 |  | -1.3773 | 0.00011715 | -1.3775 | 9.0234e-05 | 150 | 2.89
| 338 | 338 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/10-09/15/2021 | B |  | 0.004 | [6 empty] | -0.001112 | -1.4205 | 0.9997 |  |  | 2.3791e-21 | -1.5699e-17 | 4.0579e-14 | -5.3439e-11 | 4.0665e-08 | -1.9907e-05 | -1.4205 | 0.9961 |  |  | 338
| 339 | 339 | 20579 | 16083 | 986 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/17-09/21/2021 | A |  | -0.056 | [6 empty] | -0.0010481 | -1.3824 | 0.9999 |  |  | 8.3429e-22 | -4.4478e-18 | 8.0942e-15 | -6.8091e-12 | 7.2077e-09 | -1.0019e-05 | -1.3823 | 0.9977 |  |  | 339 | 12/06-12/08/2021 |  | -1.3865 | 7.5617e-05 | -1.3867 | 6.7157e-05 | 210 | 4.17
| 340 | 340 | 19006 | 13780 | 992 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/17-09/21/2021 | B |  | -0.053 | [6 empty] | -0.0010865 | -1.3833 | 0.9995 |  |  | 8.5956e-23 | -1.5194e-18 | 7.8644e-15 | -1.8861e-11 | 2.458e-08 | -1.7244e-05 | -1.3831 | 0.9924 |  |  | 340 | 12/13-12/15/2021 |  | -1.3876 | 6.9305e-05 | -1.3879 | 8.1681e-05 | 330 | 4.31
| 341 | 341 | 20602 | 13778 | 989 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/25-09/27/2021 | A |  | -0.099 | [6 empty] | -0.0010725 | -1.4284 | 0.9997 |  |  | 2.9799e-22 | -5.9912e-20 | -5.9029e-15 | 1.507e-11 | -9.7434e-09 | -4.8584e-06 | -1.4283 | 0.9962 |  |  | 341 | 12/13-12/15/2021 |  | -1.4363 | 7.7548e-05 | -1.4365 | 7.2163e-05 | 220 | 7.9
| 342 | 342 | 20728 | 16080 | 984 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/05-10/08/2021 | B |  | -0.12 | [6 empty] | -0.0011165 | -1.3691 | 0.9998 |  |  | 1.2974e-21 | -7.3365e-18 | 1.4672e-14 | -1.3484e-11 | 1.09e-08 | -1.2271e-05 | -1.369 | 0.9975 |  |  | 342 | 11/15-11/17/2021 |  | -1.3719 | 6.6754e-05 | -1.3721 | 6.4968e-05 | 190 | 2.77
| 343 | 343 | 20336 | 16131 | 987 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/05-10/08/2021 | A |  | -0.116 | [6 empty] | -0.0010699 | -1.4123 | 0.9999 |  |  | 4.2053e-21 | -2.6899e-17 | 6.5386e-14 | -7.6509e-11 | 4.7735e-08 | -2.0404e-05 | -1.4122 | 0.9983 |  |  | 343 | 12/06-12/08/2021 |  | -1.4155 | 5.1809e-05 | -1.4164 | 0.00024802 | 880 | 3.18
| 344 | 344 | 20352 | 14101 | 993 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/15-10/19/2021 | A |  | -0.07 | [6 empty] | -0.0010994 | -1.409 | 0.9999 |  |  | 2.565e-21 | -1.7887e-17 | 4.9451e-14 | -7.0045e-11 | 5.5473e-08 | -2.5337e-05 | -1.4088 | 0.9968 |  |  | 344 | 12/20-12/22/2021 |  | -1.4127 | 6.7568e-05 | -1.4133 | 7.4659e-05 | 600 | 3.7
| 345 | 345 | 20039 | 14102 | 994 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/15-10/19/2021 | B |  | -0.06 | [6 empty] | -0.001109 | -1.394 | 0.9995 |  |  | 2.6286e-21 | -1.6994e-17 | 4.2622e-14 | -5.3858e-11 | 3.9369e-08 | -1.9791e-05 | -1.3939 | 0.9941 |  |  | 345 | 12/20-12/22/2021 |  | -1.3992 | 7.5554e-05 | -1.3995 | 5.3864e-05 | 250 | 5.22
| 346 | 346 | 21891 | 16649 | 1000 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/09-12/13/2021 | A |  | -0.151 | [6 empty] | -0.00083883 | -1.2667 | 0.9998 |  |  | -2.8169e-22 | 1.2887e-18 | -6.7856e-16 | -5.5486e-12 | 1.2482e-08 | -9.6738e-06 | -1.2667 | 0.9927 |  |  | 346 | 02/01-02/02/2022 |  | -1.272 | 6.1278e-05 | -1.273 | 0.00011641 | 950 | 5.36
| 347 | 347 | 21475 | 16669 | 999 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/09-12/13/2021 | B |  | 0.001 | [6 empty] | -0.00090215 | -1.2648 | 0.9998 |  |  | -3.8374e-23 | -6.1283e-19 | 5.3739e-15 | -1.5901e-11 | 2.2737e-08 | -1.5581e-05 | -1.2649 | 0.9918 |  |  | 347 | 01/10-01/12/2022 |  | -1.2708 | 8.3079e-05 | -1.2715 | 6.6283e-05 | 660 | 6.02
| 348 | 348 | 21279 | 16667 | 1003 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/16-12/18/2021 | A |  | -0.305 | [6 empty] | -0.00079175 | -1.3046 | 0.9997 |  |  | 1.4555e-21 | -9.9095e-18 | 2.6065e-14 | -3.2938e-11 | 2.0209e-08 | -4.135e-06 | -1.3047 | 0.9699 |  |  | 348 | 02/28-03/02/2022 |  | -1.3097 | 0.00014458 | -1.3103 | 0.00036904 | 670 | 5.1
| 349 | 349 | 21146 | 17088 | 1001 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/16-12/18/2021 | B |  | -0.116 | [6 empty] | -0.00078648 | -1.29 | 0.9998 |  |  | 6.56e-24 | -6.2225e-19 | 4.4294e-15 | -1.2664e-11 | 1.7793e-08 | -1.211e-05 | -1.2901 | 0.9885 |  |  | 349 | 02/28-03/02/2022 |  | -1.2966 | 6.0648e-05 | -1.2994 | 0.00025411 | 2820 | 6.54
| 350 | 350 | 21874 | 16958 | 1005 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/23-12/25/2021 | A |  | -0.015 | [6 empty] | -0.00092824 | -1.2916 | 0.9999 |  |  | 2.0592e-21 | -1.2888e-17 | 3.1652e-14 | -4.0101e-11 | 3.113e-08 | -1.6541e-05 | -1.2915 | 0.9958 |  |  | 350 | 03/07-03/10/2022 |  | -1.2957 | 0.00010192 | -1.296 | 0.00019185 | 310 | 4.03
| 351 | 351 | 21827 | 16959 | 1008 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/30-01/03/2022 | B |  | -0.103 | [6 empty] | -0.0007499 | -1.3328 | 0.9994 |  |  | 7.0229e-22 | -5.4504e-18 | 1.6808e-14 | -2.6184e-11 | 2.1747e-08 | -8.6043e-06 | -1.3329 | 0.8776 |  |  | 351 | 03/07-03/10/2022 |  | -1.3375 | 0.00018068 | -1.3377 | 0.00025247 | 240 | 4.69
| 352 | 352 | 20265 | 16801 | 1014 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/30-01/03/2022 | A |  | -0.071 | [6 empty] | -0.00085636 | -1.3024 | 0.9995 |  |  | -1.4668e-21 | 9.8398e-18 | -2.4465e-14 | 2.5904e-11 | -6.7147e-09 | -5.5026e-06 | -1.3023 | 0.9843 |  |  | 352 | 03/21-03/23/2022 |  | -1.3074 | 0.00032443 | -1.3076 | 0.00047383 | 220 | 5.01
| 353 | 353 | 21848 | 17346 | 1070 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/03-11/07/2022 | A |  | -0.134 |  |  |  | 4.4708e-14 | -1.3666e-10 | 1.0721e-07 | -0.00082958 | -1.3296 | 0.9996 |  |  | -2.1944e-22 | 8.07e-19 | 8.3576e-16 | -8.3573e-12 | 1.5372e-08 | -1.0044e-05 | -1.3296 | 0.9727 | DF |  | 353 | 12/19-12/22/2022 |  | -1.3417 | 8.2173e-05 | -1.3423 | 8.8439e-05 | 600 | 12.11
| 354 | 354 | 22751 | 16799 | 1013 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 01/20-01/24/2022 | B |  | -0.013 | [5 empty] | 9.8035e-09 | -0.0010343 | -1.3919 | 0.9998 | 8.2911e-27 | -7.5807e-23 | 2.8935e-19 | -5.9671e-16 | 7.1832e-13 | -5.0965e-10 | 2.0778e-07 | -4.9327e-05 | -1.3913 | 0.9987 | DF |  | 354 | 03/21-03/23/2022 |  | -1.4023 | 7.3423e-05 | -1.4031 | 0.00036992 | 750 | 10.46
| 355 | 355 | 21960 | 16800 | 1015 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 01/28-01/30/2022 | A |  | -0.009 | [6 empty] | -0.0010836 | -1.4021 | 0.9999 |  |  | 1.1696e-21 | -8.9976e-18 | 2.7853e-14 | -4.4397e-11 | 3.8999e-08 | -1.8267e-05 | -1.4023 | 0.9894 |  |  | 355 | 03/21-03/24/2022 |  | -1.4065 | 6.2474e-05 | -1.4066 | 4.7586e-05 | 150 | 4.42
| 356 | 356 | 21888 | 16798 | 1016 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 01/28-01/30/2022 | B |  | 0.025 | [6 empty] | -0.001094 | -1.3943 | 0.9999 |  |  | 1.0037e-23 | -1.0257e-18 | 6.8542e-15 | -1.8606e-11 | 2.6536e-08 | -2.1248e-05 | -1.3944 | 0.9979 |  |  | 356 | 03/21-03/24/2022 |  | -1.3983 | 4.7893e-05 | -1.3985 | 0.00011714 | 280 | 3.98
| 357 | 357 | 21792 | 17507 | 1023 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/03-04/05/2022 | A |  | -0.036 | [5 empty] | -5.8958e-09 | -0.0010632 | -1.4305 | 0.9995 |  |  | 3.4662e-21 | -2.2043e-17 | 5.4115e-14 | -6.547e-11 | 4.2882e-08 | -1.6791e-05 | -1.4307 | 0.9874 | DF |  | 357 | 10/10-10/12/2022 |  | -1.4352 | 4.9414e-05 | -1.436 | 7.1651e-05 | 850 | 4.66
| 358 | 358 | 21311 | 17175 | 1017 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 02/27-03/03/2022 | B |  | -0.072 | [6 empty] | -0.0011095 | -1.4047 | 0.9997 |  |  | -7.2119e-22 | 3.4919e-18 | -3.7199e-15 | -6.4561e-12 | 1.7138e-08 | -1.1868e-05 | -1.4046 | 0.9851 | DF |  | 358 | 04/25-04/27/2022 |  | -1.408 | 9.2123e-05 | -1.4102 | 0.00010366 | 2230 | 3.27
| 359 | 359 | 21977 | 17194 | 1027 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/03-04/05/2022 | B |  | -0.108 | [6 empty] | -0.0010253 | -1.4575 | 0.9998 | 5.7837e-27 | -4.9473e-23 | 1.7326e-19 | -3.1918e-16 | 3.3115e-13 | -1.9425e-10 | 6.5439e-08 | -1.6462e-05 | -1.4572 | 0.9923 | DF |  | 359 | 05/16-05/18/2022 |  | -1.4629 | 0.00011541 | -1.4634 | 0.00023886 | 460 | 5.47
| 360 | 360 | 21286 | 17195 | 1026 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/07-04/09/2022 | A |  | -0.031 | [6 empty] | -0.001004 | -1.4099 | 0.9998 | 6.0699e-27 | -5.4163e-23 | 1.9942e-19 | -3.895e-16 | 4.3105e-13 | -2.6792e-10 | 9.0751e-08 | -2.0867e-05 | -1.4099 | 0.998 | DF |  | 360 | 05/16-05/18/2022 |  | -1.4142 | 9.303e-05 | -1.4142 | 8.1912e-05 | 40 | 4.3
| 361 | 361 | 21026 | 17213 | 1025 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/15-04/17/2022 | B |  | -0.049 | [6 empty] | -0.0010265 | -1.4251 | 0.9999 | 1.1695e-26 | -1.0023e-22 | 3.5319e-19 | -6.5958e-16 | 7.0241e-13 | -4.3029e-10 | 1.5096e-07 | -3.3079e-05 | -1.4248 | 0.9955 | DF |  | 361 | 05/16-05/18/2022 |  | -1.4279 | 8.4249e-05 | -1.4282 | 7.4316e-05 | 280 | 2.82
| 362 | 362 | 21216 | 17214 | 1028 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/15-04/17/2022 | A |  | 0.008 |  |  |  |  | 8.4712e-12 | -5.9144e-09 | -0.0010779 | -1.4161 | 0.9996 |  |  | 1.2257e-21 | -9.4274e-18 | 2.9463e-14 | -4.8146e-11 | 4.3947e-08 | -2.0367e-05 | -1.4159 | 0.9805 | DF |  | 362 | 05/16-05/18/2022 |  | -1.4185 | 0.00010255 | -1.4188 | 8.4444e-05 | 280 | 2.46
| 363 | 363 | 21803 | 16961 | 1030 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/23-04/25/2022 | A |  | -0.035 |  |  |  |  | -1.5811e-11 | -1.161e-08 | -0.0010129 | -1.4136 | 0.9995 |  |  | 2.0337e-21 | -1.3887e-17 | 3.7214e-14 | -4.9961e-11 | 3.6801e-08 | -1.697e-05 | -1.4137 | 0.9812 | DF |  | 363 | 05/31-06/03/2022 |  | -1.4178 | 6.2369e-05 | -1.418 | 9.0558e-05 | 190 | 4.25
| 364 | 364 | 21246 | 17422 | 1029 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 04/23-04/25/2022 | B |  | -0.091 | [6 empty] | -0.0010012 | -1.4258 | 0.9998 |  |  | 3.2303e-21 | -2.0772e-17 | 5.186e-14 | -6.4334e-11 | 4.3947e-08 | -1.9323e-05 | -1.4258 | 0.9954 | DF |  | 364 | 11/07-11/09/2022 |  | -1.4323 | 0.00012604 | -1.4334 | 0.00013823 | 1150 | 6.46
| 365 | 365 | 21839 | 17506 | 1046 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/10-07/14/2022 | A |  | -0.119 |  | -1.1099e-19 | 6.2398e-16 | -1.2221e-12 | 9.7887e-10 | -3.0347e-07 | -0.0010467 | -1.449 | 0.9988 |  | 1.0534e-24 | -8.5971e-21 | 2.9089e-17 | -5.284e-14 | 5.4978e-11 | -3.1706e-08 | 9.0061e-06 | -1.4491 | 0.6739 | DF |  | 365 | 10/10-10/12/2022 |  | -1.4525 | 3.7012e-05 | -1.4545 | 6.2689e-05 | 1950 | 3.48
| 366 | 366 | 21910 | 17351 | 1045 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/10-07/14/2022 | B |  | -0.056 | [6 empty] | -0.001175 | -1.4016 | 0.9994 |  |  | 1.3338e-21 | -1.0032e-17 | 2.974e-14 | -4.3284e-11 | 3.1304e-08 | -9.3622e-06 | -1.4014 | 0.8954 | DF |  | 366 | 10/10-10/12/2022 |  | -1.4057 | 4.3052e-05 | -1.4068 | 5.7167e-05 | 1080 | 4.12
| 367 | 367 | 21267 | 17352 | 1044 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/26-06/30/2022 | A |  | -0.062 |  | 9.6525e-20 | -5.2946e-16 | 1.0356e-12 | -8.282e-10 | 2.3583e-07 | -0.0010664 | -1.3946 | 0.9997 |  |  | 2.5664e-21 | -1.7751e-17 | 4.8222e-14 | -6.4894e-11 | 4.4367e-08 | -1.4e-05 | -1.3945 | 0.936 | DF |  | 367 | 10/10-10/12/2022 |  | -1.4005 | 0.00019986 | -1.4014 | 0.00019297 | 910 | 5.87
| 368 | 368
| 369 | 369 | 22739 | 18991 | 1047 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/03-08/07/2022 | A |  | -0.05 | [6 empty] | -0.0010302 | -1.4349 | 0.9997 | -2.1321e-27 | 1.9434e-23 | -7.3657e-20 | 1.4992e-16 | -1.7668e-13 | 1.2132e-10 | -4.6269e-08 | 8.0588e-06 | -1.435 | 0.8446 | DF |  | 369 | 10/16-10/18/2023 |  | -1.4386 | 9.1169e-05 | -1.439 | 0.00011234 | 430 | 3.68
| 370 | 370 | 21832 | 17420 | 1049 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/03-08/07/2022 | B |  | -0.252 | [6 empty] | -0.0010224 | -1.3895 | 0.9998 |  |  | -5.8574e-22 | 4.1895e-18 | -1.184e-14 | 1.668e-11 | -1.2315e-08 | 4.5517e-06 | -1.3893 | 0.7572 | DF |  | 370 | 10/24-10/26/2022 |  | -1.3957 | 9.5169e-05 | -1.3975 | 7.2282e-05 | 1820 | 6.21
| 373 | 373 | 21972 | 17509 | 1048 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 08/24-08/28/2022 | A |  | -0.109 | [6 empty] | -0.0010239 | -1.3724 | 0.9996 |  |  | 2.2013e-21 | -1.4887e-17 | 3.9617e-14 | -5.2942e-11 | 3.7679e-08 | -1.2734e-05 | -1.3725 | 0.9372 | DF |  | 373 | 10/28-11/01/2022 |  | -1.3782 | 0.00015478 | -1.3786 | 0.00017327 | 440 | 5.74
| 374 | 374 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  | Sept 2023: MSC memory corrupted and returned to MBARI. MSC reflashed and sensors recalibrated (K0 and ISUS) | 08/24-08/28/2022 | B |  | 0.075 |  |  |  | -3.6897e-14 | 1.1546e-10 | -8.9795e-08 | -0.0010724 | -1.3658 | 0.9992 |  | -3.9205e-24 | 2.9063e-20 | -8.5316e-17 | 1.2645e-13 | -1.0131e-10 | 4.6886e-08 | -1.6132e-05 | -1.3658 | 0.9875 | DF |  | 374 | 10/02-10/03/2023 |  | -1.3728 | 6.1525e-05 | -1.3745 | 0.00010344 | 1620 | 7.07
| 375 | 375 | 20209 | 16411 | 1055 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/29-10/03/2022 | A |  | -0.054 |  |  |  |  | 2.6448e-13 | 2.4907e-08 | -0.0010756 | -1.394 | 0.9993 |  |  | 2.7561e-21 | -1.8954e-17 | 5.1057e-14 | -6.7932e-11 | 4.5984e-08 | -1.4458e-05 | -1.3939 | 0.867 | DF |  | 375 | 11/07-11/09/2022 |  | -1.3983 | 7.35e-05 | -1.3985 | 6.4545e-05 | 220 | 4.31
| 376 | 376 | 20293 | 16412 | 1056 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/29-10/03/2022 | B |  | -0.047 |  |  |  | 1.7028e-14 | -7.7543e-11 | 9.7649e-08 | -0.0010641 | -1.3927 | 0.9997 |  |  | 1.1026e-21 | -7.6405e-18 | 2.0456e-14 | -2.6135e-11 | 1.5683e-08 | -2.7762e-06 | -1.3927 | 0.9435 | DF |  | 376 | 11/07-11/09/2022 |  | -1.3966 | 9.8753e-05 | -1.3968 | 9.6762e-05 | 210 | 3.9
| 377 | 377 | 21535 | 17384 | 1066 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/03-11/07/2022 | B |  | -0.161 |  |  |  |  | 4.0976e-12 | 1.6621e-09 | -0.00093248 | -1.3009 | 0.9999 |  |  | 1.6351e-21 | -1.1344e-17 | 3.1096e-14 | -4.3127e-11 | 3.2118e-08 | -1.1075e-05 | -1.3008 | 0.9853 | DF |  | 377 | 12/12-12/14/2022 |  | -1.2976 | 4.4989e-05 | -1.2983 | 4.7832e-05 | 780 | -3.36
| 378 | 378 | 21939 | 17382 | 1068 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/11-11/15/2022 | A |  | -0.034 | [5 empty] | 4.8956e-09 | -0.00093326 | -1.2875 | 0.9997 |  |  | 1.3451e-22 | 1.6095e-19 | -3.5756e-15 | 8.9225e-12 | -8.9575e-09 | 4.0167e-06 | -1.2874 | 0.7799 | DF |  | 378 | 12/19-12/22/2022 |  | -1.2903 | 5.1839e-05 | -1.2905 | 5.5887e-05 | 210 | 2.79
| 379 | 379 | 21857 | 17385 | 1067 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/11-11/15/2022 | B |  | -0.009 | [5 empty] | 1.4839e-08 | -0.0010121 | -1.2802 | 0.9998 | 1.841e-27 | -1.662e-23 | 6.2868e-20 | -1.2956e-16 | 1.5811e-13 | -1.1553e-10 | 4.7638e-08 | -9.148e-06 | -1.2801 | 0.6282 | DF |  | 379 | 12/12-12/14/2022 |  | -1.2837 | 5.6347e-05 | -1.2841 | 4.6497e-05 | 390 | 3.49
| 380 | 380 | 22218 | 18483 | 1075 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/18-11/22/2022 | A |  | -0.267 |  |  |  |  | 2.2232e-11 | -3.6855e-08 | -0.0009435 | -1.2664 | 0.9977 |  |  | -4.6496e-22 | 3.0965e-18 | -8.4199e-15 | 1.2175e-11 | -1.1044e-08 | 7.8929e-06 | -1.2665 | 0.9477 | DF |  | 380 | 01/30-02/01/2023 |  | -1.2698 | 8.6996e-05 | -1.2698 | 8.1477e-05 | 60 | 3.43
| 381 | 381 | 22372 | 18529 | 1078 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 11/18-11/22/2022 | B |  | -0.089 |  |  |  |  | 2.0475e-11 | -2.1311e-08 | -0.00099986 | -1.2721 | 0.9991 |  |  | 2.5897e-21 | -1.8606e-17 | 5.2696e-14 | -7.4404e-11 | 5.3746e-08 | -1.7249e-05 | -1.272 | 0.874 | DF |  | 381 | 01/30-02/01/2023 |  | -1.276 | 5.6334e-05 | -1.2764 | 6.0117e-05 | 320 | 3.89
| 382 | 382 | 22912 | 18532 | 1079 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/07-12/12/2022 | A |  | 0.002 |  |  |  | -1.4054e-14 | 6.575e-11 | -6.5504e-08 | -0.0009517 | -1.2668 | 0.9996 |  |  | -7.7614e-22 | 6.0635e-18 | -1.9183e-14 | 3.1086e-11 | -2.6837e-08 | 1.0527e-05 | -1.2669 | 0.9674 | DF |  | 382 | 03/06-03/08/2023 |  | -1.2699 | 8.4037e-05 | -1.2717 | 0.00018069 | 1760 | 3.15
| 383 | 383 | 22715 | 18531 | 1080 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/07-12/12/2022 | B |  | -0.02 |  |  |  | -6.3573e-14 | 2.2765e-10 | -1.7493e-07 | -0.0010226 | -1.285 | 0.9986 |  |  | -9.2788e-23 | 9.1443e-19 | -3.4701e-15 | 6.9091e-12 | -8.5738e-09 | 6.3513e-06 | -1.2851 | 0.8723 | DF |  | 383 | 03/06-03/08/2023 |  | -1.2879 | 6.7313e-05 | -1.288 | 5.4382e-05 | 110 | 2.98
| 384 | 384 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/15-12/20/2022 | A |  | -0.059 |  |  |  | 7.3048e-15 | -4.6245e-12 | -1.6106e-09 | -0.00096415 | -1.27 | 0.9997 |  |  | 1.5008e-21 | -9.9264e-18 | 2.5448e-14 | -3.195e-11 | 1.9605e-08 | -3.8321e-06 | -1.27 | 0.9576 | DF |  | 384
| 385 | 385 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/31-01/04/2023 | A |  | -0.064 |  |  | -9.99e-17 | 3.9758e-13 | -4.4255e-10 | 1.2529e-07 | -0.00098106 | -1.2822 | 0.9995 |  |  | 6.5109e-22 | -4.4859e-18 | 1.2941e-14 | -2.0008e-11 | 1.6029e-08 | -3.8779e-06 | -1.2823 | 0.9807 | DF |  | 385
| 386 | 386 | 22658 | 18684 | 1082 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/31-01/04/2023 | B |  | -0.039 |  |  |  | -5.7569e-14 | 2.0849e-10 | -1.7901e-07 | -0.0010181 | -1.444 | 0.9998 |  |  | 9.5455e-23 | -7.7909e-19 | 2.5106e-15 | -3.7278e-12 | 1.3884e-09 | 1.4597e-06 | -1.444 | 0.9062 | DF |  | 386 | 04/04-04/06/2023 |  | -1.4472 | 7.527e-05 | -1.4475 | 0.00016891 | 330 | 3.21
| 387 | 387 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 02/16-02/21/2023 | A |  | -0.009 | -1.1199e-22 | 5.0428e-19 | -6.0788e-16 | -8.3775e-14 | 4.0273e-10 | -6.5673e-08 | -0.00094885 | -1.2831 | 0.9992 |  | -2.05e-24 | 1.7551e-20 | -5.9982e-17 | 1.0393e-13 | -9.6281e-11 | 4.6624e-08 | -1.0535e-05 | -1.2831 | 0.6446 | DF |  | 387
| 388 | 388 | 22958 | 18530 | 1086 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 02/16-02/21/2023 | B |  | -0.021 |  |  |  |  | 1.1681e-11 | -9.5383e-09 | -0.0010052 | -1.29 | 0.9995 |  |  | 2.651e-21 | -1.8104e-17 | 4.7901e-14 | -6.1306e-11 | 3.8282e-08 | -8.5857e-06 | -1.29 | 0.9796 | DF |  | 388 | 04/17-04/19/2023 |  | -1.293 | 9.5114e-05 | -1.2931 | 0.0001029 | 160 | 3.01
| 389 | 389 | 22278 | 18685 | 1087 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 03/02-03/06/2023 | A |  | -0.036 |  |  |  |  | 4.5468e-11 | -8.3836e-08 | -0.00096636 | -1.2874 | 0.9994 |  |  | 1.0846e-22 | -1.1152e-18 | 4.2667e-15 | -7.6723e-12 | 5.9829e-09 | -2.7461e-07 | -1.2874 | 0.956 | DF |  | 389 | 04/17-04/19/2023 |  | -1.2907 | 9.3329e-05 | -1.2908 | 8.1436e-05 | 160 | 3.28
| 390 | 390 | 22972 | 18686 | 1088 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 03/02-03/06/2023 | B |  | -0.133 |  |  | 1.8255e-17 | -5.4018e-14 | 1.0561e-11 | 3.5598e-08 | -0.00093133 | -1.2791 | 0.9996 |  |  | 1.1204e-21 | -7.8249e-18 | 2.1026e-14 | -2.7213e-11 | 1.7579e-08 | -4.4765e-06 | -1.2792 | 0.9433 | DF |  | 390 | 04/17-04/19/2023 |  | -1.2823 | 7.1041e-05 | -1.2825 | 5.8195e-05 | 200 | 3.19
| 391 | 391 | 22956 | 18689 | 1089 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 03/10-03/14/2023 | A |  | -0.028 | [6 empty] | -0.0010198 | -1.259 | 0.9993 | 1.7378e-27 | -1.564e-23 | 5.9037e-20 | -1.213e-16 | 1.4689e-13 | -1.0594e-10 | 4.4181e-08 | -9.385e-06 | -1.259 | 0.8722 | DF |  | 391 | 05/01-05/03/2023 |  | -1.2632 | 6.5505e-05 | -1.2636 | 7.5796e-05 | 430 | 4.2
| 392 | 392 | 22555 | 18480 | 1090 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 03/10-03/14/2023 | B |  | -0.025 | [6 empty] | -0.00092811 | -1.2866 | 0.9998 |  | -1.4031e-24 | 1.2953e-20 | -4.8375e-17 | 9.4216e-14 | -1.027e-10 | 6.2455e-08 | -1.875e-05 | -1.2862 | 0.9638 | DF |  | 392 | 05/01-05/03/2023 |  | -1.2946 | 7.629e-05 | -1.2948 | 0.00014131 | 240 | 8.04
| 393 | 393 | 21730 | 16960 | 1092 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 03/29-04/02/2023 | A |  | 0.024 |  |  | 2.2467e-17 | -8.5356e-14 | 8.8069e-11 | 7.1552e-09 | -0.0010101 | -1.2754 | 0.9998 |  |  | -2.0325e-22 | 1.4378e-19 | 4.3452e-15 | -1.5271e-11 | 2.1376e-08 | -1.2954e-05 | -1.2752 | 0.9873 | DF |  | 393 | 05/08-05/10/2023 |  | -1.2788 | 8.1149e-05 | -1.2792 | 0.000202 | 380 | 3.41
| 394 | 394 | 22916 | 18690 | 1093 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 03/29-04/02/2023 | B |  | 0.006 |  |  |  | 2.1355e-14 | -8.2412e-11 | 1.1305e-07 | -0.0010466 | -1.2863 | 0.9991 |  |  | 2.2076e-21 | -1.5228e-17 | 4.1227e-14 | -5.545e-11 | 3.8807e-08 | -1.3639e-05 | -1.2862 | 0.9172 | DF |  | 394 | 05/08-05/10/2023 |  | -1.2915 | 4.5186e-05 | -1.2936 | 4.239e-05 | 2100 | 5.18
| 395 | 395 | 23596 | 18482 | 1095 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 05/13-05/18/2023 | A |  | 0.051 |  |  |  | 4.3809e-14 | -1.7352e-10 | 1.8073e-07 | -0.0010614 | -1.2683 | 0.9997 | 7.9977e-27 | -7.0644e-23 | 2.5947e-19 | -5.124e-16 | 5.8643e-13 | -3.896e-10 | 1.4229e-07 | -2.6797e-05 | -1.2683 | 0.9713 | DF |  | 395 | 06/26-06/28/2023 |  | -1.2706 | 4.7559e-05 | -1.2707 | 4.9276e-05 | 70 | 2.29
| 396 | 396 | 22760 | 18691 | 1096 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 05/13-05/18/2023 | B |  | -0.004 | [6 empty] | -0.0011366 | -1.3866 | 0.9988 | 6.6003e-28 | -7.5806e-24 | 3.5689e-20 | -8.9796e-17 | 1.3128e-13 | -1.1336e-10 | 5.6291e-08 | -1.441e-05 | -1.3864 | 0.544 | DF |  | 396 | 06/26-06/28/2023 |  | -1.3887 | 5.9429e-05 | -1.3889 | 5.7276e-05 | 130 | 2.15
| 397 | 397 | 22537 | 18595 | 1097 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 05/24-05/27/2023 | A |  | -0.048 |  |  |  | 1.8597e-14 | -5.854e-11 | 3.9724e-08 | -0.00097665 | -1.3007 | 0.9999 |  |  | 1.7271e-21 | -1.1856e-17 | 3.1428e-14 | -4.049e-11 | 2.6271e-08 | -6.5432e-06 | -1.3008 | 0.9843 | DF |  | 397 | 06/26-06/28/2023 |  | -1.3045 | 4.3025e-05 | -1.3047 | 3.6938e-05 | 180 | 3.82
| 398 | 398 | 22953 | 18819 | 1100 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 05/24-05/27/2023 | B |  | 0.005 | [6 empty] | -0.00089572 | -1.2747 | 0.9999 |  | 1.0103e-25 | -1.4108e-21 | 6.7602e-18 | -1.5671e-14 | 1.9912e-11 | -1.4596e-08 | 6.1287e-06 | -1.2748 | 0.8374 | DF |  | 398 | 07/17-07/19/2023 |  | -1.2787 | 3.5387e-05 | -1.2788 | 3.5049e-05 | 140 | 3.97
| 399 | 399 | 22835 | 18484 | 1102 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/05-06/09/2023 | A |  | -0.093 | [6 empty] | -0.00093011 | -1.2666 | 0.9999 |  |  | -7.2388e-22 | 5.4236e-18 | -1.6175e-14 | 2.4207e-11 | -1.8856e-08 | 8.284e-06 | -1.2666 | 0.9857 | DF |  | 399 | 07/24-07/26/2023 |  | -1.272 | 4.2065e-05 | -1.2722 | 8.5186e-05 | 240 | 5.4
| 400 | 400 | 22565 | 18594 | 1103 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/14-06/18/2023 | B |  | 0.009 | [5 empty] | -2.8694e-09 | -0.00095761 | -1.2853 | 0.9999 |  |  | -4.0066e-23 | 2.0525e-19 | -6.9184e-16 | 1.8771e-12 | -2.7447e-09 | 3.3404e-06 | -1.2853 | 0.9888 | DF |  | 400 | 07/24-07/26/2023 |  | -1.2874 | 4.6475e-05 | -1.2876 | 6.6004e-05 | 250 | 2.06
| 401 | 401 | 22743 | 18821 | 1104 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/14-06/18/2023 | A |  | -0.082 |  |  |  | -1.1851e-14 | 3.2077e-11 | -2.2294e-08 | -0.00096867 | -1.2688 | 1 |  |  | -2.6338e-22 | 1.4033e-18 | -2.7429e-15 | 2.9447e-12 | -2.7976e-09 | 2.8477e-06 | -1.2688 | 0.9798 | DF |  | 401 | 07/24-07/26/2023 |  | -1.2726 | 4.1962e-05 | -1.2728 | 5.3115e-05 | 170 | 3.74
| 402 | 402 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/30-07/04/2023 | A |  | -0.05 |  |  |  | 4.9911e-15 | -8.2605e-12 | -9.5427e-09 | -0.00096304 | -1.2911 | 0.9998 |  |  | 2.7093e-23 | 6.1068e-21 | -3.3676e-16 | 1.0704e-13 | 1.12e-09 | 7.6887e-07 | -1.2911 | 0.9897 | DF |  | 402
| 403 | 403 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 06/30-07/04/2023 | B |  | -0.097 | [5 empty] | 5.5003e-09 | -0.00093648 | -1.2633 | 0.9999 |  |  | 3.1556e-22 | -1.4329e-18 | 1.4102e-15 | 2.2614e-12 | -5.1244e-09 | 4.0301e-06 | -1.2632 | 0.9713 | DF |  | 403
| 404 | 404 | 22790 | 19036 | 1105 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/22-07/25/2023 | A |  | -0.026 |  |  |  | 2.2843e-14 | -5.9162e-11 | 3.025e-08 | -0.0009453 | -1.2969 | 0.9999 |  |  |  | -8.3143e-19 | 4.3953e-15 | -8.9986e-12 | 8.5894e-09 | -4.5452e-07 | -1.2969 | 0.9888 | DF |  | 404 | 08/21-08/23/2023 |  | -1.3002 | 8.4874e-05 | -1.3004 | 5.9902e-05 | 130 | 3.31
| 405 | 405 | 22081 | 19024 | 1107 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 07/22-07/25/2023 | B |  | -0.132 | [5 empty] | -5.5101e-09 | -0.00093425 | -1.2603 | 0.9996 |  |  | -8.6587e-22 | 4.9042e-18 | -1.0317e-14 | 9.9271e-12 | -4.282e-09 | 2.4128e-06 | -1.2603 | 0.9785 | DF |  | 405 | 08/21-08/23/2023 |  | -1.264 | 6.5183e-05 | -1.2643 | 6.4077e-05 | 240 | 3.79
| 406 | 406 | 22243 | 18927 | 1124 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 09/26-09/30/2023 | A |  | 0.004 |  |  |  |  | -7.0852e-12 | 8.5538e-09 | -0.00096828 | -1.2709 | 0.9999 |  |  |  | 4.8218e-19 | -2.9126e-15 | 6.6335e-12 | -6.6743e-09 | 3.8539e-06 | -1.2709 | 0.9853 | DF |  | 406 | 11/13-11/15/2023 |  | -1.2746 | 4.301e-05 | -1.2757 | 5.1897e-05
| 407 | 407 | 22578 | 18600 | 1121 |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/03-10/07/2023 | A |  | -0.054 |  |  |  |  | 7.6164e-12 | -1.3722e-08 | -0.00096551 | -1.3067 | 0.9998 |  |  |  | -3.6879e-19 | 2.3432e-15 | -5.6588e-12 | 5.8517e-09 | 1.8063e-07 | -1.3067 | 0.9949 | DF |  | 407 | 11/14-11/15/2023 |  | -1.311 | 3.9908e-05 | -1.3115 | 5.8612e-05
| 408 | 408 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/03-10/07/2023 | B |  | -0.132 | [5 empty] | -9.9985e-09 | -0.0010351 | -1.3944 | 0.9981 |  | 2.5861e-24 | -2.1127e-20 | 6.9744e-17 | -1.1947e-13 | 1.1371e-10 | -5.9822e-08 | 1.5779e-05 | -1.3946 | 0.8124 | DF |  | 408
| 409 | 409 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/26-10/30/2023 | A |  | -0.205 | [6 empty] | -0.0010013 | -1.3529 | 0.9998 |  |  | -2.137e-21 | 1.6468e-17 | -5.0947e-14 | 8.0492e-11 | -6.7516e-08 | 2.6892e-05 | -1.3532 | 0.99 | DF |  | 409
| 410 | 410 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 10/26-10/30/2023 | B |  | -0.196 | [5 empty] | -3.91e-09 | -0.0010116 | -1.397 | 0.9998 |  | 8.0465e-25 | -6.9927e-21 | 2.4969e-17 | -4.7587e-14 | 5.2409e-11 | -3.2882e-08 | 9.8868e-06 | -1.397 | 0.9094 | DF |  | 410
| 411 | 411 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 12/05-12/08/2023 | A |  | 0.035 |  |  |  |  | -9.7001e-12 | 1.8522e-08 | -0.0010513 | -1.3978 | 0.9998 |  |  | -4.7037e-22 | 4.1868e-18 | -1.4991e-14 | 2.7979e-11 | -2.8828e-08 | 1.5892e-05 | -1.398 | 0.9958 | DF |  | 411
new Honeywell isfets | 412 | 412 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 01/12-01/16/2024 | A |  | -0.129 | [5 empty] | 5.2149e-10 | -0.00094179 | -1.5165 | 0.9995 | -3.4553e-27 | 3.0841e-23 | -1.1363e-19 | 2.2386e-16 | -2.5656e-13 | 1.7677e-10 | -7.5084e-08 | 1.8869e-05 | -1.5166 | 0.9578 | DF |  | 412
412-419 (8) blue heatshrink | 413 | 413 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 01/12-01/16/2024 | B |  | -0.108 | [5 empty] | -7.414e-09 | -0.0010261 | -1.4845 | 0.9995 | -4.6832e-27 | 4.2056e-23 | -1.5683e-19 | 3.1393e-16 | -3.6413e-13 | 2.4691e-10 | -9.4059e-08 | 1.7853e-05 | -1.4844 | 0.9393 | DF |  | 413
| 414 | 414 |  |  |  |  | sbe-3.5ksi | Ver6-PEEK, Pt |  |  | 01/19-01/22/2024 | A |  | -0.195 | [6 empty] | -0.00096342 | -1.4971 | 0.999 | -6.1286e-27 | 5.5291e-23 | -2.0752e-19 | 4.196e-16 | -4.9517e-13 | 3.4662e-10 | -1.4149e-07 | 3.1587e-05 | -1.497 | 0.9653 | DF |  | 414
| 415 | 415 | [8 empty] | 01/19-01/22/2024 | B |  | -0.203 | [5 empty] | -7.7395e-09 | -0.00090422 | -1.4803 | 0.9997 |  | 1.6677e-24 | -1.4689e-20 | 5.3324e-17 | -1.0304e-13 | 1.1406e-10 | -7.1635e-08 | 2.2948e-05 | -1.4802 | 0.9886 | DF |  | 415
| 416 | 416 | [8 empty] | 02/09-02/13/2024 | A |  | -0.032 | [6 empty] | -0.00095209 | -1.498 | 0.9994 |  | 2.9485e-24 | -2.3502e-20 | 7.6402e-17 | -1.3086e-13 | 1.2751e-10 | -7.1137e-08 | 2.0251e-05 | -1.4979 | 0.9416 | DF |  | 416
| 417 | 417 | [8 empty] | 02/09-02/13/2024 | B |  | -0.112 | [5 empty] | -4.9776e-09 | -0.00095736 | -1.4868 | 0.9998 |  |  | -3.0676e-22 | 3.0936e-18 | -1.2196e-14 | 2.4347e-11 | -2.5871e-08 | 1.3478e-05 | -1.487 | 0.9788 | DF |  | 417
| 418 | 418
| 419 | 419
CPF001 | 207 | 207 |  | 5952 | 1012 |  |  |  |  | pH sensor on CPF 001 | 07/02-07/06/2017 | C |  | 0.068 | [6 empty] | -0.00097247 | -1.422 | 0.9985 |  |  | -5.5486e-22 | 1.9826e-18 | 1.3306e-15 | -1.4474e-11 | 2.3673e-08 | -1.5426e-05 | -1.4219 | 0.9658 |  |  | 207 | 03/14-03/16/2022 |  | -1.4425 | 5.9325e-05 | -1.4426 | 6.1777e-05 | 110 | 20.53
GDF | 4 | 4 | 19719 | 9021 | 850 |  |  | GDF 04 |  | test gdf in MB non std cal | 12/10-12/12/2020 | [9 empty] | -0.00096427 | -1.42 | [6 empty] | -8.4518e-13 | 6.2525e-09 | -1.7213e-05 | -1.4 |  |  |  | 4 |  |  | -1.429 | 0.000116 | -1.4291 | 0.000101 | 94 | 9.049 | early GDF deployment. jp entered cals 12/082022
GDF | 9 | 9 | 17465 | 17424 | 1034 |  |  | GDF 09 |  | hand broached corners | 05/18-05/21/2022 |  |  | -0.033 |  |  |  | -3.955e-14 | 4.7625e-11 | -5.9057e-09 | -0.0010394 | -1.4008 | 0.9998 |  |  | -6.1682e-21 | 3.3549e-17 | -6.4166e-14 | 5.3274e-11 | -2.2685e-08 | 2.0248e-06 | -1.4007 | 0.9978 | GDF |  | 9 | 07/20-07/25/2022 |  | -1.4061 | 0.0001994 | -1.4062 | 0.00019955 | 200 | 5.28
GDF | 10 | 10 | 20109 | 12454 | 1032 |  |  | GDF 10 |  | hand broached corners | 05/18-05/21/2022 |  |  | -0.011 |  |  |  | -2.0518e-14 | 4.1977e-11 | -4.7477e-08 | -0.0010109 | -1.4195 | 0.9998 |  |  | -1.2826e-21 | 6.8517e-18 | -1.3348e-14 | 1.3487e-11 | -9.6292e-09 | 7.1396e-07 | -1.4194 | 0.9975 | GDF |  | 10 | 08/10-08/16/2022 |  | -1.4223 | 0.00011209 | -1.4225 | 0.00010864 | 190 | 2.82
GDF | 11 | 11 | 20148 | 15478 | 980 |  |  | GDF 11 |  |  | 10/19-10/25/2021 |  |  | 0.087 |  |  |  |  | -1.9109e-11 | -2.2004e-08 | -0.0010097 | -1.4316 | 0.9993 |  |  | -3.7167e-21 | 2.2249e-17 | -4.8113e-14 | 4.5213e-11 | -1.6083e-08 | -4.0003e-06 | -1.4315 | 0.9957 | GDF |  | 11 | 11/09-11/11/2021 |  | -1.4347 | 0.00018007 | -1.435 | 0.00017683 | 250 | 3.1
GDF | 12 | 12 | 20675 | 14604 | 979 |  |  | GDF 12 |  |  | 10/11-10/17/2021 |  |  | 0.039 |  |  |  | -3.8374e-14 | 9.5262e-11 | -1.0033e-07 | -0.0010142 | -1.4306 | 0.9995 |  |  | -4.4898e-21 | 2.9399e-17 | -7.0372e-14 | 7.5502e-11 | -3.5567e-08 | 2.3151e-06 | -1.4308 | 0.9947 | GDF |  | 12 | 11/09-11/11/2021 |  | -1.4314 | 0.00012886 | -1.4317 | 0.00014215 | 280 | 0.83
GDF | 13 | 13 | 21489 | 16670 | 998 |  |  | GDF 13 |  |  | 12/02-12/07/2021 |  |  | 0.01 |  |  |  | -2.3398e-14 | 5.3825e-11 | -7.2755e-08 | -0.0010283 | -1.402 | 0.9998 |  | -5.5215e-24 | 3.8915e-20 | -1.0608e-16 | 1.4128e-13 | -9.517e-11 | 2.936e-08 | -7.0986e-06 | -1.4017 | 0.9987 | GDF |  | 13 | 01/10-01/12/2022 |  | -1.4029 | 8.4123e-05 | -1.4033 | 8.0137e-05 | 390 | 0.87
GDF | 14 | 14 | 21141 | 16668 | 997 |  |  | GDF 14 |  |  | 12/10-12/15/2021 |  |  | -0.024 |  |  |  |  | -1.501e-11 | -3.6025e-08 | -0.0010901 | -1.4149 | 0.9997 |  |  | 3.9658e-22 | -8.7455e-19 | -1.9394e-15 | 6.2773e-12 | -2.6588e-09 | -5.1365e-06 | -1.4148 | 0.9915 | GDF |  | 14 | 01/10-01/13/2022 |  | -1.4197 | 0.00011783 | -1.4203 | 0.00011463 | 640 | 4.78
GDF | 15 | 15 | [5 empty] | GDF 15 |  |  | 08/23-08/28/2022 | A |  | -0.018 |  |  |  |  | 8.0985e-12 | -5.4376e-08 | -0.0016933 | -1.0447 | 0.9998 |  |  | -9.3838e-22 | 6.2604e-18 | -1.4765e-14 | 1.4e-11 | -2.1553e-09 | -4.4706e-06 | -1.0445 | 0.9275 | GDF |  | 15
GDF | 16 | 16 | 21844 | 17171 | 1019 |  |  | GDF 16 |  |  | 04/06-04/09/2022 | A |  | -0.33 |  |  |  |  | 6.8679e-12 | -5.9782e-08 | -0.0010498 | -1.3724 | 0.9998 |  |  | -2.2608e-21 | 1.3778e-17 | -3.0494e-14 | 2.9983e-11 | -1.1904e-08 | -2.6389e-06 | -1.3723 | 0.9928 | GDF |  | 16 | 04/25-04/27/2022 |  | -1.3834 | 0.00038109 | -1.3849 | 0.00040847 | 1460 | 11.01
GDF | 17 | 17 | 21206 | 17193 | 1022 |  |  | GDF 17 |  |  | 04/16-04/19/2022 | A |  | -0.191 |  |  | 9.3471e-17 | -3.9362e-13 | 4.8126e-10 | -1.9757e-07 | -0.0010086 | -1.3754 | 0.9997 |  |  | -2.4427e-21 | 1.3145e-17 | -2.495e-14 | 2.1533e-11 | -1.0645e-08 | -8.5368e-07 | -1.3754 | 0.9984 | GDF |  | 17 | 05/03-05/04/2022 |  | -1.3818 | 6.0515e-05 | -1.3827 | 7.7683e-05 | 930 | 6.41
GDF | 18 | 18 | 21811 | 17177 | 1021 |  |  | GDF 18 |  |  | 04/06-04/09/2022 | B |  | -0.527 |  |  |  | -4.6645e-14 | 1.1009e-10 | -1.1098e-07 | -0.0010407 | -1.3911 | 0.9998 |  |  | -2.8235e-21 | 1.5201e-17 | -2.7316e-14 | 1.8296e-11 | -3.7057e-09 | -3.4161e-06 | -1.391 | 0.9973 | GDF |  | 18 | 05/02-05/04/2022 |  | -1.4005 | 6.5575e-05 | -1.4027 | 4.0872e-05 | 2230 | 9.36
GDF | 19 | 19 | 21984 | 17174 | 1018 |  |  | GDF 19 |  |  | 03/24-03/27/2022 | A |  | -0.116 |  |  |  |  | -1.3083e-11 | -3.7311e-08 | -0.00097567 | -1.4169 | 0.9997 |  |  | -5.8354e-22 | 2.499e-18 | -1.9538e-15 | -2.8186e-12 | 3.998e-09 | -5.4226e-06 | -1.4167 | 0.9959 | GDF |  | 19 | 04/25-04/27/2022 |  | -1.4251 | 8.2989e-05 | -1.4255 | 7.2773e-05 | 380 | 8.24
GDF | 20 | 20 | 20040 | 14606 | 1024 |  |  | GDF 20 |  |  | 04/16-04/19/2022 | B |  | -0.349 |  |  | 6.1146e-17 | -2.8307e-13 | 3.9712e-10 | -1.8409e-07 | -0.001026 | -1.399 | 0.9994 |  |  | -5.206e-22 | 1.5589e-18 | 4.3554e-16 | -3.4128e-12 | 1.4609e-09 | -4.1413e-06 | -1.3989 | 0.9963 | GDF |  | 20 | 05/02-05/04/2022 |  | -1.4054 | 5.9719e-05 | -1.4057 | 7.1754e-05 | 330 | 6.33
GDF | 21 | 21 | 21142 | 17176 | 1020 |  |  | GDF 21 |  |  | 03/17-03/20/2022 | B |  | -0.021 |  |  |  |  | -2.2669e-11 | -1.9772e-08 | -0.0010128 | -1.4259 | 0.9998 |  |  | 2.7423e-21 | -1.5871e-17 | 3.4166e-14 | -3.3735e-11 | 1.6129e-08 | -7.6692e-06 | -1.4257 | 0.9954 | GDF |  | 21 | 04/25-04/27/2022 |  | -1.4307 | 0.00013175 | -1.4317 | 0.00012796 | 900 | 4.84
GDF | 22 | 22 | 20572 | 13779 | 1031 |  |  | GDF 22 |  |  | 05/20-05/25/2022 |  |  | -0.102 |  |  |  | -4.3847e-14 | 1.123e-10 | -1.0295e-07 | -0.0010637 | -1.4082 | 0.9994 |  |  | -8.0527e-22 | 3.9512e-18 | -5.6232e-15 | 2.3062e-12 | -7.1322e-10 | -2.9938e-06 | -1.408 | 0.9959 | GDF |  | 22 | 06/20-06/22/2022 |  | -1.4153 | 0.00011816 | -1.4157 | 0.00011437 | 360 | 7.14
GDF | 23 | 23 | [5 empty] | GDF 23 |  |  | 07/09-07/14/2022 | B |  | -0.134 |  |  |  |  | -4.0724e-11 | 1.425e-08 | -0.0011065 | -1.4016 | 0.9998 |  |  | 6.3914e-22 | -5.2563e-18 | 1.5566e-14 | -1.8147e-11 | 5.3739e-09 | -1.2411e-06 | -1.4015 | 0.9918 | GDF |  | 23
GDF | 24 | 24 | 21302 | 17432 | 1036 |  |  | GDF 24 |  |  | 06/29-07/04/2022 |  |  | -0.167 |  |  |  | -4.6352e-14 | 1.2448e-10 | -9.4551e-08 | -0.0010061 | -1.3749 | 0.9993 |  |  | -1.1764e-21 | 7.6521e-18 | -1.7372e-14 | 1.7395e-11 | -8.094e-09 | -3.392e-06 | -1.3748 | 0.9947 | GDF |  | 24 | 08/22-08/23/2022 |  | -1.3798 | 6.517e-05 | -1.3801 | 6.4891e-05 | 290 | 4.92
GDF | 25 | 25 | 21546 | 16962 | 1051 |  |  | GDF 25 |  |  | 09/03-09/08/2022 |  |  | 0.136 |  |  | 3.3617e-17 | -1.4206e-13 | 1.668e-10 | -6.2745e-08 | -0.00096869 | -1.4169 | 0.9998 |  | 2.6985e-24 | -1.8921e-20 | 4.993e-17 | -6.1446e-14 | 3.7648e-11 | -1.3257e-08 | -1.4818e-06 | -1.4168 | 0.9975 | GDF |  | 25 | 10/28-11/01/2022 |  | -1.4178 | 0.0001741 | -1.4179 | 0.00015921 | 60 | 0.93
GDF | 26 | 26 | 21285 | 17463 | 1033 |  |  | GDF 26 |  |  | 06/23-06/28/2022 | A |  | -0.085 |  |  |  | -2.9205e-14 | 8.0865e-11 | -9.2788e-08 | -0.00099699 | -1.3832 | 0.9995 |  |  | 1.6499e-21 | -1.1511e-17 | 3.0283e-14 | -3.545e-11 | 1.7538e-08 | -6.1457e-06 | -1.3829 | 0.9876 | GDF |  | 26 | 07/20-07/25/2022 |  | -1.3888 | 0.0001882 | -1.389 | 0.00018724 | 160 | 5.65
GDF | 27 | 27 | 21298 | 17431 | 1035 |  |  | GDF 27 |  |  | 06/12-06/17/2022 | A |  | -0.13 |  |  |  | -5.2156e-14 | 1.624e-10 | -1.6441e-07 | -0.00097313 | -1.3958 | 0.9995 |  |  | -3.8233e-21 | 2.2398e-17 | -4.8162e-14 | 4.7788e-11 | -2.2482e-08 | 1.1273e-06 | -1.3959 | 0.99 | GDF |  | 27 | 07/20-07/25/2022 |  | -1.402 | 0.00015397 | -1.4022 | 0.00015346 | 170 | 6.26
GDF | 28 | 28 | 21951 | 17433 | 1037 |  |  | GDF 28 |  |  | 07/09-07/14/2022 | A |  | -0.086 |  |  |  | -4.0694e-14 | 1.1363e-10 | -1.161e-07 | -0.00098619 | -1.3887 | 0.9994 |  |  | -1.8362e-22 | 1.757e-18 | -5.3158e-15 | 7.7556e-12 | -5.7439e-09 | -1.9248e-06 | -1.3888 | 0.9919 | GDF |  | 28 | 08/10-08/16/2022 |  | -1.3952 | 0.00016323 | -1.3953 | 0.00016073 | 70 | 6.46
GDF | 29 | 29 | 21519 | 17592 | 1062 |  |  | GDF 29 |  |  | 11/05-11/10/2022 |  |  | -0.039 |  |  |  | -5.6167e-14 | 1.8614e-10 | -2.0622e-07 | -0.00087272 | -1.397 | 0.9995 |  |  | 2.7233e-21 | -1.8059e-17 | 4.6733e-14 | -5.8081e-11 | 3.4545e-08 | -1.0378e-05 | -1.3967 | 0.9425 | GDF |  | 29 | 12/05-12/08/2022 |  | -1.4006 | 0.00013008 | -1.4014 | 6.6983e-05 | 770 | 3.56
GDF | 30 | 30 | 21441 | 17350 | 1061 |  |  | GDF 30 |  |  | 11/05-11/10/2022 |  |  | -0.019 |  |  |  |  | 3.0527e-12 | -5.3858e-08 | -0.00083442 | -1.2765 | 0.9997 |  | -3.9719e-24 | 2.7985e-20 | -7.5409e-17 | 9.5128e-14 | -5.3598e-11 | 1.0647e-08 | -3.7919e-06 | -1.2764 | 0.9777 | GDF |  | 30 | 12/05-12/07/2022 |  | -1.2825 | 9.6274e-05 | -1.2827 | 6.5121e-05 | 270 | 5.99
GDF | 31 | 31 | 21479 | 17423 | 1038 |  |  | GDF 31 |  | died cycle 2, built by M Bif | 08/20-08/25/2022 |  |  | -0.1 |  |  | 5.7418e-17 | -2.7677e-13 | 4.0435e-10 | -1.9867e-07 | -0.0010599 | -1.3712 | 0.9998 |  |  | 1.0781e-21 | -7.6449e-18 | 2.0557e-14 | -2.5018e-11 | 1.4478e-08 | -8.5004e-06 | -1.3709 | 0.9977 | GDF |  | 31 | 09/19-09/22/2022 |  | -1.3801 | 0.00032129 | -1.3802 | 0.00032189 | 90 | 8.92
GDF | 32 | 32 | 21485 | 17683 | 1039 |  |  | GDF 32 |  | died cycle 2, built by Taylor W | 08/20-08/25/2022 |  |  | -0.123 |  |  |  |  | -1.486e-11 | 5.6168e-09 | -0.0010041 | -1.3587 | 0.9995 |  |  | -4.0772e-22 | -1.2377e-18 | 1.664e-14 | -4.2737e-11 | 4.294e-08 | -1.6619e-05 | -1.3581 | 0.9406 | GDF |  | 32 | 09/16-09/21/2022 |  | -1.3661 | 0.00040014 | -1.3662 | 0.00039247 | 150 | 7.34
GDF | 33 | 33 | spray glider deployment Mont Bay |  |  |  |  | GDF 33 |  | oil filled GDF | 04/12-04/17/2024 |  |  | -0.064 | [5 empty] | -2.4858e-08 | -0.00080319 | -1.2734 | 0.9993 |  |  |  | -1.0699e-18 | 6.0341e-15 | -1.194e-11 | 1.0184e-08 | -2.7915e-06 | -1.2733 | 0.9471 | GDF |  | 33
GDF | 34 | 34 | Spray Glider deployment Mont Bay | [7 empty] | 12/21-12/27/2024 |  |  | -0.078 | [5 empty] | -5.6905e-09 | -0.0013546 | -0.14509 | 0.9998 | [7 empty] | 2.5631e-06 | -0.14511 | 0.9931 | GDF
AE GDF | 2000 | 2000 | 22554 | 18688 | 1091 |  |  | AE GDF |  | Ti no shoulder coverplate | 04/13-04/18/2023 |  |  | 0.108 |  | 2.5497e-20 | -1.517e-16 | 3.1987e-13 | -2.817e-10 | 7.8439e-08 | -0.00090111 | -1.2805 | 0.9999 |  |  | -4.6453e-22 | 1.8916e-18 | -9.6698e-16 | -3.9069e-12 | 5.3522e-09 | -7.0192e-06 | -1.2805 | 0.9994 | GDF |  | 2000 | 05/08-05/10/2023 |  | -1.2841 | 0.0001699 | -1.2842 | 0.00013293 | 150 | 3.6 | Ti no shoulder coverplate
AE GDF | 2001 | 2001 | 22725 | 18550 | 1077 |  |  | AE GDF |  |  | 12/29-01/03/2023 |  |  | 0.023 |  |  | 3.5782e-17 | -1.7568e-13 | 2.78e-10 | -2.0788e-07 | -0.00087152 | -1.3143 | 0.9992 |  |  | -6.3031e-22 | 4.7783e-18 | -1.2804e-14 | 1.4753e-11 | -4.3634e-09 | -1.0309e-05 | -1.3142 | 0.9991 | GDF |  | 2001 | 05/01-05/03/2023 |  | -1.3177 | 8.5987e-05 | -1.3186 | 0.00025765 | 870 | 3.39
AE GDF | 2002 | 2002 | 22219 | 18820 | 1099 |  |  | AE GDF |  |  | 05/25-05/29/2023 |  |  | -0.053 |  |  |  | -4.9052e-14 | 1.744e-10 | -1.9861e-07 | -0.0010663 | -1.3968 | 0.9998 |  | -2.4422e-24 | 1.7791e-20 | -5.0044e-17 | 6.973e-14 | -5.4105e-11 | 2.9738e-08 | -1.907e-05 | -1.3963 | 0.9991 | GDF |  | 2002 | 07/17-07/19/2023 |  | -1.4003 | 8.1086e-05 | -1.4004 | 6.7971e-05 | 100 | 3.5
AE GDF | 2003 | 2003 | 22720 | 18683 | 1094 |  |  | AE GDF |  | Ti no shoulder coverplate | 04/01-04/10/2023 |  |  | -0.041 |  |  | -1.2678e-17 | 8.5319e-14 | -1.9629e-10 | 1.4669e-07 | -0.0011652 | -1.4391 | 0.9998 |  |  | 3.5661e-22 | -3.2246e-18 | 1.2156e-14 | -2.3675e-11 | 2.5131e-08 | -1.8253e-05 | -1.4388 | 0.998 | GDF |  | 2003 | 08/21-08/23/2023 |  | -1.4461 | 7.195e-05 | -1.4463 | 8.9139e-05 | 220 | 6.95 | Ti no shoulder coverplate
AE GDF | 2004 | 2004 | 22615 | 18986 | 1106 |  |  | AE GDF |  | Ti no shoulder coverplate | 07/09-07/14/2023 |  |  | 0.086 |  |  | -4.0534e-17 | 1.9439e-13 | -3.1192e-10 | 1.6586e-07 | -0.0010921 | -1.3747 | 0.9999 |  |  | -5.1819e-22 | 3.7593e-18 | -9.8592e-15 | 9.3143e-12 | 4.0055e-09 | -1.2907e-05 | -1.3744 | 0.9866 | GDF |  | 2004 | 08/22-08/24/2023 |  | -1.3766 | 6.47e-05 | -1.3767 | 6.2926e-05 | 90 | 1.9 | Ti no shoulder coverplate
AE GDF | 2005 | 2005 | 22127 | 18987 | 1098 |  |  | AE GDF |  |  | 05/25-05/29/2023 |  |  | -0.07 |  |  |  | -3.4796e-14 | 1.1662e-10 | -1.5759e-07 | -0.00095018 | -1.4023 | 0.9999 |  | -1.8396e-24 | 1.2876e-20 | -3.5105e-17 | 4.8483e-14 | -3.8886e-11 | 2.3107e-08 | -1.5881e-05 | -1.4019 | 0.9992 | GDF |  | 2005 | 07/17-07/19/2023 |  | -1.4044 | 7.635e-05 | -1.4044 | 7.3055e-05 | 60 | 2.06
AE GDF | 2007 | 2007 | 22127 | 18601 | 1101 |  |  | AE GDF |  | Ti no shoulder coverplate | 04/13-04/18/2023 |  |  | 0.018 |  |  | -2.3686e-17 | 1.2987e-13 | -2.3295e-10 | 1.0562e-07 | -0.00095962 | -1.2856 | 0.9998 |  | -9.5297e-25 | 8.1237e-21 | -2.7696e-17 | 4.8654e-14 | -4.9049e-11 | 3.3077e-08 | -1.7689e-05 | -1.2855 | 0.9986 | GDF |  | 2007 | 07/17-07/19/2023 |  | -1.2882 | 9.6068e-05 | -1.2882 | 9.0732e-05 | 30 | 2.6 | Ti no shoulder coverplate
AE GDF | 2008 | 2008 | 21986 | 17508 | 1050 |  |  | AE GDF |  | Ti no shoulder coverplate, Sept 2023: MSC 1050 memory corrupted and returned to MBARI. MSC reflashed and sensors recalibrated (K0 and ISUS) | 06/08-06/13/2023 |  |  | 0.139 |  |  | -3.4843e-17 | 1.5554e-13 | -2.2739e-10 | 9.2849e-08 | -0.0010436 | -1.4458 | 0.9996 |  |  | -6.5553e-22 | 4.9411e-18 | -1.3878e-14 | 1.5967e-11 | -9.0064e-10 | -1.28e-05 | -1.4458 | 0.9967 | GDF |  | 2008 | 10/16-10/18/2023 |  | -1.4429 | 9.5109e-05 | -1.443 | 9.4841e-05 | 60 | -2.89 | Ti no shoulder coverplate
AE GDF | 2009 | 2009 | 21648 | 18549 | 1076 |  |  | AE GDF |  |  | 12/11-12/16/2022 |  |  | -0.116 |  |  |  | -3.5449e-14 | 1.189e-10 | -1.3113e-07 | -0.00076933 | -1.3217 | 0.9997 |  |  | 2.6067e-21 | -1.6142e-17 | 3.8949e-14 | -4.7831e-11 | 3.4334e-08 | -1.8344e-05 | -1.3215 | 0.9953 | GDF |  | 2009 | 05/01-05/03/2023 |  | -1.3258 | 0.00017116 | -1.3265 | 0.00034575 | 740 | 4.11
AE GDF | 2010 | 2010 | [5 empty] | AE GDF
AE GDF | 2011 | 2011 | 22344 | 18923 | 1112 |  |  | AE GDF |  |  | 08/04-08/09/2023 |  |  | 0.006 |  |  | 3.3352e-17 | -1.0553e-13 | 5.2854e-11 | 1.7702e-08 | -0.00085691 | -1.2845 | 0.9998 |  | 3.2194e-24 | -2.3745e-20 | 6.7726e-17 | -9.4214e-14 | 6.7692e-11 | -2.4108e-08 | -1.6098e-06 | -1.2846 | 0.9983 | GDF |  | 2011 | 10/16-10/18/2023 |  | -1.2878 | 0.0001017 | -1.2878 | 0.00010157 | 30 | 3.26
AE GDF | 2012 | 2012 | 22893 | 19031 | 1118 |  |  | AE GDF |  | Ti no shoulder coverplate | 08/04-08/09/2023 |  |  | -0.057 |  |  | -6.695e-17 | 2.7864e-13 | -3.3876e-10 | 5.3066e-08 | -0.00095491 | -1.2966 | 0.9998 |  | -4.1928e-24 | 3.2157e-20 | -9.7091e-17 | 1.455e-13 | -1.1183e-10 | 4.1959e-08 | -8.2233e-06 | -1.2963 | 0.8139 | GDF |  | 2012 | 11/06-11/08/2023 |  | -1.301 | 8.416e-05 | -1.3011 | 7.9059e-05 | 50 | 4.44
AE GDF | 2013 | 2013 | 22732 | 18822 | 1113 |  |  | AE GDF |  |  | 08/11-08/16/2023 |  |  | 0.068 |  |  |  | -2.1706e-14 | 6.4973e-11 | -7.3654e-08 | -0.00086554 | -1.3336 | 0.9997 |  |  | -1.049e-21 | 6.8772e-18 | -1.6465e-14 | 1.7557e-11 | -7.2414e-09 | -1.6126e-06 | -1.3336 | 0.9803 | GDF |  | 2013 | 10/23-10/25/2023 |  | -1.3358 | 7.1805e-05 | -1.3359 | 5.6508e-05 | 150 | 2.18
AE GDF | 2014 | 2014 | 22044 | 18995 | 1114 |  |  | AE GDF |  |  | 08/11-08/16/2023 |  |  | 0.126 |  |  | 2.5335e-17 | -1.584e-13 | 2.8946e-10 | -1.5374e-07 | -0.0010087 | -1.2814 | 0.9989 | 4.8301e-27 | -4.1187e-23 | 1.4057e-19 | -2.4476e-16 | 2.3104e-13 | -1.1748e-10 | 2.9994e-08 | -5.4849e-06 | -1.2814 | 0.9876 | GDF |  | 2014 | 10/23-10/25/2023 |  | -1.2834 | 0.00010722 | -1.2836 | 6.3359e-05 | 200 | 1.98
AE GDF | 2015 | 2015 | 22551 | 19032 | 1115 |  |  | AE GDF |  |  | 08/26-08/31/2023 |  |  | 0.046 |  | 1.3847e-19 | -6.5087e-16 | 9.8393e-13 | -4.9989e-10 | 2.8182e-08 | -0.0010182 | -1.3046 | 0.9991 | 7.9696e-27 | -6.2744e-23 | 1.9439e-19 | -3.0074e-16 | 2.4667e-13 | -1.0913e-10 | 2.6218e-08 | -6.0564e-06 | -1.3045 | 0.9917 | GDF |  | 2015 | 10/23-10/25/2023 |  | -1.3084 | 5.366e-05 | -1.3087 | 8.7432e-05 | 210 | 3.81
AE GDF | 2016 | 2016 | 22279 | 18994 | 1116 |  |  | AE GDF |  |  | 09/03-09/08/2023 |  |  | 0.086 |  |  | -1.464e-17 | 4.5412e-14 | -1.326e-11 | -6.0985e-08 | -0.00091462 | -1.3135 | 0.9986 |  | 1.6275e-24 | -1.1797e-20 | 3.2433e-17 | -4.1068e-14 | 2.3134e-11 | -6.1119e-09 | 8.7758e-07 | -1.3135 | 0.9858 | GDF |  | 2016 | 10/23-10/25/2023 |  | -1.3177 | 6.7355e-05 | -1.318 | 6.3061e-05 | 290 | 4.26
AE GDF | 2017 | 2017 | 22849 | 19033 | 1117 |  |  | AE GDF |  |  | 09/16-09/22/2023 |  |  | 0.085 |  | 2.2926e-20 | -7.459e-17 | 2.9474e-14 | 5.0656e-11 | -3.2262e-08 | -0.00088155 | -1.3024 | 0.9991 |  |  |  | -1.925e-18 | 7.5091e-15 | -5.4508e-12 | -5.0543e-09 | -2.5998e-06 | -1.3024 | 0.9975 | GDF |  | 2017 | 11/06-11/09/2023 |  | -1.3058 | 9.0542e-05 | -1.3062 | 9.4777e-05 | 390 | 3.41
AE GDF | 2018 | 2018 | 22274 | 18924 | 1120 |  |  | AE GDF |  |  | 09/16-09/22/2023 |  |  | -0.138 | [5 empty] | -5.1686e-08 | -0.00086409 | -1.2821 | 0.9997 |  |  |  | 6.4803e-19 | -3.0113e-15 | 4.7583e-12 | -1.2585e-09 | -4.6969e-06 | -1.282 | 0.9927 | GDF |  | 2018 | 11/06-11/09/2023 |  | -1.2863 | 7.8584e-05 | -1.2865 | 7.4951e-05 | 190 | 4.22
AE GDF | 2019 | 2019 | 22923 | 18993 | 1126 |  |  | AE GDF |  |  | 10/29-11/03/2023 |  |  | -0.089 | [5 empty] | -4.1362e-08 | -0.00084112 | -1.3271 | 0.9997 |  |  | -4.233e-21 | 2.757e-17 | -6.769e-14 | 7.552e-11 | -3.3044e-08 | -2.4311e-06 | -1.3269 | 0.9702 | GDF |  | 2019 | 11/20-11/22/2023 |  | -1.3314 | 4.1898e-05 | -1.3315 | 4.3186e-05 | 90 | 4.29
AE GDF | 2020 | 2020 | 22287 | 18599 | 1123 |  |  | AE GDF |  | Ti no shoulder coverplate | 10/06-10/11/2023 |  |  | -0.047 |  |  |  | 3.2193e-14 | -9.8368e-11 | 1.6933e-08 | -0.00087245 | -1.3349 | 0.9999 |  |  |  | -1.4624e-18 | 7.4901e-15 | -1.3091e-11 | 9.7954e-09 | -3.8312e-06 | -1.3347 | 0.9516 | GDF |  | 2020 | 11/13-11/15/2023 |  | -1.3374 | 6.6832e-05 | -1.3375 | 5.6659e-05 | 40 | 2.58
AE GDF | 2021 | 2021 | [5 empty] | AE GDF
AE GDF | 2022 | 2022 | 22438 | 18597 | 1122 |  |  | AE GDF |  | Ti no shoulder coverplate | 10/06-10/11/2023 |  |  | -0.045 |  |  |  | 1.0294e-14 | -4.4865e-11 | 2.9118e-08 | -0.00087181 | -1.3088 | 0.9997 |  |  |  | -1.715e-18 | 1.067e-14 | -2.32e-11 | 2.165e-08 | -1.1486e-05 | -1.3085 | 0.9963 | GDF |  | 2022 | 11/13-11/15/2023 |  | -1.3136 | 4.329e-05 | -1.3137 | 4.0127e-05 | 90 | 4.81
AE GDF | 2023 | 2023 | 22769 | 18928 | 1130 |  |  | AE GDF |  |  | 12/06-12/11/2023 |  |  | 0.044 |  |  |  | -2.358e-14 | 8.2312e-11 | -9.8557e-08 | -0.00086072 | -1.2965 | 1 |  | -1.6981e-24 | 1.3775e-20 | -4.3328e-17 | 6.5434e-14 | -4.5766e-11 | 1.0448e-08 | -3.8862e-06 | -1.2965 | 0.998 | AE GDF |  | 2023 | 01/08-01/10/2024 |  | -1.2985 | 6.5169e-05 | -1.2987 | 7.1085e-05 | 170 | 2.02
AE GDF | 2024 | 2024 | [5 empty] | AE GDF
AE GDF | 2025 | 2025 | 22687 | 19034 | 1127 |  |  | AE GDF |  |  | 10/23-10/26/2023 |  |  | 0.089 |  | -7.5587e-20 | 4.1051e-16 | -7.7603e-13 | 5.9246e-10 | -1.8491e-07 | -0.00087703 | -1.3318 | 0.9997 |  | -1.6344e-24 | 1.3264e-20 | -4.1832e-17 | 6.2705e-14 | -4.1404e-11 | 4.987e-09 | 1.4044e-06 | -1.3317 | 0.9934 | GDF |  | 2025 | 11/20-11/22/2023 |  | -1.3348 | 7.492e-05 | -1.3348 | 6.6829e-05 | 70 | 2.99
AE GDF | 2026 | 2026 |  | 20216 | 1132 |  |  | AE GDF |  |  | 03/21-03/26/2024 |  |  | -0.049 |  |  |  |  | -3.1647e-11 | 1.799e-08 | -0.00091835 | -1.3008 | 0.9997 |  |  | 1.1124e-21 | -1.0715e-17 | 3.5844e-14 | -4.9964e-11 | 2.6371e-08 | -9.013e-06 | -1.3007 | 0.9971 | AE GDF |  | 2026 | 05/28-05/30/2024 |  | -1.3058 | 6.5735e-05 | -1.3058 | 6.4932e-05 | 30 | 4.97
AE GDF | 2027 | 2027 | 22175 | 19037 | 1131 |  |  | AE GDF |  |  | 12/14-12/19/2023 |  |  | 0.011 |  |  | 3.8933e-17 | -1.8519e-13 | 2.6777e-10 | -1.2368e-07 | -0.00092633 | -1.3179 | 0.9999 |  |  | 1.9379e-21 | -1.1455e-17 | 2.4198e-14 | -1.9917e-11 | 4.3464e-09 | -5.3038e-06 | -1.3179 | 0.9973 | AE GDF |  | 2027 | 01/08-01/10/2024 |  | -1.3216 | 8.2047e-05 | -1.3217 | 8.7409e-05 | 60 | 3.71
AE GDF | 2028 | 2028 | 22474 | 19035 | 1125 |  |  | AE GDF |  |  | 10/28-11/02/2023 |  |  | -0.068 |  |  |  | -7.9928e-15 | 2.603e-12 | -2.5225e-08 | -0.00086532 | -1.2905 | 0.9996 |  |  | -4.5145e-22 | 1.1816e-18 | 1.8458e-15 | -6.764e-12 | 5.1113e-09 | -6.9682e-06 | -1.2903 | 0.9978 | GDF |  | 2028 | 11/20-11/22/2023 |  | -1.295 | 8.673e-05 | -1.2951 | 7.6947e-05 | 70 | 4.51
AE GDF | 2029 | 2029 | [5 empty] | AE GDF
AE GDF | 2030 | 2030 | 22653 | 18992 | 1128 |  |  | AE GDF |  |  | 10/29-11/03/2023 |  |  | -0.068 |  | 8.8456e-20 | -4.1159e-16 | 6.1899e-13 | -3.2668e-10 | 5.5277e-09 | -0.00088102 | -1.3101 | 0.9993 |  | 5.5161e-24 | -4.3565e-20 | 1.3269e-16 | -1.9388e-13 | 1.3769e-10 | -4.3336e-08 | 2.1216e-06 | -1.3102 | 0.9953 | GDF |  | 2030 | 11/20-11/22/2023 |  | -1.3146 | 6.2575e-05 | -1.3147 | 6.1673e-05 | 60 | 4.5
AE GDF | 2031 | 2031 | [5 empty] | AE GDF
AE GDF | 2032 | 2032 | [5 empty] | AE GDF
AE GDF | 2033 | 2033 | [5 empty] | AE GDF
AE GDF | 2034 | 2034 |  | 18598 | 1142 |  |  | AE GDF |  |  | 04/23-04/29/2024 |  |  | -0.094 |  |  |  | -1.7887e-14 | 5.7377e-11 | -6.4256e-08 | -0.0008133 | -1.3204 | 0.9998 |  | -3.0741e-26 | -5.3759e-22 | 2.9856e-18 | -2.9935e-15 | -5.0308e-12 | 1.0824e-08 | -9.851e-06 | -1.3205 | 0.9984 | AE GDF |  | 2034 | 06/17-06/19/2024 |  | -1.323 | 4.8864e-05 | -1.3232 | 6.4844e-05 | 120 | 2.63
AE GDF | 2035 | 2035 |  | 20218 | 1135 |  |  | AE GDF |  |  | 03/30-04/04/2024 |  |  | -0.147 |  |  | 2.594e-17 | -1.3241e-13 | 2.2233e-10 | -1.3856e-07 | -0.00082424 | -1.2638 | 0.9998 | 7.6504e-27 | -6.4174e-23 | 2.188e-19 | -3.8913e-16 | 3.8697e-13 | -2.1755e-10 | 6.9876e-08 | -1.6771e-05 | -1.2638 | 0.9987 | AE GDF |  | 2035 | 05/28-05/30/2024 |  | -1.2673 | 6.1007e-05 | -1.2673 | 5.8794e-05 | 30 | 3.5
AE GDF | 2036 | 2036 |  | 20263 | 1139 |  |  | AE GDF |  |  | 04/30-05/06/2024 |  |  | -0.096 |  |  |  | -2.8207e-14 | 8.1569e-11 | -5.5323e-08 | -0.00085002 | -1.3203 | 0.9999 |  | -1.8935e-24 | 1.4849e-20 | -4.6236e-17 | 7.3881e-14 | -6.6481e-11 | 3.6722e-08 | -1.7052e-05 | -1.3202 | 0.9989 | AE GDF |  | 2036 | 06/17-06/19/2024 |  | -1.3232 | 9.2835e-05 | -1.3235 | 0.00014127 | 370 | 2.85
AE GDF | 2037 | 2037 | [5 empty] | AE GDF
AE GDF | 2038 | 2038 | [5 empty] | AE GDF
AE GDF | 2039 | 2039 | [5 empty] | AE GDF
AE GDF | 2040 | 2040 |  | 20279 | 1145 |  |  | AE GDF |  |  | 12/30-01/04/2024 |  |  | -0.053 |  |  |  | -2.0002e-14 | 5.293e-11 | -4.542e-08 | -0.0008189 | -1.3061 | 0.9999 | 4.1693e-27 | -3.4121e-23 | 1.1296e-19 | -1.959e-16 | 1.9689e-13 | -1.2432e-10 | 5.2955e-08 | -1.8308e-05 | -1.306 | 0.9993 | AE GDF |  | 2040 | 07/15-07/17/2024 |  | -1.3095 | 4.7197e-05 | -1.3096 | 4.8854e-05 | 70 | 3.36
AE GDF | 2041 | 2041 | [5 empty] | AE GDF
AE GDF | 2042 | 2042 | [5 empty] | AE GDF
AE GDF | 2043 | 2043 | [5 empty] | AE GDF
AE GDF | 2044 | 2044 | [5 empty] | AE GDF
AE GDF | 2045 | 2045 |  | 20266 | 1144 |  |  | AE GDF |  |  | 01/05-01/11/2024 |  |  | -0.071 | [5 empty] | -4.5863e-08 | -0.0008206 | -1.2779 | 0.9998 |  | -5.8433e-24 | 4.8371e-20 | -1.5853e-16 | 2.6244e-13 | -2.3276e-10 | 1.0887e-07 | -2.7506e-05 | -1.278 | 0.9962 | AE GDF |  | 2045 | 07/15-07/17/2024 |  | -1.2814 | 7.1108e-05 | -1.2815 | 8.0803e-05 | 110 | 3.53
AE GDF | 2046 | 2046 | [5 empty] | AE GDF
AE GDF | 2047 | 2047 |  | 20282 | 1143 |  |  | AE GDF |  |  | 03/21-03/27/2024 |  |  | -0.055 | [6 empty] | -0.00083174 | -1.3016 | 0.9998 |  |  | -4.4102e-21 | 2.9048e-17 | -7.2712e-14 | 8.6047e-11 | -4.7625e-08 | 5.0756e-06 | -1.3015 | 0.9964 | AE GDF |  | 2047 | 07/15-07/17/2024 |  | -1.3053 | 4.4582e-05 | -1.3053 | 4.8689e-05 | 30 | 3.66
AE GDF | 2048 | 2048 | [5 empty] | AE GDF
AE GDF | 2049 | 2049 | [5 empty] | AE GDF
AE GDF | 2050 | 2050 | [5 empty] | AE GDF
AE GDF | 2051 | 2051 | [5 empty] | AE GDF
AE GDF | 2052 | 2052 | [5 empty] | AE GDF
Note on V7: Hans Jannasch: K2 run 1/23-24/2012 w/Ref VL#22
Note on G10: Luke renamed the original 7033 (probe leaked) to 7038 (new probe). Carole recaled new probe as 7033, which was then remnamed to 7038. Dana uploaded renamed "7038" file on 3/27/13.
Note on AI10: Hans Jannasch: CE on DF
Note on AI11: Hans Jannasch: CE on DF
Note on AI12: Hans Jannasch: CE on DF
Note on AI13: Hans Jannasch: CE on DF
Note on AA41: Hans Jannasch: Still doesn't get <50psi curvature.
Note on H45: Hans Jannasch: Printed Clearview flow cell for pH only painted black.
Note on D48: Hans Jannasch: Wa APF#9096
Note on E48: Hans Jannasch: Was CTD#5588
Note on D67: Hans Jannasch: Was 9265 for MRY tests
Note on E67: Hans Jannasch: Originally 5584 had bad S readings. Still deployed in Monterey Bay, but rebuild w/5590 before real deployment.
Note on E84: Hans Jannasch: Leaked during P-test in Seattle. Will need to rebuild w/new logger.
Note on AH91: Hans Jannasch: An unusually flat pressure curve (<1 mV), thus the low r^2
Note on D93: Hans Jannasch: was 9634 now is 12552
Note on D97: Hans Jannasch: was 9752 12382??
Note on E102: Hans Jannasch: NO FLOW CELL
Note on E105: Hans Jannasch: CF hull
Note on E115: Hans Jannasch: 4521 was #6& #4 port
Note on AP138: Hans Jannasch: High?
Note on E140: Hans: recalibrated with CTD#7973
Note on AP140: Hans Jannasch: High?
Note on F146: Hans Jannasch: 746 replaced damaged 738.
Note on AP150: Hans Jannasch: Low?
Note on D152: Hans Jannasch: 12575
Note on D153: Hans Jannasch: 12559
Note on D155: Hans Jannasch: 12542
Note on D167: Hans Jannasch: 12361
Note on AP167: Hans Jannasch: Low?
Note on D168: Hans Jannasch: 12369
Note on D169: Hans Jannasch: 12363
Note on F174: Hans Jannasch: 12544 7914 766
Note on A257: Stems built with Dow Corning Molykote 55 oring grease. No longer using vac grease for o-ring lube
Note on E262: Peter Walz: Ib railed after stem built. Replaced chip and gave to Yui
Note on A279: Peter Walz: ApfID 12544 Recovered by CSIRO with CTD 11434 and MSC 838. Redeployed with pH SBE 12581, CTD 11434 and MSC 838
Note on A284: SBE Fuzzbutton sensors first run with longer (0.060") fuzzbutton connectors, Re-ran in May-June with shorter fuzzbuttions. July 2018 55 durometer o-rings swapped out with 75 durometer orings. All sensors and stems returned to SBE for rebuild with durometer 70 orings and shorter fuzzbuttons (0.055"), July 2018 SBE to run f(P) and MBARI to spot check a few sensor pairs to compare f(P) calcs. MBARI to do all K0 runs on completed stems from SBE.
Note on AK292: SBE built and f(P) and K2 at SBE. From Charlie Branham
Note on AK293: SBE built stem and f(P) and K2 data from SBE Charlie Branham
Note on B334: Peter Walz: Began using matlab for data calcs
Note on C346: Peter Walz hi k2 stems - tested ok at MBARI and ok for use on floats
Note on C347: Peter Walz: hi k2 stems - tested ok at MBARI and ok for use on floats
Note on A359: Peter Walz: low pressure hook sensor from SBE. Recalibrated at MBARI and ok for use
Note on A360: Peter Walz: low pressure hook sensor from SBE. Recalibrated at MBARI and ok for use.
Note on A361: Peter Walz: drifty sensor from SBE. Recalibrated at MBARI and ok.
Note on A366: Peter Walz: Starting in Nov 2022: Raw cal data provided from SBE for MBARI to use in Matlab to calculate sensor coeffs. No longer using cal sheets from SBE.
Note on B368: Peter Walz: Ben D using for lab work Nov 2022
Note on B374: Peter Walz: Yui and Joe used for lab work Nov 2022
Note on B381: Peter Walz: Feb 2023 being used by Yui T lab
Note on A412: Peter Walz: MSC 856 on CTD 11766 flooded during testing at UW
Note on C446: Peter Walz: O-ring groove damaged on tombstone PEEK part. Save for lab or local deployments
Note on K451: DF 280 and MSC 891 were built onto CTD 12445. CTD had bad salinity so these sensors were moved to another CTD.
Note on B455: Peter Walz: K0 run Vrs noisy - use for lab expts not deployed on soccom
Note on B459: blk/whi/red ISFET wires offset and epoxied without heat shrink on solder joints
Note on B464: Josh Plant: in situ f(P) different from cal f(P). JP used deployment bottle pH to regenerate k0 & Pcoefs 09/24/21. K2 from matlab GUI
Note on V464: Josh Plant: jp adjusted usin GUI output
Note on AA464: Josh Plant: new pcoefs from refit to deployment bottle pH
Note on AP464: Josh Plant: jp adjuste as part of pcoef refit
Note on AK491: Peter Walz: Started using MATLAB for calcs of K0
Note on B499: Began using MATLAB scripts for calculating PT and K0 data
Note on B513: Peter Walz: contact to ISFET chip damaged during stem assembly. May not be reparable. Dec 2021.
Note on C521: Peter Walz: DF346-353 built using chips from May 2021 shipment
Note on V521: Peter Walz: May 2021 lot of ISFET chips from SBE. They also saw "out of family" k2 values for this recent batch of chips.
Note on V522: Peter Walz: May 2021 lot of ISFET chips from SBE. They also saw "out of family" k2 values for this recent batch of chips.
Note on C528: Peter Walz: Sent to SBE for testing in their PT chambers. Returned and rerun in PTC Nov 2022
Note on B540: Peter Walz: Odd fP for 365 366 and 367
Note on A546: Peter Walz: DSD sensor pairs 371 and 372 given to Yui and were not fP or k2 calibrated in our lab
Note on A547: Peter Walz: Very high K0 pump offset. Did not use for floats
Note on AR547: Peter Walz: very high pump offset - did not use for floats
Note on B558: Peter Walz: odd cal data - save for lab use?
Note on B560: Peter Walz: odd k2 f(P) response. Recal in PTC?
Note on AR568: Peter Walz: stopped sanding ref pellet after seawater equilibration. Reduces pump offsets during k0
Note on D572: Peter Walz: AOML float
Note on D573: Peter Walz: AOML float
Note on B575: DF 402 and 403 built into nanoFET sensor for Yui T
Note on B581: Peter Walz: To Yui T lab for their use
Note on B582: Peter Walz: DF 409 and 410 built for Yui T use, NanoFET builds
Note on B585: Peter Walz: DF 412-419 using the new run of Honeywell ISFET chips (6/20/23). Blue heat shrink on paddle wires
Note on A602: Peter Walz: GDF sensors starting with GDF11 now have machined corners for the isfet pocket.
Note on C629: Peter Walz: used for coeverplate testing. Ti coverplate leaked during PTC run Jan 27 2023. Sensor flooded. Rebuilt and recal with new parts. Ti no shoulder coverplate
Note on C632: Peter Walz: Ti no shoulder coverplate
Note on C633: Peter Walz: Ti no shoulder coverplate
Note on A634: Peter Walz: GDF2006 does not exist. That number was skipped during the Applied Engineering builds.
Note on I634: Peter Walz: GDF2006 does not exist. That number was skipped during the Applied Engineering builds.
Note on C635: Peter Walz: Ti no shoulder coverplate
Note on C636: Peter Walz: Ti no shoulder coverplate
Note on AR636: Peter Walz: ref pellet not polished after sw equilibration
Note on C638: Peter Walz: Flooded during PTC cal, ISFET chip cracked
Note on AR640: Peter Walz: ref pellets not polished prior to k0 calibrations
Note on C641: Peter Walz: Ti no shoulder coverplate
Note on C649: Peter Walz: Ti coverplate
Note on C651: Peter Walz: Ti coverplate
Note on F656: Peter Walz: older MSC board, blue mask (Rev B)