## Slide 1: 20200908- Scott, Thom, Yui, JKW meeting

## Slide 2: Naming and Notes
Mfet B and mFET C (marked B)
  C has better solder masks for component assembly
‘Trin’ has been hand assembling, needs more parts to use automation
Silkscreen Serial numbers?
  White corner w/ handwritten numeration? Sticker?
NAMING CONVENTION: Board rev, Build##, ###
  Example, RevC, build 2, board 42 (C02042)
Software dependent on mCAP so make sure CAP numerated as well.
Thom: Fast sampling on RevA mFET?
  NO: Just order the new version and remove RevA from circulation
More notes of Failure as they come off:
  We want to avoid 25% failures (Find cause: Oscillator, other problem?)

## Slide 3: Manufacturing
Scott: mFET/mCAP is best array as you then order ‘40 units (ex)’
  Yui: Keep separate for what we need
Already Failed boards:
  Finding the common issue
  Thom: Trin does some tuning on the code loaded
  Scott: if all just fabrication errors we get that repaired
    Thom: Had Jim and TM look under microscope.
Default assembly should take away all options/default jumpers
  Reworked boards had Do not sleep jumper installed
  Jumpered for ‘fast reset’ currently
Option to not sleep explanation: Processor has ability to turn off and hold power supply off while sleeping. Jumper left in as a way to bypass for bring-up
  Wake switch different, as it goes to processor. Jumper bypasses HIB line. Keep location but don’t populate w/ jumper.

## Slide 4: Noise Issue:
Still to test: Single bias battery Jumper.
  Currently: leave default as 2x AA batteries.
Scotts thoughts:
  Thermocouples and dissimilar materials
  What metal did Gene use? What material in adapters
Vk follows temperature
Email Greg and Josh for HVAC schedule.
Ken J theory: Microbubble reacting to room issue? (Setra didn’t show)
Thom: RevA mCAP noise issues historically?
  YT: was put on glider, and different electrode makes it go away.
Get EM noise meter from Thom
NEXT TEST: RevA/RevA
  1. Box shield for mFET? Would help w/ EM noise.
  2. Heat gun on board/jumper (turn on and not heating, just noise, then heating CAP/FET.)
  3. ESP Thermal chamber

## Slide 5: Noise Issue: (Continued)
Analog Simulator: Insulate from temperature.
  Make more Jensenian Simulators
  Add a FET?
    Way to test w/o ISFET?
    Ideally: Multiple boards getting checked simultaneously
Scott: Ranges of values for component checks

## Slide 6: Odds and Ends
Nano-FET: Jose status?
  Scott Isolation testing (needs an ISFET for this) and Noise assesment
  Once good: Hook up pH Sensor and test.
Timeline of mFETC
  Scott: Nearing end of September before checkout (1wk) so October
    If all passes, ready to fabricate C
Boards for Thom to explore
  038- from Canary deployment
  147- reworked and deployed in lab