Yui and JKW ran through the assembly of the mpHOx SCSP on 20220124:

TO DO:

- EE shop order for cabling and potting

  - Bulkhead termination (pinout) needed first

- ADAPT Bulkhead wires to Harwin wire

    - Need lengths

    - Ray Hom can do in future

- Clear fender washers for battery terminals?

    - Acrylic rod vs stock

- Drill divots thru battery end cap and use sealing screws

- Solution for Battery Rattle

- FLOW CELL:

    - Ream ¼-20 drill holes with:

      - Close fit: 0.2010 F

      - Free fit: 0.2570	H(17/64)

- SBE LOOKUP

    - XSG to MCBH conversion/adapter:

- Order:

    - Barrel/butt splice terminal crimper

    - Antifouling spray paint

      - https://www.amazon.com/Interlux-YBA068A-16-Aerosol-Antifouling/dp/B003E1Y31O

    - 8-32 x ⅝”

    - IV bags w/ same tubing as SBE pump (1/2” ID /3/4” OD tube)

BUILD NOTES: (Used 200JD_mphoxhousing.pdf)

- Bulkheads and Endcap:

  - Loctite 242 used to secure 8 and 4 pin male bulkheads

  - Cable management:

    - 4-40x 1/4” for cable tiedowns

    - 8-32 x ½” PAN head phillips w/ Grey Fender Washers

      - When installing this ensure enough slack and NO PINCH

    - Wire terminal (Ring terminal #8 w/ star grip), with Spring for Negative batt

      - Make sure ring is inserted into spring so that when screw is installed good continuity is maintained between spring/terminal

- Durafet:

  - Grease O-ring w/ Vacuum grease

  - Place 2x White delrin washers in receptacle (PN?!?)

  - Push Durafet into proper hole, and push til sensor has clicked into its position.

  - Clock the FET towards the flow path

- OPTODE:

  - Grease O-ring w/ Vacuum grease

- ASSEMBLED:

- Sensor/Battery support block

  - No fender washers are needed to attach the positive Bias terminal.

    - 8-32 x ⅝” w/ Nylock (SOCKET MAKES BETTER CONNECTION)

      - Watch amount of solder on terminals, can potentially rattle

  - Insert D-Cell Bias battery first (FLAT ON SPRING, NIPPLE ON SOCKET) and ensure the cables are in the middle of the channel and not being pinched

  - ¼-20 2.25” 4x used to secure backing block to endcap and screw til it bottoms out

  - Mount mpHOx board onto clamps that affix to tube:

    - 4-40X 1” for ring clamp assembly

    - 4-40x ¼” stand offs and screws for board mounting

  - Screw in Battery Tube until it bottoms out

  - Insert D-Cell batteries flat end first into tube.

    - Socket for positive side = 8-32x 1”

    - Battery will rattle somewhat even when tightened down all the way

  - Ideal Wire Slack:

- FLOW CELL

  - Align flow cell as indicated in schematics

    - NOTE: 3D printed cell kind of ‘soft’ not best tolerances

- MOUNTS/BRACKETS

  - Install lower bracket assembly as indicated by JD schematics

  - PUMPS:

    - Remove cap and place cap into flow path with tubing

      - ⅛” Drive for nylon tipped set screws x2

    - MICROCAT:

      - Modifications:

        - Remove Neck bracket near bulkhead

          - 9/64” hex driver

            - NOTE: replace w/ threaded inserts?

        - Remove foot bracket at far end

          - 10mm nut driver

        - Flow Path and TBT Blocks

          - Remove Flat head screws and keep. Replace with 3D printed component

  - Clamp Mounts together

    - 1/4-20 2.5” or 2.25” TBD

    - Ensure tolerances are left for taping/painting the instruments

      - NOTE: Split bottom bracket in half?