20210610

Nanofet testing

Scott Jensen, Joseph Warren

Goal: To prove that the latest unipolar bias works (no virtual Ground). We will test solutions with varying pH, temperature, and add/remove electrical noise in the form of a heat gun as well as a stir plate.

Setup:

- Nanofet board:

  - rB mFET: Q27812/ 112. V3.3.1 mfet V2 firmware.

  - Isfet adapter board rA

- Extech PSU: 12V 0.25A

- AA: PKCell 3.6V ER14505

- Durafet II: 1914Y200000087947

  - Durafet UDA

Note on differences in data fields:

- Standard = This board

- Extrefoff = Extref V

  - Read Extrefoff in test menu as if it were Vint usually (not Vint_off).

  - There is not Vint_offset output in this hardware.

- Intrefoff = cntr V (Vk)

- ExtrefV = NA

- IntrefV (Vint) = NA

- Cntr I = Cntr I

  - Difference between Subst Curr and Subst Vol = Cntr I

- Cntr V = NA

- Subs I = SubsV2Gnd

- SubsV = SourceV

- Vtherm=Vtherm

Test 1: 3, 8, S Testmenu with FSW of known pH (Vint ~ 0.08 and 0.1M HCl ~-0.3x).

Notes:

- Vk similar to other mFETS

- Vext_off is dissimilar to other mFETS

- When Durafet was hooked into K2 Box/DB15 MF 146, voltages were correct.

- THERE IS NO NERNST NUMBER

  - = Vint, not Vintoffset

- Allowed us to see that actual Vint is the same across the boards (146 and ISFETADAPT)

TEST 2: Isfet adapter with MF146, Excel tab NFTST02_Response time

Notes:

- Do not have Subst Curr or Subst Vol in the config data fields options, but are available in the test menu.

- File is configured properly, see excel spreadsheet for notes.

pH Solution Tests:

Test 1: File NFTST02.txt

- TRIS in FSW: 8.2, 7.6, 7.4. 60s each solution, manual stirring 3x

- Then: FSW/0.1MHCL mix (pH~3) to 8.2 60s each, 3x.

Summary: We made a rough calibration curve for 8.2, 7.6, 7.4, and tested the FSW/HCl unknown mixture, to see if we could determine its pH, and were able to (1.588). This experiment also shows a nominal response time of 30s til stable readings.

Noise Testing:

Test 1: File NFTST03.txt, excel tab NFTST03_Noise

- Tris in FSW: 8.2, stir bar. 60s resting, 60s stirred, 3x. Stir at 500 rpm.

Summary: No noise was able to be detected throughout the experiment.