Notes

- nanoFET Vrse_stdev was ~200 uV in over holiday testing. Why?

  - Waterbath on/off?

  - Wiring may not be shielded

  - Do a battery test: power mFET + nanoFET with battery. Let it log internally. Basically we need to understand if this is a limitation in the circuitry, or if it is external noise getting into this experiment.

    - Repeat this battery test in several locations in the lab.

    - Replicate with 2x nanoFET.

    - Plug in 2x DSD into both mFET and nanoFET.

  - Vk_std was at 10uV. So it seems like noise is coming into reference line.

- Mock-FET: P-FET attached to nanoFET.

  - Test results?

- Will Jensenian work with the nanoFET?

  - Maybe.

  - There were rough estimates on the output voltages. They will be different.

Circuit

- Cap for RC filter is ground ref, but Vrse is relative to source.

- Can cap be tied to source? -> need further thought.

- Potentially, we can remove the cap to see if that removes noise.

Experiments

- Short term room temp HCl noise test. (Joe + Irene)

- Get familiar with nanoFET circuit, and then have a meeting to go over gory details (everyone. Early February?)

- Send ISFET documents (Yui)

- Update project library with most updated schematics (Scott)

- Underway pH measurements (mFET/nanoFET, batt powered, internally logged) (Joe); Only when lab tests are done.

- Nernstian experiment: titrate SW over large range of pH, compare Vrse from mFET/mCAP vs nanoFET. (Joe + Irene)

- WireWalker when it’s ready (Irene)

- When Firmware is ready for SIO, then we can send them a controller + API library. (Communicate any issues to Thom).