Coral Macro Operational Procedures

September 20, 2021

Version 1

- Deployment

  - Once the tripod is assembled and all the cables (strobes, battery, lasers, and DSC (DeepSea Connect) and connected to the Camera housing, assemble the following to prepare the unit for deployment:

    - Laptop computer with: uCon, VirtualHere, Sony Imaging Edge Desktop

    - Power supply capable of +24V at 5A

    - Deck cable with MCIL10F at one end, and banana plugs, RJ45 and DB9 at the other end

  - Set laptop network connection before turning system on:

    - Network addresses:

      - RaspberryPi in Camera Houing

        - IP address: 192.168.1.33

        - Subnet mask: 255.255.255.0

        - Gateway: 192.168.1.1

      - Ensure laptop set to:

        - IP address: 192.168.1. X

        - Subnet mask: 255.255.255.0

        - Gateway: 192.168.1.1

  - Plug deck cable into Camera housing MCBH10F connector and to power supply and laptop. The DB9 is not currently used in this system.

  - Turn the power supply on. At +28V the system should draw around 150 mA.

  - Test network connection by opening a command window and “ping”ing the camera IP address. (command is “ping –t 192.168.1.33”, you can also “ping raspberrypi.local”)  To stop pinging press “ctrl”+”c” at the same time.

  - Open up VirtualHere window, when the Raspberry Pi is booted up, you should see:

  - Right click on USB-RS232 and click on “Use This Device”

  - Now it should look like this:

  - Next open up the terminal software like uCon to talk to the Snappy Camera controller. 9600, 8,N, 1. You can use the Portscan option in uCon or Device Manager to determine the com port added via Virtual Here.

  - Open uCon and implement logging and timestamping as described in section 1.m. .

  - Once uCon is open, press enter and you should see the software version on the Snappy Controller.  As of Sept 20, 2021 the version is 1.4.8.

  - There is no menu option for Snappy so use the command set at the end of this document.

  - Before programming the mission verify the function of the following:

    - Test strobes

      - Type “26”  to turn on charging to the strobes. Note if the battery plugged in you won’t see a change in the power supply, if not it should increase by up to 200 mA during charging – which happens over a few seconds.

      - Type “22” to see voltage on strobe, see if it is charged now ~35 V

      - Type “30” to trigger the strobes, type another “22” right after to see voltage is significantly less. These strobes charge really fast (in a few seconds) so if you wait to press “22” they will recharge before you read the voltage.

      - Type “27” to turn power to the strobes off. They will slowly discharge.

    - Test the camera

      - Type “19” plus enter to turn the cameras on

      - Unlike other time-lapse camera systems this alone does not connect the Camera via the USB. There is a board with latching relays that has to be turned on. To do this you need to open CMD window.

      - SSH into the raspberry pi by ssh pi@192.168.1.33

      - The password is disco123

      - Once you have a promt type USBOn (see image below)

      - the camera is on it should appear in the VirtualHere window:

      - Right Click on “ILCE-7RM3” and select “Use this Device” from menu, this should result in opening up an empty file explorer window, just close the window and open the Sony Remote software. Here is the icon:

      - When you double click on the Remote icon you will see this window:

      - Double click where it says ILCE-7RM3, which should open up this window:

      - This software allows you to see the Live View and take an image by pressing the camera icon inside the circle under the “Shooting” menu, you can also change the ISO, F-stop, Shutter speed.

      - Take an image using this software and verify an image is collected.

      - Turn on power to one strobe and take an image using the Remote software to verify the camera is triggering the strobes.

      - You can also use the command “32” to fire the shutter through Snappy, the image will still be uploaded to the laptop. Don’t use “34” in this mode as it may power cycle the camera.

      - Use EOS utility to verify camera settings:

        - ISO= 500

        - Shutter = 1/25

        - F = 14

        - Image quality = Large Fine

        - Images= Raw

I

      - Once this is tested and the settings are complete you can close the Remote Utility window and then in the cmd window type in USBOff and hit enter. This should disconnect the camera from the raspberry pi. It should no longer appear in the  Virtual Here window.

    - Once camera is set and tested type “20” to turn camera off

    - Read battery voltage by typing in “14” and enter. The battery pack should start out around +30V

    - Once you are satisfied that all subsystems are operational, it is time to start a mission

  - Starting the mission

    - Set Clk by typing “17”, this will lead you through setting the date and time. Use only 2 digits for the year.

    - Type “18”to read clock and confirm correct time.

    - Set  delay: type “9” to set the delay. This is the time between sending the deploy command and starting the mission. I set for 120 seconds to allow a shutdown to the raspberry pi and turn off the deck cable. The Sony camera can still work even if the pi is fired up as long as the USB connection is turned off through the cmd window. Confirm delay by entering “10” to read the delay value

    - Set sleep: use command “7” to set the sleep. Sleep is the number of minutes between taking a picture. Set this to 1 minute less than you want the elapsed time to be. Read that it is set correctly by typing 8

STOP HERE!!! IF you are going to set the camera up in situ with the ROV move to step 3.

    - Once sleep and delay are set you can start the deployment by pressing “3”. Once the deployment is started if you want to interrupt it, you will need to either disconnect battery power or wait until the PIC wakes from sleep to take the next picture. See step “1. o”

    - Disconnect deck cable, and replace dummy plug (clean and lubed with medium 44).

    - Fill lasers with water to get air bubbles out.

    - Plug in lasers

    - Make sure beacons have fresh batteries and are turned on.

    - Make sure cables zipped up and no tools on stand

    - Make sure pin holding camera to stabs is in place

    - Make sure ROV DSC (Deepsea connect) can is connected to MINK10FCRL science port (120VAC, Eth) and the paddle is accessible to the ROV manip

  - To “halt” the system you need to type “5” either:

      - If you can disconnect the battery, turn deck power off for a minute – so that it is completely powered down. Have your ucon window up and ready, same with virtual here. As soon as you power it up wait for the virtual here to bring up the USB-RS232 cable then select “Use this device”, then jump to ucon and use the button with the two arrows pointing to each other to connect. You should see the system counting down from 70, type “5” to halt. Do this twice.

      - IF you cant disconnect the battery you need to connect through the deck cable or the Deepsea connect with the test box. Once you are connected and powered up, use Virtual Here to enable the USB to serial server. Open uCon and patiently wait for the PIC to wake up. You then have 20 seconds to halt the program by pressing “5”.

- Setting up Camera and Starting Deployment in situ via the ROV and Deepsea Connect

  - Follow the deployment testing procedure steps 2a-2n.iv DO NOT start Deployment.

  - Turn on the Sci port providing the 120 VAC and Ethernet, to which the DSC can on the ROV is connected. Ping the ROV DSC 192.168.1.21

  - The ROV will need to connect the two DeepSeaConnect (DSC) paddles, one from the camera tripod and one from the ROV. This is putting the “toast” into the “toaster”.

  - Next the pilots need to align the magnet on the boomerang with the magnet inside the Camera DSC can.

  - Once DSC pair are firmly connected you should try to ping the Camera DSC 192.168.1.20, if you get a response then ping the raspberry pi 192.168.1.33

  - Once this is connected repeat from the Deployment section steps 1f-1n.iv. Now it is time to set up the camera for your target.

  - Reminder, to get a good picture you need to charge both strobes up before firing the shutter trigger.  This takes a few seconds and you can just turn on “26” left strobe on, after the image you can turn it off “27”. If you want to be sure the strobes are left off (which is important to not waste the battery) typr “30” to fire, it will usually have enough for one burst before the capicitors are depleted. So type “30” a couple of times and you should see no light by the second time if the strobe charge circuit is truly off.

  - Go into the LiveView on the remote shooting menu. With the ROV lights and by turning up the ISO you should be able to get an idea what the camera will see. Charge strobes and have the ROV pilots turn off lights and take an image.

  - To change the angle of the camera have the pilots use the handle on the side of the tripod, use the live view to help direct.

  - The lasers overlap in the focal plane. Make sure the subject of most interest is in the plane where the laser dots converge.

  - When you are happy with the images it is time to start the mission.

  - Check the delay (interval before starting mission) is still set to 120, to check delay type “10”, if you need to set it type “9”

  - Check the clock by typing “18” if you need to set it type “17”, remember use just two digits for the year

  - Set the sleep  by typing “7” to one minute shorter than the desired interval, to check sleep type “8”

  - Start the mission type “3” to start the mission, once started Snappy will not respond to any more typed commands until delay is over. Immediately disconnect from the USB-RS232 ports in Virtual Here and have the pilots turn off the 120V to the Camera tripod DSC connection. This is imperative to see that the mission is working correctly.

  - In about 2.5-3 minutes you should see the strobes flash which means the first image was successful

  - Have the pilots reseat the DSC connector on the tripod frame

  - Have the pilots release the recovery line

  - Take screen grab of ucon window with communications to snappy before closing then check your log file worked.

  - The end!

- Recovery Procedures

  - Once camera tripod is on deck thoroughly rinse the system with fresh water.

  - Dry connector on battery can with towel, then disconnect the cable from the battery housing to the camera housing (blue cable). This removes power from the camera controller. Use a metal shell dummy cap to seal the connector on the battery housing, and a plastic or metal plug on the cable side to protect the connector.

  - Start up a laptop that has the following software: VirtualHere, uCon, and Canon EOS Utility. These three programs will allow you to download all data and images.

  - Connect the deck cable to the MCBH10F connector on the back of the camera housing. Secure with ziptie to the frame near the camera. In the lab, connect the RJ45 of the deck cable either directly to the laptop or to a network switch to which the laptop is already connected. DO NOT have switch attached to any external network (ship, shore, etc)

  - Network addresses:

    - RaspberryPi in Camera Housing

      - IP address: 192.168.1.23

      - Subnet mask: 255.255.255.0

      - Gateway: 192.168.1.1

    - Ensure laptop set to:

      - IP address: 192.168.1. X

      - Subnet mask: 255.255.255.0

      - Gateway: 192.168.1.1

  - Power Supply: Set up a power supply capable of +28 V and at least 3 A. If you plan to test the strobes as well, you will need a power supply capable of 5A. Confirm that voltage and current settings and then plug in banana plugs to the power supply.

  - Turn power on to deck cable. You may see an initial voltage dip, but the power supply should draw about 120 - 150 mA @28V when the pi is booted up.

  - Test network connection by opening a command window and “ping”ing the camera IP address. (command is “ping –t 192.168.1.33”, you can also “ping raspberrypi.local”)

    - If you do not get a response after 40 seconds:

      - Check deck cable is drawing correct current

      - In command window type “ipconfig” then hit enter, check that your laptop is set at the right IP address

      - Confirm there are lights coming from the RJ45 connection from the deck cable and your laptop if using a switch

    - If you get a reply proceed to the next step

  - Open Virtual Here software (icon on desktop says “vhui64”

  - Ideally you see something like this:

  - You may see one or two entries for USB-RS232 Cable. Right click on each one you see and click on “Use This Device”, the order you click the devices effects which com port it is assigned. The com port number also varied by laptop.

  - Now it should look like this:

  - Next open up the terminal software uCon. The configuration window looks like this:

    - Click the “PortScan” button to update the serial ports in the “Serial Back End” section of the configuration. You should see an unused serial port

    - Connect at COMX 9600,8,n,1. Press OK and a terminal window will open.

    - If you press return and do not see a prompt you chose the wrong com port, to change it go into go to the configuration menu at the top of the screen, select “Backend” then “Com port” and you will be taken to a menu where you can change the port number.

    - Turn on time-stamping and logging in ucon

      - In the Config menu select “Timestamp”

      - Select “Enabled” and “Stamp at toggle” and “Local only”

      - Press OK and continue to setting up logging

      - Under the “Logging” menu select “Standard” then you will see this window:

      - Select “Enabled” and then using the “Browse button” you can name and save the file in a place you won’t forget. Press Ok when finished.

    - Press Enter, most likely nothing will happen because the PIC on Snappy which you are communicating with is asleep or it is trying to take a picture. Now that you have everything configured right you will need to restart Snappy.

- Fresh restart:

    - Disconnect the serial port in the uCon window by clicking on the button with two arrows pointing away from each other. See red arrow above for help finding the button in the figure below.

    - Turn power off to the deck cable for at least a full minute. The next steps are time critical for stopping the program so read fully before starting.

    - With ucon and VirtualHere windows open and visible start turn power supply back on.

    - As soon as the USB-RS232 Cable appears in the VirtualHere window, press the connect button on in uCon (two arrows pointing to each other)

    - When the com port is connected you should see that the program is counting down, as quick as you can type “5” and press enter. This should halt the program. For good measure I always do this step 2 times.

    - Once the program is halted typing the command “36” then “Enter” will spit out the stats of the last deployments. Make sure to print out twice then you can disable logging. Take a screen shot incase log file does not work.

- Now that you have the strobe data, the next step is to download the images. You need to open the camera housing to get the cards out of the camera and offload the images.

- After images are all downloaded, which make take hours, back the up on another disk drive.

- Now you can close all the windows and turn off the power supply and disconnect the deck cable.

- To reformat and clear both flash and SD cards – there should be an option in the “Select and Download” window, maybe the lower left hand corner. Double check this.

Snappy command set

- Sleep

- Read Wake

- Deploy

- Resume

- Halt

- # Images

- Set sleep

- Read sleep

- Set delay

- Read delay

- Set timeout

- Read timeout

- Set batt

- Read batt

- Set SN

- Read SN

- Set clk

- Read clk

- Cam on

- Cam off

- Cam power?

- Left strb V?

- Right strb V?

- Chg Left strb

- Chg Right strb

- Left strb on

- Left strb off

- Right strb on

- Right strb off

- Trig left strb

- Trig right strb

- Trig far shut

- Trg near shut

- Snap far cam

- Snap near cam

- Get status