Coral Macro Operational Procedures

December 1, 2021

Version 3

- Deployment Checklist

  - Plug in battery cable (check o-ring)

  - Plug in lasers cable (use medium 44 grease)

  - Remove lens cap

  - Cut ziptie limiting tilt

  - Connect magnet for DSC

  - Make sure all tools clear

- Setting up Camera and Starting Deployment in situ via the ROV and Deepsea Connect

  - Set laptop network connection before turning system on:

    - Network addresses:

      - RaspberryPi in Camera Houing

        - IP address: 192.168.1.33

        - Subnet mask: 255.255.255.0

        - Gateway: 192.168.1.1

      - Ensure laptop set to:

        - IP address: 192.168.1. X

        - Subnet mask: 255.255.255.0

        - Gateway: 192.168.1.1

  - Plug Ethernet cable into port where the ROV pilots tell you to

  - Turn on the ROV Sci port providing the 120 VAC and Ethernet, to which the DSC can on the ROV is connected. Ping the ROV DSC 192.168.1.20. To do this type “cmd” in the search window to open up a command window. Ant the prompt type >ping  -t 192.168.1.20. A ctrl+C will halt the program.

  - The ROV will need to connect the two DeepSeaConnect (DSC) paddles, one from the camera tripod and one from the ROV. This is putting the “toast” into the “toaster”.

  - Check that the magnet is still aligned on the DSC housing.

  - Once DSC pair are firmly connected you should try to ping the Camera DSC 192.168.1.21, if you get a response then ping the raspberry pi 192.168.1.33

  - Open up VirtualHere window, when the Raspberry Pi is booted up, you should see: (you will see the datalogger too, just ignore)

  - Right click on USB-RS232 and click on “Use This Device”

  - Now it should look like this:

  - Next open up the terminal software like uCon to talk to the Snappy Camera controller. 9600, 8,N, 1. You can use the Portscan option in uCon or Device Manager to determine the com port added via Virtual Here. It will be the only one listed as “COMX” usually X=3.

  - Turn on time-stamping and logging in ucon

      - In the Config menu select “Timestamp”

      - Select “Enabled” and “Stamp at toggle” and “Local only”

      - Press OK and continue to setting up logging

      - Under the “Logging” menu select “Standard” then you will see this window:

      - Select “Enabled” and then using the “Browse button” you can name and save the file in a place you won’t forget. Press Ok when finished.

  - Once uCon is open, press enter and you should see the software version on the Snappy Controller.  As of Sept 20, 2021 the version is 1.4.9.

  - There is no menu option for Snappy so use the command set at the end of this document.

  - Test all functionality once you have connected to the camera in situ:

    - Test strobes

      - Type “26”  to turn on charging to the strobes. Note if the battery plugged in you won’t see a change in the power supply, if not it should increase by up to 200 mA during charging – which happens over a few seconds.

      - Type “22” to see voltage on strobe, see if it is charged now ~35 V

      - Type “30” to trigger the strobes, type another “22” right after to see voltage is significantly less. These strobes charge really fast (in a few seconds) so if you wait to press “22” they will recharge before you read the voltage.

      - Type “27” to turn power to the strobes off. They will slowly discharge. Type “22” to make sure the voltage is dropping.

    - Test the camera

      - Unlike other time-lapse camera systems this alone does not connect the Camera via the USB. There is a board with latching relays that has to be turned on. To do this you need to open CMD window.

      - SSH into the raspberry pi by ssh pi@192.168.1.33

      - The password is disco123

Pause before the next two steps. The next two steps are time sensitive. After you turn the camera on (step 4) you must turn the USB port on (step 5) within 30 seconds.

      - Type “19” plus enter to turn the cameras on

      - Once you have a prompt type USBOn (see image below)

      - the camera is on it should appear in the VirtualHere window:

      - Right Click on “ILCE-7RM3” and select “Use this Device” from menu, this should result in opening up an empty file explorer window, just close the window and open the Sony Remote software. Here is the icon:

      - When you double click on the Remote icon you will see this window:

      - Double click where it says ILCE-7RM3, which should open up this window:

      - This software allows you to see the Live View and take an image by pressing the camera icon inside the circle under the “Shooting” menu, you can also change the ISO, F-stop, Shutter speed.

      - Use EOS utility to verify camera settings:

        - ISO= 800

        - Shutter = 1/25

        - F = 22

        - Image quality = Large Fine

        - Images= Raw

I

  - Reminder, to get a good picture you need to charge both strobes up before firing the shutter trigger.  This takes a few seconds and you can just turn on “26” left strobe on, after the image you can turn it off “27”. If you want to be sure the strobes are left off (which is important to not waste the battery) typr “30” to fire, it will usually have enough for one burst before the capicitors are depleted. So type “30” a couple of times and you should see no light by the second time if the strobe charge circuit is truly off.

  - Go into the LiveView on the remote shooting menu. With the ROV lights and by turning up the ISO you should be able to get an idea what the camera will see. Charge strobes and have the ROV pilots turn off lights and take an image.

  - To change the angle of the camera have the pilots use the handle on the side of the tripod, use the live view to help direct.

  - The lasers overlap in the focal plane. Make sure the subject of most interest is in the plane where the laser dots converge.

  - When you are happy with the images and the settings are complete you can close the Remote Utility window and then in the cmd window type in USBOff and hit enter. This should disconnect the camera from the raspberry pi. It should no longer appear in the  Virtual Here window.

  - Once camera is set and tested wait one minute then type “20” to turn camera off

  - Double check the strobes are off by sending a “27”

  - Read battery voltage by typing in “14” and enter. The battery pack should start out around +30V

  - Starting the mission

    - Set Clk by typing “17”, this will lead you through setting the date and time. Use only 2 digits for the year.

    - Type “18”to read clock and confirm correct time.

    - Set  delay: type “9” to set the delay. This is the time between sending the deploy command and starting the mission. I set for 120 seconds to allow a shutdown to the raspberry pi and turn off the ROV cable. The Sony camera can still work even if the pi is fired up as long as the USB connection is turned off through the cmd window. Confirm delay by entering “10” to read the delay value

    - Set sleep: use command “7” to set the sleep. Sleep is the number of minutes between taking a picture. Set this to 1 minute less than you want the elapsed time to be. Read that it is set correctly by typing 8

  - Start the mission type “3” to start the mission, once started Snappy will not respond to any more typed commands until delay is over. Immediately disconnect from the USB-RS232 ports in Virtual Here and have the pilots turn off the 120V to the Camera tripod DSC connection. This is imperative to see that the mission is working correctly.

  - In about 2.5-3 minutes you should see the strobes flash which means the first image was successful

  - Have the pilots reseat the DSC connector on the tripod frame

  - Have the pilots remove the magnet from the DSC housing

  - Have the pilots release the recovery line

  - Take screen grab of ucon window with communications to snappy before closing then check your log file worked.

  - The end!

- Recovery Procedures

  - Once camera tripod is on deck thoroughly rinse the system with fresh water.

  - Dry connector on battery can with towel, then disconnect the cable from the battery housing to the camera housing (blue cable). This removes power from the camera controller. Use a metal shell dummy cap to seal the connector on the battery housing, and a plastic or metal plug on the cable side to protect the connector.

Snappy command set

- Sleep

- Read Wake

- Deploy

- Resume

- Halt

- # Images

- Set sleep

- Read sleep

- Set delay

- Read delay

- Set timeout

- Read timeout

- Set batt

- Read batt

- Set SN

- Read SN

- Set clk

- Read clk

- Cam on

- Cam off

- Cam power?

- Left strb V?

- Right strb V?

- Chg Left strb

- Chg Right strb

- Left strb on

- Left strb off

- Right strb on

- Right strb off

- Trig left strb

- Trig right strb

- Trig far shut

- Trg near shut

- Snap far cam

- Snap near cam

- Get status