DiSCO Multispec RevA

Design Review Notes

Irene Hu

Feb. 6, 2024

System Overview

DiSCO Multispec is a multispectral ROV-operated camera for imaging deep-sea corals.  Large portions of its design are based on EyeRIS / Chiton and other designs.  A system diagram (lite) is on the first page of the control board schematic.

Main housing:

- Contains the camera, lens, zoom control, AC/DC converter, and controller board

- Four connectors to the outside:

  - Optical comms for camera data and user comms with camera

  - High voltage / ethernet comms to/from ROV (MINK10)

  - Junction box (MINK8)

  - Spare (MCBH)

Junction box:

- Oil filled box with lots of connectors

- Controls the activity of 10 strobe pucks

Strobe pucks

- Wires soldered directly (from 6-pin cables), then potted

- Already made, please refrain from suggesting changes

We are still taking nominations for a clever name.

Board Functionality and Notes

Control board

- Ethernet interface to ROV for comms with user, through a mounted Adafruit module

- Power and trigger camera (trigger enables synching with strobes)

- Power and comms with stepper motor for zoom control (via external Pololu module)

- Control strobe pucks via junction box

- Power strobes via jbox, with help from a small capacitor farm

- Environmental data, RTC, Flash

- Note: board split into two to accommodate housing interior. Both boards are in the same schematic; nets shared across b2b connector have _T and _B labels

- Note 2: portions of schematic copied from SINKER controller, hopefully in a way that preserves layout

Junction box

- Enable/disable power to individual strobe pucks (no fault detection, use trial and error)

- Control on/off of different LED colors for all enabled strobe pucks

- LED control lines are fanned out to all strobe pucks through optoisolators

- 6 signal lines (from 8 pin connector) are demultiplexed to control above two functions

Strobe pucks

- 8 colors of LEDs

- Same form factor as EyeRIS strobes