Pre-Deployment Checklist

Version 3Jun26

For clock sync:

ssh: sintela@192.168.1.236

user: sintela

password: 5intela_440

Instrument set-up

- Check Engineering set-up (decimator, leaky cutoff frequency settings). If changes are made need to click save chain in the top right of the page.

- Check that correct optical set-up configuration is applied and automatically applied on reboot.

- Check noise floor is reasonable.

- Run OTDR to make sure there is no significant energy loss along cable or reflection at termination. Make sure data recorders are off before running.

- Check that GPS time setting is correctly enabled/disabled.

- Check that there is sufficient memory storage available: _______ MB available.

- Check that data recorders are set-up and files are being recorded to the right place.

- Check RAID is active (especially after deleting files, may need to configure again).

Instrument test via deck cable

- Install shorting plug.

- Apply power to deck cable with 30 V bench power supply.

- Confirm communication with DAS (ping 192.168.1.236, login to Sintela web-server dashboard 192.168.1.236:8080).

- Remove deck cable and plug in DeepSea Connect can.

Instrument test via DeepSea Connect (DSC)

- Remove foam and zipties binding instrument DSC can.

- Ask ROV pilots to power ROV side of DSC.

- Test communication to ROV side of DSC by pinging address 192.168.1.20

- Engage magnet to turn instrument DSC can ON.

- Test communication with instrument DSC can by pinging address 192.168.1.21

- Test communication with DAS by pinging address 192.168.1.236 or by using a browser and enter address 192.168.1.236:8080

- Power OFF DSC can by pushing ROV handle in.

MOB Day

- Test instrument via deck cable (see above).

- Test instrument via DeepSea Connect (see above).

- Test DAS recorders.

Pre-Dive

- Test instrument via deck cable and check set-up (see above).

- Synchronize time using the procedure in the next section below.

- Confirm shorting plug has o-rings and is in place.

- Connect batteries after applying fresh grease to connectors.

- Verify working, start data recorders: recording started at ________PDT

- Turn off power supply. Verify working on battery power.

- Close DAS application.

- Disconnect and stow deck cable and connect DSC after applying fresh grease to connector.

- Test instrument via DeepSea Connect.

- Test DSC ROV side working.

- Make sure data recorders are on.

- Confirm all connectors are snug and all cables are tied up.

- Confirm DSC ROV handle is in OFF position.

- Confirm deck cable is removed and DSC can is plugged in.

- Check comp levels: levels are __________________________

- Record beacon serial numbers and RF frequencies.

- Turn ON beacons (e.g. RF beacon and strobe) and pull any safety pins (e.g., drop weights).

- Notify deck crew to deploy instrument.

Synchronize time on the DAS using OLDSTARDUST Ubuntu laptop

The OLDSTARDUST Ubuntu laptop is setup to allow the DAS system to share the wi-fi adapter network while physically connected with an Ethernet cable. We just need to configure the DAS system to use OLDSTARDUST as a gateway

- Connect OLDSTARDUST to the DAS using the Ethernet cable on the deck test cable.

- Ensure OLDSTARDUST is on a wi-fi network with access to the web. Ping a well-known NTP server to test.
 $ ping time.cloudfare.com

- Request OLDSTARDUST to sync the system to that NTP server.
 $ sudo ntpdate time.cloudfare.com

- Use SSH to log in to the DAS instrument.
 $ ssh sintela@192.168.1.236  (passwd 5intela_440)

- Ping an address on the web and it should fail.
 $ ping time.cloudfare.com

- Add a gateway route through OLDSTARDUST (ip address 192.168.1.227).
 $ sudo route add default gw 192.168.1.227

- Repeat the failed ping command above. This time it should succeed.

- Request the DAS system to sync the system time using the NTP server.
 $ sudo ntpdate time.cloudfare.com

- Use the web browser on OLDSTARDUST to ensure that the laptop and DAS system clocks are in sync.