## Slide 1: Araon Expedition
Suggested TriggerBox Settings (full-duty mode with pauses):
c f 1 # Full-duty mode (square wave, 5 Hz)
c l 200 # Period = 200 ms → 5 Hz
c d 60 # Duration: 60 seconds of transmission
c p 30 (60) # Pause: 30 or 60 seconds between cycles
c w 1 # Wait full minute (sync to full minute)
c s # Save settings
s # Start transmission
Recommended vessel speed: 1.0 knots (0.5 m/s)
Sampling Frequency: 10 kHz

## Slide 2
Tests before measurements
1. Data Loggers synchronization test
Connecting one logger to the TriggerBox and to the transmitter box power them on
Open SelfPotLogger_SocketClient.
Press s → this sends a sync pulse to all connected loggers, aligning their clocks to the PC or to one designated master logger (set in the .cfg file with an “S” on the master device).
After synchronization, the software should display:
>>>>>>>> Sync successful ! <<<<<<<<
You can verify synchronization in Compare mode (c key) — all loggers should show identical timestamps.
Short recording using triggerbox measurement set up ( 15 minutes of recording)
Download the data and check them (send me raw data)
Repeat for all the data loggers
IMPORTANT: Loggers have to be synchronised to UTC time

## Slide 3
Tests before measurements
Gain Calibration Test
Connecting a logger to the TriggerBox and to the transmitter box power them on
Set the initial gain to 0 for all channels (±4.096 V input range).
Via software (SelfPotLogger_SocketClient): 0 261 0
Short recording using triggerbox measurement set up ( 15 minutes of recording)
Download the data and check them (send me raw data)
Repeat for all the data loggers