## GDF 2000
GDF Production Traveler
Serial Number: |  | 2000
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.297 | 0.096 | 0.061 | 0.018 | 0.303
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-07-05 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 2022-12-04 |  | He Leak check - pass.
8 | VL pellet resistance |  | 15 KΩ | RB | ~50 KΩ |  | Deployed with Ti no shoulder coverplate
9 | ISFET diode measurement |  | 0.655/0.654 | RB | 0.602/0.602 PW
10 | Counter electrode resistance |  | 0.23 Ω | RB | 0.2 Ω |  | used for coeverplate testing. Ti coverplate leaked during PTC run Jan 27 2023. Sensor flooded. Rebuilt and recal with new parts.
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name) | [5 empty] | CTD: 18688
14 | Install on endcap | [5 empty] | MSC: 1091
15 | Test / Calibration K0
16 | Shipped to | [5 empty] | PT CAL: 04/13 - 04/18/2023
17 | Deployed | [5 empty] | K0 CAL: 05/08 - 05/10/2023
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2001
GDF Production Traveler
Serial Number: |  | 2001
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.2975 | 0.096 | 0.06 | 0.018 | 0.302
4 | Gasket thickness |  | 0.032
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 |  |  |  | 12/7/2022 PW |  | comments:
7 | Other photos3 | [5 empty] | Very short green wire to CE plate. Difficult to replace CE plate for vac test
8 | VL pellet resistance |  | 16 KΩ | RB | ~60 KΩ |  | Pass He leak test
9 | ISFET diode measurement |  | 0.655/0.655 | RB | 0.606/0.606 PW
10 | Counter electrode resistance |  | 0.27 Ω | RB | 0.3 Ω
11 | Seawater Cond. (num. days) | [5 empty] | CTD: 18550
12 | K2 cal (file name) | [5 empty] | MSC: 1077
13 | f(p) cal (file name)
14 | Install on endcap | [5 empty] | PT CAL: 12/29 - 01/03/2023
15 | Test / Calibration K0 | [5 empty] | K0 CAL: 05/01 - 05/03/2023
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2002
GDF Production Traveler
Serial Number: |  | 2002
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.298 | 0.096 | 0.08 | 0.018 | 0.302
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | P. Walz 12/7/2022 |  | Pass He leak test
8 | VL pellet resistance |  | 16 KΩ | RB | 65 KΩ
9 | ISFET diode measurement |  | 0.658/0.658 V | RB | 0.607 V
10 | Counter electrode resistance |  | 0.24 Ω | RB | 0.3 Ω
11 | Seawater Cond. (num. days) | [5 empty] | CTD: 18820
12 | K2 cal (file name) | [5 empty] | MSC: 1099
13 | f(p) cal (file name)
14 | Install on endcap | [5 empty] | PT Cal: 05/25 - 05/29/2023
15 | Test / Calibration K0 | [5 empty] | K0 CAL: 07/17 - 07/19/2023
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2003
GDF Production Traveler
Serial Number: |  | 2003
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.036 | 0.299 | 0.096 | 0.06 | 0.018 | 0.301
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 1/10/2023 PW |  | No comments on He Leak test
8 | VL pellet resistance |  | 29 KΩ | RB | ~90 KΩ |  | No He leak test done. Delievered later than others in first batch of sensors from AE
9 | ISFET diode measurement |  | 0.648/0.648 V | RB | 0.601/0.601 V
10 | Counter electrode resistance |  | 0.46 Ω | RB | 0.2 Ω |  | Ti No Shoulder Coverplate
11 | Seawater Cond. (num. days)
12 | K2 cal (file name) | [5 empty] | CTD: 18683
13 | f(p) cal (file name) | [5 empty] | MSC: 1094
14 | Install on endcap
15 | Test / Calibration K0 | [5 empty] | PT CAL: 04/01 - 04/10/2023
16 | Shipped to | [5 empty] | K0 CAL: 08/21 - 08/23/2023
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2004
GDF Production Traveler
Serial Number: |  | 2004
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | sn
2 | Visual Inspection |  | 2022-07-05 | rb
3 | Coverplate |  | 2022-09-14 | vs
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.298 | 0.096 | 0.069 | 0.018 | 0.302
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 12/7/22 PW |  | Pass He Leak test
8 | VL pellet resistance |  | 30 KΩ | RB | ~90 KΩ
9 | ISFET diode measurement |  | 0.657/0.656 V | NL | 0.608/0.608 V |  | Ti No Shoulder Coverplate
10 | Counter electrode resistance |  | 0.18 Ω | RB | 0.2 Ω
11 | Seawater Cond. (num. days) | [5 empty] | CTD: 18986
12 | K2 cal (file name) | [5 empty] | MSC: 1106
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0 | [5 empty] | PT CAL: 07/09 - 07/14/2023
16 | Shipped to | [5 empty] | K0 CAL: 08/22 - 08/24/2023
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2005
GDF Production Traveler
Serial Number: |  | 2005
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.299 | 0.096 | 0.06 | 0.018 | 0.302
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 12/7/2022 PW |  | Pass He leak test
8 | VL pellet resistance |  | 29 KΩ | RB | 110 KΩ
9 | ISFET diode measurement |  | 0.659/0.658 V | NL | 0.608/0.608 V |  | CTD: 18987
10 | Counter electrode resistance |  | 0.19 Ω | RB | 0.3 Ω |  | MSC: 1098
11 | Seawater Cond. (num. days)
12 | K2 cal (file name) | [5 empty] | PT CAL: 05/25 - 05/29/2023
13 | f(p) cal (file name) | [5 empty] | KO CAL: 07/17 - 07/19/2023
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2007
GDF Production Traveler
Serial Number: |  | 2007
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-04-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.299 | 0.096 | 0.06 | 0.018 | 0.301
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 12/7/2022 PW |  | Pass He leak test
8 | VL pellet resistance |  | 31 KΩ | RB | 80 KΩ
9 | ISFET diode measurement |  | 0.651/0.651 V | NL | 0.606/0.606 V |  | Ti No Shoulder Coverplate
10 | Counter electrode resistance |  | 0.19 Ω | RB | 0.12 Ω
11 | Seawater Cond. (num. days) | [5 empty] | CTS: 18601
12 | K2 cal (file name) | [5 empty] | MSC: 1101
13 | f(p) cal (file name)
14 | Install on endcap | [5 empty] | PTC CAL: 04/13 - 04/18/2023
15 | Test / Calibration K0 | [5 empty] | K0 CAL: 7/17 - 07/19/2023
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2008
GDF Production Traveler
Serial Number: |  | 2008
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.035 | 0.298 | 0.096 | 0.06 | 0.018 | 0.301
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 12/7/2022 PW |  | Pass He leak test
8 | VL pellet resistance |  | 30 KΩ | RB | 85 KΩ
9 | ISFET diode measurement |  | 0.650/0.649 V | NL | 0.603/0.603 |  | Ti no shoulder coverplate
10 | Counter electrode resistance |  | 0.19 Ω | RB | 0.2 Ω
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2009
GDF Production Traveler
Serial Number: |  | 2009
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.036 | 0.299 | 0.096 | 0.061 | 0.019 | 0.301
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 12/7/2022 PW |  | Pass He Leak test
8 | VL pellet resistance |  | 24 KΩ | RB | 95 KΩ
9 | ISFET diode measurement |  | 0.658/0.656 V | NL | 0.607/0.607
10 | Counter electrode resistance |  | 0.19 Ω | RB | 0.3 Ω
11 | Seawater Cond. (num. days) | [5 empty] | CTD: 18549
12 | K2 cal (file name) | [5 empty] | MSC: 107F6
13 | f(p) cal (file name)
14 | Install on endcap | [5 empty] | PT CAL: 12/11 - 12/16/2022
15 | Test / Calibration K0 | [5 empty] | K0 CAL: 05/01 - 05/03/2023
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2010
GDF Production Traveler
Serial Number: |  | 2010
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2022-07-01 | SN
2 | Visual Inspection |  | 2022-07-05 | RB
3 | Coverplate |  | 2022-09-14 | VS
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.025 | 0.036 | 0.297 | 0.096 | 0.06 | 0.018 | 0.301
4 | Gasket thickness |  | 0.033
| Gasket material |  | Viton
Gasket compression
5 | Build sensor |  | 2022-11-17 | BN
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 |  |  |  | 12/7/2022 PW |  | 12/13/2022 Passed initial He leak test but then flooded during PTC run - isfet chip was likely cracked after He leak test. To be rebuilt and reused…
8 | VL pellet resistance |  | 29 KΩ | RB | 85 KΩ
9 | ISFET diode measurement |  | 0.655/0.657 V | NL | 0.604/0.604 V
10 | Counter electrode resistance |  | 0.19 Ω | RB | 0.3 Ω |  | Feb 2024: sensor rebuilt with oil fill port to test pressure balance GDF design. Saunders gasket and tall shoulder PEEK coverplate.
11 | Seawater Cond. (num. days) | [5 empty] | Feb 28 2024: Built with oil port, and capped off with blanking plate for pressure testing. Sensor flooded during first pressure cycle.
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2011
GDF Production Traveler
Serial Number: |  | GDF 2011
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection |  | 2023-07-13 | RB
3 | Coverplate |  | 2023-07-13 | RB
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.036 | 0.298 | 0.098 | 0.061 | 0.0183 | 0.3015
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | He leak test pass
8 | VL pellet resistance |  | 200 K ¦ 16 K | RB | 7/13/2023
9 | ISFET diode measurement |  | 0.64/0.64 V | RB | 7/13/2023
10 | Counter electrode resistance |  | 0.44 Ω | RB | 7/13/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2012
GDF Production Traveler
Serial Number: |  | 2012
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection |  | 2023-07-13 | RB
3 | Coverplate |  | 2023-07-13 | RB
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.036 | 0.298 | 0.098 | 0.0609 | 0.0183 | 0.3015
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | He leak test pass
8 | VL pellet resistance |  | 200 K ¦ 20 KΩ | RB | 7/17/2023
9 | ISFET diode measurement |  | 0.64/0.64 V | RB | 7/14/2023
10 | Counter electrode resistance |  | 0.45 Ω | RB | 7/17/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2013
GDF Production Traveler
Serial Number: |  | 2013
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection |  |  | RB
3 | Coverplate |  | 2023-07-26 | RB
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0351 | 0.2973 | 0.0975 | 0.0604 | 0.0179 | 0.0604
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (8/1/2023)
8 | VL pellet resistance |  | 25 KΩ | RB | 7/31/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | RB | 7/27/2023
10 | Counter electrode resistance |  | 0.45 Ω | RB | 7/31/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2014
GDF Production Traveler
Serial Number: |  | 2014
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0241 | 0.0356 | 0.2972 | 0.098 | 0.0608 | 0.018 | 0.301
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (8/1/2023)
8 | VL pellet resistance |  | 0.23 KΩ | RB | 7/31/2023
9 | ISFET diode measurement |  | 0.64/0.64 V | RB | 7/27/2023
10 | Counter electrode resistance |  | 0.49 Ω | RB | 7/31/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2015
GDF Production Traveler
Serial Number: |  | 2015
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0353 | 0.2974 | 0.0976 | 0.0608 | 0.018 | 0.301
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (8/1/2023)
8 | VL pellet resistance |  | 0.24 kΩ | RB | 7/31/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | RB | 7/27/2023
10 | Counter electrode resistance |  | 0.30 Ω | RB | 7/31/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2016
GDF Production Traveler
Serial Number: |  | 2016
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0242 | 0.0354 | 0.2972 | 0.098 | 0.0605 | 0.0178 | 0.301
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (8/1/2023)
8 | VL pellet resistance |  | 0.22 KΩ | RB | 7/31/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | RB | 7/31/2023
10 | Counter electrode resistance |  | 0.28 Ω | RB | 7/31/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2017
GDF Production Traveler
Serial Number: |  | 2017
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0351 | 0.2974 | 0.0976 | 0.0607 | 0.0186 | 0.3027
4 | Gasket thickness
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (8/1/2023)
8 | VL pellet resistance |  | 0.23 KΩ | RB | 7/31/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | RB | 7/27/2023
10 | Counter electrode resistance |  | 0.43 Ω | RB | 7/31/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2018
GDF Production Traveler
Serial Number: |  | 2018
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0353 | 0.2973 | 0.0975 | 0.061 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (8/10/2023)
8 | VL pellet resistance |  | 21-22 KΩ | BN | 8/10/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | BN | 8/10/2023
10 | Counter electrode resistance |  | 0.54-0.55 Ω | BN | 8/10/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
GDFs 2018 - 2022 paper sign off (on who did meaasurments/ tests) is hard to decipher
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet) | [6 empty] | VL pellet resitences reported are missing decimel place (e.g. should be 0.2 KΩ, or 200 Ω)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2019
GDF Production Traveler
Serial Number: |  | 2019
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2973 | 0.0975 | 0.061 | 0.0179 | 0.3011
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He Leak test (8/10/2023)
8 | VL pellet resistance |  | 22-23 KΩ | BN | 2023-08-10
9 | ISFET diode measurement |  | 65/65 V | BN | 2023-08-10
10 | Counter electrode resistance |  | 0.34-0.35 Ω | BN | 2023-08-10
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
GDFs 2018 - 2022 paper sign off (on who did meaasurments/ tests) is hard to decipher
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet) | [6 empty] | VL pellet resitences reported are missing decimel place (e.g. should be 0.2 KΩ, or 200 Ω)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2020
GDF Production Traveler
Serial Number: |  | 2020
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0355 | 0.2972 | 0.0975 | 0.0605 | 0.0186 | 0.3037
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | He Leak Test Pass (8/10/23)
8 | VL pellet resistance |  | 18-19 KΩ | BN | 8/10/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | BN | 8/10/2023
10 | Counter electrode resistance |  | 0.30-0.31 Ω | BN | 8/10/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
GDFs 2018 - 2022 paper sign off (on who did meaasurments/ tests) is hard to decipher
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet) | [6 empty] | VL pellet resitences reported are missing decimel place (e.g. should be 0.2 KΩ, or 200 Ω)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2021
GDF Production Traveler
Serial Number: |  | 2021
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.035 | 0.2973 | 0.0974 | 0.0607 | 0.018 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | He Leak test Pass (8/10/23)
8 | VL pellet resistance |  | 21-22 KΩ | BN | 2023-08-10
9 | ISFET diode measurement |  | 0.65/0.65 V | BN | 2023-08-10
10 | Counter electrode resistance |  | 0.26-0.27 Ω | BN | 2023-08-10
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
GDFs 2018 - 2022 paper sign off (on who did meaasurments/ tests) is hard to decipher
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet) | [6 empty] | VL pellet resitences reported are missing decimel place (e.g. should be 0.2 KΩ, or 200 Ω)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2022
GDF Production Traveler
Serial Number: |  | 2022
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0351 | 0.2973 | 0.0975 | 0.0607 | 0.018 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | He Leak test Pass (8/10/23)
8 | VL pellet resistance |  | 22-23 KΩ | BN | 8/10/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | BN | 8/10/2023
10 | Counter electrode resistance |  | 0.33 -0.34 Ω | BN | 8/10/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
GDFs 2018 - 2022 paper sign off (on who did meaasurments/ tests) is hard to decipher
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet) | [6 empty] | VL pellet resitences reported are missing decimel place (e.g. should be 0.2 KΩ, or 200 Ω)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2023
GDF Production Traveler
Serial Number: |  | 2023
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.03535 | 0.2973 | 0.0975 | 0.0604 | 0.018 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He Leak test
8 | VL pellet resistance |  | 99 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.64 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.36 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2024
GDF Production Traveler
Serial Number: |  | 2024
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2972 | 0.0974 | 0.0605 | 0.0179 | 0.3012
4 | Gasket thickness |  | 0.0325
Gasket material
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 68 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.32 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2025
GDF Production Traveler
Serial Number: |  | 2025
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0351 | 0.2972 | 0.0974 | 0.0603 | 0.018 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 63 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.39 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2026
GDF Production Traveler
Serial Number: |  | 2026
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0352 | 0.2972 | 0.0974 | 0.0603 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 89 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.29 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2027
GDF Production Traveler
Serial Number: |  | 2027
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2973 | 0.0974 | 0.0603 | 0.0179 | 0.3011
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 83 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.40 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2028
GDF Production Traveler
Serial Number: |  | 2028
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0351 | 0.2973 | 0.0975 | 0.0605 | 0.0179 | 0.3012
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 130 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.29 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2029
GDF Production Traveler
Serial Number: |  | 2029
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2973 | 0.0974 | 0.0603 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 110 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.27 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2030
GDF Production Traveler
Serial Number: |  | 2030
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0351 | 0.2973 | 0.0975 | 0.0605 | 0.0179 | 0.3012
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 62 kΩ | RB | 9/19/23 9/19/23 9/19/23
9 | ISFET diode measurement |  | 0.65 / 0.65 V | RB | 9/19/23 9/19/23 9/19/23
10 | Counter electrode resistance |  | 0.30 Ω | RB | 9/19/23 9/19/23 9/19/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2031
GDF Production Traveler
Serial Number: |  | 2031
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0357 | 0.297 | 0.097 | 0.0607 | 0.0181 | 0.301
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 21-22 KΩ | 12/6/23
9 | ISFET diode measurement |  | 0.65/0.65 V | 12/6/23
10 | Counter electrode resistance |  | 0.26 Ω | 12/6/23
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2032
GDF Production Traveler
Serial Number: |  | 2032
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0357 | 0.297 | 0.0976 | 0.0604 | 0.0182 | 0.3015
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 27 KΩ | AE
9 | ISFET diode measurement |  | 0.65/0.65 V | AE
10 | Counter electrode resistance |  | 0.30 Ω | AE
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2033
GDF Production Traveler
Serial Number: |  | 2033
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0351 | 0.2972 | 0.0974 | 0.0608 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 21-23 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/ 0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.26 - 0.27 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2034
GDF Production Traveler
Serial Number: |  | 2034
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.036 | 0.2972 | 0.0973 | 0.0607 | 0.018 | 0.301
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 12/6/2023
9 | ISFET diode measurement |  | 0.65/0.65 V | 12/6/2023
10 | Counter electrode resistance |  | 0.29 Ω | 12/6/2023
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2035
GDF Production Traveler
Serial Number: |  | 2035
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.035 | 0.2973 | 0.0975 | 0.0608 | 0.018 | 0.3016
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.31 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2036
GDF Production Traveler
Serial Number: |  | 2036
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2974 | 0.0975 | 0.0606 | 0.0186 | 0.3029
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.38 Ω | 2023-12-07
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2037
GDF Production Traveler
Serial Number: |  | 2037
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0353 | 0.2973 | 0.0975 | 0.3023 | 0.019 | 0.061
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-06 |  |  | 4/4/2024: sensor failed during day three of PTC calibration. Moisture found inside of pcb pocket.
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.30 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2038
GDF Production Traveler
Serial Number: |  | 2038
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0352 | 0.2973 | 0.0975 | 0.0605 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.27 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2039
GDF Production Traveler
Serial Number: |  | 2039
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0351 | 0.2973 | 0.0975 | 0.0608 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 21-22 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.31 Ω | incomplete date
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2040
GDF Production Traveler
Serial Number: |  | 2040
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0355 | 0.2973 | 0.0976 | 0.0609 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 20-21 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.27 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2041
GDF Production Traveler
Serial Number: |  | 2041
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0353 | 0.2973 | 0.0975 | 0.0604 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 20-21 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.29 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2042
GDF Production Traveler
Serial Number: |  | 2042
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2972 | 0.0974 | 0.0604 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.29 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2043
GDF Production Traveler
Serial Number: |  | 2043
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0351 | 0.2972 | 0.0974 | 0.0603 | 0.018 | 0.3011
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.33 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2044
GDF Production Traveler
Serial Number: |  | 2044
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0352 | 0.2972 | 0.0974 | 0.0601 | 0.018 | 0.3011
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 23-24 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.31 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2045
GDF Production Traveler
Serial Number: |  | 2045
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.035 | 0.2973 | 0.0974 | 0.0605 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test
8 | VL pellet resistance |  | 20-21 KΩ | 2023-12-06
9 | ISFET diode measurement |  | 0.65/0.65 V | 2023-12-06
10 | Counter electrode resistance |  | 0.33 Ω | 2023-12-06
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2046
GDF Production Traveler
Serial Number: |  | 2046
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0241 | 0.0355 | 0.2972 | 0.0978 | 0.0603 | 0.0179 | 0.3011
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test 12/13/2023
8 | VL pellet resistance |  | 21-22 KΩ | 2023-12-13 | PDV
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 2023-12-13 | PDV
10 | Counter electrode resistance |  | 0.33 Ω | 2023-12-15 | PDV
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2047
GDF Production Traveler
Serial Number: |  | 2047
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0355 | 0.2976 | 0.098 | 0.0601 | 0.0185 | 0.3013
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (1/16/2024)
8 | VL pellet resistance |  | 21-22 kΩ | 2024-01-16 | PDV
9 | ISFET diode measurement |  | 65 / 65 V | 2024-01-16 | PDV
10 | Counter electrode resistance |  | 0.33 Ω | 2024-01-17 | PDV
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2048
GDF Production Traveler
Serial Number: |  | 2048
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0354 | 0.2973 | 0.0976 | 0.0694 | 0.0178 | 0.3011
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (12/13/2023)
8 | VL pellet resistance |  | 22-23 KΩ | 2023-12-13 | PDV
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 2023-12-13 | PDV
10 | Counter electrode resistance |  | 0.35 Ω | 2023-12-15 | PDV
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number
12/13/23

## GDF 2049
GDF Production Traveler
Serial Number: |  | 2049
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0359 | 0.2972 | 0.0976 | 0.0604 | 0.0179 | 0.301
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (12/13/2023)
8 | VL pellet resistance |  | 21-22 KΩ | 2023-12-13 | PDV
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 2023-12-13 | PDV
10 | Counter electrode resistance |  | 0.41 Ω | 2023-12-15 | PDV
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number
PDV

## GDF 2050
GDF Production Traveler
Serial Number: |  | 2050
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0353 | 0.2973 | 0.0979 | 0.0604 | 0.018 | 0.3012
4 | Gasket thickness |  | 0.03 | CSM
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (12/13/2023)
8 | VL pellet resistance |  | 23-24 KΩ | 2023-12-13 | PDV
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 2023-12-13 | PDV
10 | Counter electrode resistance |  | 0.36 Ω | 2023-12-15 | PDV
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2051
GDF Production Traveler
Serial Number: |  | 2051
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0366 | 0.2974 | 0.098 | 0.0599 | 0.0176 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 22 kΩ | 6/13/24 |  |  | Pass He leak test (06/12/24)
9 | ISFET diode measurement |  | 0.65/0.65 V | 6/13/24
10 | Counter electrode resistance (200Ω) |  | 0.41-0.42 | 6/13/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2052
GDF Production Traveler
Serial Number: |  | 2052
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.03679 | 0.2973 | 0.0978 | 0.0602 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 22 - 23 k | 6/13/24 |  |  | Pass He Leak test 6/12/24
9 | ISFET diode measurement |  | 0.65/0.65 V | 6/13/24
10 | Counter electrode resistance |  | 0.43 - 0.44 Ω | 6/13/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2053
GDF Production Traveler
Serial Number: |  | 2053
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0238 | 0.0366 | 0.2973 | 0.0977 | 0.0602 | 0.0178 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 21 - 22 k | 6/13/24 |  |  | Pass He leak test (06/12/24)
9 | ISFET diode measurement |  | 0.65 /0.65 V | 6/13/24
10 | Counter electrode resistance |  | 0.40 -0.41 | 6/13/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2054
GDF Production Traveler
Serial Number: |  | 2054
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0238 | 0.0368 | 0.2973 | 0.0979 | 0.06 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 9 - 10 kΩ | 6/13/24 |  |  | Pass He leak test (06/12/24)
9 | ISFET diode measurement |  | 0.65/0.65 V | 6/13/24
10 | Counter electrode resistance |  | 0.40-0.41 | 6/13/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2055
GDF Production Traveler
Serial Number: |  | 2055
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0363 | 0.2979 | 0.0981 | 0.0603 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 24 - 25 | 6/13/24 |  |  | Pass He leak test (06/12/24)
9 | ISFET diode measurement |  | 0.65/0.65 V | 6/13/24
10 | Counter electrode resistance |  | 0.37 -0.38 | 6/13/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2056
GDF Production Traveler
Serial Number: |  | 2056
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0236 | 0.0353 | 0.2974 | 0.0973 | 0.06 | 0.178 | 0.3016
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 23-24 kΩ | 6/18/24 |  |  | Pass He leak test (6/21/24)
9 | ISFET diode measurement |  | 65 / 65 | 6/18/24
10 | Counter electrode resistance |  | 0.41 / 0.42 | 6/18/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2057
GDF Production Traveler
Serial Number: |  | 2057
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0239 | 0.0368 | 0.2972 | 0.0978 | 0.0601 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (6/20/24)
8 | VL pellet resistance |  | 24 - 25 KΩ | 6/18/24
9 | ISFET diode measurement |  | 65 / 65 | 6/18/24
10 | Counter electrode resistance |  | 0.41 / 0.42 | 6/18/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2058
GDF Production Traveler
Serial Number: |  | 2058
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0238 | 0.0368 | 0.2973 | 0.0977 | 0.0602 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (06/20/24)
8 | VL pellet resistance |  | 11 - 12 KΩ | 6/18/24
9 | ISFET diode measurement |  | 65 / 65 | 6/18/24
10 | Counter electrode resistance |  | 0.51 / 0.52 | 6/18/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2059
GDF Production Traveler
Serial Number: |  | 2059
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0239 | 0.0366 | 0.2972 | 0.0977 | 0.0596 | 0.0183 | 0.3013
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He leak test (06/20/24)
8 | VL pellet resistance |  | 25 - 26 KΩ | 6/18/24
9 | ISFET diode measurement |  | 65 / 65 | 6/18/24
10 | Counter electrode resistance |  | 0.55 / 0.56 | 6/18/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2060
GDF Production Traveler
Serial Number: |  | 2060
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0237 | 0.0369 | 0.2974 | 0.098 | 0.0601 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.03
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3 | [5 empty] | Pass He Leak test (06/20/24)
8 | VL pellet resistance |  | 30 - 31 KΩ | 6/18/24
9 | ISFET diode measurement |  | 65 / 65 | 6/18/24
10 | Counter electrode resistance |  | 55 / 56 | 6/18/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2061
GDF Production Traveler
Serial Number: |  | 2061
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0363 | 0.2974 | 0.0978 | 0.0604 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 20-21 kΩ | 7/8/24 |  |  | Pass He leak test (07/10/24)
9 | ISFET diode measurement |  | 0.65/0.65 V | 7/8/24
10 | Counter electrode resistance |  | 46 ± 47 | 7/10/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2062
GDF Production Traveler
Serial Number: |  | 2062
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0241 | 0.0353 | 0.2975 | 0.098 | 0.0596 | 0.0181 | 0.3011
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 23 - 24 kΩ | 7/8/24 |  |  | Pass He leak test (07/10/24)
9 | ISFET diode measurement |  | 0.65/0.65 | 7/8/24
10 | Counter electrode resistance |  | 48 ± 49 | 7/10/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF2063
GDF Production Traveler
Serial Number: |  | 2063
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0368 | 0.2973 | 0.098 | 0.0603 | 0.0178 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 23 - 24 kΩ | 7/8/24 |  |  | Pass He leak test (07/10/24)
9 | ISFET diode measurement |  | 0.65/0.65 V | 7/8/24
10 | Counter electrode resistance |  | 50 ± 51 | 7/10/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2064
GDF Production Traveler
Serial Number: |  | 2065
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0241 | 0.0368 | 0.2975 | 0.0979 | 0.0596 | 0.0181 | 0.3011
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance | [5 empty] | Pass He leak test (07/10/24)
9 | ISFET diode measurement |  | 24 - 25 kΩ | 7/8/24
10 | Counter electrode resistance |  | 0.65/0.65 V | 7/8/24
11 | Seawater Cond. (num. days) |  | 51 ± 52 | 7/10/24
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2065
GDF Production Traveler
Serial Number: |  | 2065
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0238 | 0.0366 | 0.298 | 0.0981 | 0.0603 | 0.0177 | 0.3017
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 23 ± 24 kΩ | 7/8/24 |  |  | Pass He leak test (07/10/24)
9 | ISFET diode measurement |  | 0.65/0.65 V | 7/8/24
10 | Counter electrode resistance |  | 51 ± 52 | 7/10/24
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2066
GDF Production Traveler
Serial Number: |  | 2066
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0365 | 0.2974 | 0.0979 | 0.0596 | 0.0184 | 0.3012
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 0.22 - 0.21 kΩ | 7/23/2024 |  |  | Pass He leak test (07/22/24)
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 7/23/2024
10 | Counter electrode resistance |  | 45 - 46 | 7/23/2024
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2067
GDF Production Traveler
Serial Number: |  | 2067
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0239 | 0.0367 | 0.2973 | 0.0982 | 0.0601 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 0.20 -0.21 kΩ | 7/23/2024 |  |  | Pass He leak test (07/23/24)
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 7/23/2024
10 | Counter electrode resistance |  | 0.45 - 0.46 | 7/23/2024
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2068
GDF Production Traveler
Serial Number: |  | 2068
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0241 | 0.0366 | 0.2975 | 0.0979 | 0.0603 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 0.26-0.27 kΩ | 7/23/2024 |  |  | Pass He leak test (07/22/24)
9 | ISFET diode measurement |  | 0.65 /0.65 V | 7/23/2024
10 | Counter electrode resistance |  | 43 -44 | 7/23/2024
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2069
GDF Production Traveler
Serial Number: |  | 2069
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.024 | 0.0367 | 0.2974 | 0.0982 | 0.06 | 0.0177 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 0.25 - 0.26 kΩ | no date |  |  | Pass He leak test (07/22/24)
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 7/23/2024
10 | Counter electrode resistance |  | 46 - 47 | no date
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2070
GDF Production Traveler
Serial Number: |  | 2070
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0241 | 0.0366 | 0.2979 | 0.098 | 0.0636 | 0.1777 | 0.3016
4 | Gasket thickness |  | 0.0325
| Gasket material |  | Viton
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance |  | 0.25 - 0.24 kΩ | no date |  |  | Pass He leak test (07/23/24)
9 | ISFET diode measurement |  | 0.65 / 0.65 V | 7/23/2024
10 | Counter electrode resistance |  | 44 - 45 | no date
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF 2071
GDF Production Traveler
Serial Number:
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
4 | Gasket thickness
Gasket material
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance
9 | ISFET diode measurement
10 | Counter electrode resistance
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number

## GDF template
GDF Production Traveler
Serial Number:
Version2: |  | rev A
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
4 | Gasket thickness
Gasket material
Gasket compression
5 | Build sensor
6 | Photos front and back3 | [5 empty] | comments:
7 | Other photos3
8 | VL pellet resistance 200K
9 | ISFET diode measurement
10 | Counter electrode resistance 200Ω
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)
3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos,
YYYY-xxx is the serial number