## GDF15_Calibration_20211227_all
PH CALIBRATION FILE GENERATED ON: 12/27/2021 14:58
** START CONFIG FILE COPY BLOCK **
Durafet SN = 15, APEX# XXXX
8, number of calibration coefficients
NO K0 DATA WAS PEOCESSED YET OR LOADED INTO GUI!!!!
-1.15296e-14,1.68484e-11,-7.17136e-08,-1.02883e-03, =k2(fP)*T
-7.12904e-06, =k3*P
2.97914e-10, =k4*P^3
1.57641e-11, =k5*P^4
-2.14683e-14, =k6*P^5
1.06335e-17, =k7*P^6
-1.79920e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2(T) | k0(HCl) | R^2 | k6 | k5 | k4 | k3 | k2 | k1 | k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
12/18-12/24/2021 |  |  | NaN | -0.0010828 | -1.4355 | 0.9998 | -1.7992e-21 | 1.0633e-17 | -2.1468e-14 | 1.5764e-11 | 2.9791e-10 | -7.129e-06 | -1.4352 | 0.9832
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 15
pH stem mfg: | MBARI
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF15_cal_every2C.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2021-12-18 01:07 | 2021-12-24 06:01
Data subset: | ALL
Vrs drift rate: | NO DATA | mV/day
Vk drift rate: | NO DATA | mV/day
Avg std Vk | 190.3806 | uVolts
Average k2: | -0.0010828 | V/degC
Std dev k2: | 5.2845e-05 | V/degC
k2 f(P) c0 | -0.0010288
k2 f(P) c1 | -7.1714e-08
k2 f(P) c2 | 1.6848e-11
k2 f(P) c3 | -1.153e-14
k0 HCL: | -1.4355 | Volts
Avg Ik: | 20.32 | nano amps
Std dev Ik: | 0.64 | nano amps
Avg Ib: | 3.39 | nano amps
Std dev Ib: | 36.11 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
21 | -0.0010241 | 1.6803e-07 | -1.4355 | 2.0938e-06 | 0.9998 | 8070 | 1.9 | 19.7
37.8 | -0.0010305 | 1.2042e-06 | -1.4355 | 1.5262e-05 | 0.9997 | 210 | 1.9 | 19.7
62 | -0.0010316 | 1.321e-06 | -1.4357 | 1.6306e-05 | 0.9995 | 285 | 1.9 | 19.7
88.3 | -0.0010387 | 1.8387e-06 | -1.4358 | 2.318e-05 | 0.9992 | 245 | 1.9 | 19.7
124.7 | -0.0010428 | 2.0198e-06 | -1.4361 | 2.5106e-05 | 0.9983 | 463 | 1.9 | 19.7
173.6 | -0.0010432 | 2.5844e-06 | -1.4364 | 3.2311e-05 | 0.9972 | 466 | 1.9 | 19.7
249 | -0.0010454 | 2.5679e-06 | -1.4368 | 3.1915e-05 | 0.9945 | 911 | 1.9 | 19.7
349.1 | -0.0010529 | 2.7993e-06 | -1.4373 | 3.4651e-05 | 0.9937 | 895 | 1.9 | 19.8
448.5 | -0.0010582 | 2.8186e-06 | -1.4376 | 3.4609e-05 | 0.9937 | 889 | 1.9 | 19.8
548.6 | -0.0010612 | 2.6575e-06 | -1.438 | 3.289e-05 | 0.9945 | 889 | 1.9 | 19.8
649.4 | -0.0010741 | 2.5796e-06 | -1.4381 | 3.2484e-05 | 0.9952 | 840 | 1.9 | 19.8
749.9 | -0.0010785 | 2.4936e-06 | -1.4384 | 3.0935e-05 | 0.9954 | 858 | 1.9 | 19.8
899.5 | -0.001088 | 2.0061e-06 | -1.4387 | 2.4752e-05 | 0.9941 | 1734 | 1.9 | 19.8
1099.5 | -0.0011003 | 1.9221e-06 | -1.4392 | 2.3688e-05 | 0.9948 | 1709 | 1.9 | 19.8
1299.5 | -0.0011156 | 1.3813e-06 | -1.4399 | 1.7046e-05 | 0.9974 | 1708 | 1.9 | 19.8
1499.6 | -0.0011383 | 1.6832e-06 | -1.4405 | 2.0727e-05 | 0.9963 | 1688 | 1.9 | 19.8
1699.3 | -0.0011635 | 1.7563e-06 | -1.4407 | 2.174e-05 | 0.9962 | 1678 | 1.9 | 19.8
1898.2 | -0.0011833 | 1.3237e-06 | -1.4407 | 1.6472e-05 | 0.998 | 1615 | 1.9 | 19.8
2059.1 | -0.0012036 | 5.1239e-07 | -1.4405 | 6.4571e-06 | 0.9995 | 2523 | 2 | 19.8 |  |  | dVrs = | 0.0052
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 7.6886e-08
f(P) R*R: | 0.9832
f(P) c0 | -1.4352
f(P) c1 | -7.129e-06
f(P) c2 | 2.9791e-10
f(P) c3 | 1.5764e-11
f(P) c4 | -2.1468e-14
f(P) c5 | 1.0634e-17
f(P) c6 | -1.7992e-21
NO KO DATA FILE LOADED INTO CALIBRATION GUI!!!:
Has the k0 calibration be completed yet?:
Is the MSC # correct? Check MSC / SN match up:

## GDF15_Ids 70 Calibration_202201
PH CALIBRATION FILE GENERATED ON: 01/06/2022 15:53
** START CONFIG FILE COPY BLOCK **
Durafet SN = 15, APEX# XXXX
8, number of calibration coefficients
NO K0 DATA WAS PEOCESSED YET OR LOADED INTO GUI!!!!
1.75568e-11,-5.71035e-08,4.45698e-06, =k2(fP)*T
-1.90625e-06, =k3*P
-3.40888e-09, =k4*P^3
1.54471e-11, =k5*P^4
-1.81354e-14, =k6*P^5
8.65497e-18, =k7*P^6
-1.45816e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2(T) | k0(HCl) | R^2 | k6 | k5 | k4 | k3 | k2 | k1 | k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
12/31-01/05/2022 |  |  | NaN | -2.0471e-05 | -1.0946 | 0 | -1.4582e-21 | 8.655e-18 | -1.8135e-14 | 1.5447e-11 | -3.4089e-09 | -1.9062e-06 | -1.0946 | 0.6011
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 15
pH stem mfg: | MBARI
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF15_cal_70Ids.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2021-12-31 07:10 | 2022-01-05 15:30
Data subset: | ALL
Vrs drift rate: | NO DATA | mV/day
Vk drift rate: | NO DATA | mV/day
Avg std Vk | 181.1775 | uVolts
Average k2: | -2.0471e-05 | V/degC
Std dev k2: | 1.7482e-05 | V/degC
k2 f(P) c0 | 4.457e-06
k2 f(P) c1 | -5.7104e-08
k2 f(P) c2 | 1.7557e-11
k0 HCL: | -1.0946 | Volts
Avg Ik: | 20.66 | nano amps
Std dev Ik: | 0.65 | nano amps
Avg Ib: | -0.59 | nano amps
Std dev Ib: | 2.86 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
21 | 3.4336e-07 | 6.5198e-07 | -1.0946 | 6.6117e-06 | 0 | 5700 | 1.9 | 17.8
37.9 | 9.8157e-08 | 4.3495e-06 | -1.0947 | 4.0339e-05 | 0 | 145 | 2 | 17.7
62.3 | 1.68e-06 | 3.4284e-06 | -1.0947 | 3.1936e-05 | 0.0012 | 209 | 2 | 17.7
88.3 | -1.5788e-06 | 4.0722e-06 | -1.0948 | 3.8267e-05 | 0.001 | 159 | 2 | 17.7
124.1 | -6.4192e-06 | 3.1346e-06 | -1.0948 | 2.93e-05 | 0.0127 | 327 | 2 | 17.8
174.1 | -6.1661e-06 | 3.3813e-06 | -1.095 | 3.1269e-05 | 0.0099 | 335 | 2 | 17.8
249.4 | -1.0437e-05 | 2.9703e-06 | -1.0951 | 2.8125e-05 | 0.0195 | 624 | 2 | 17.8
348.7 | -1.2202e-05 | 3.5283e-06 | -1.0953 | 3.3056e-05 | 0.0194 | 608 | 1.9 | 17.8
448.2 | -1.0211e-05 | 4.0803e-06 | -1.0954 | 3.8866e-05 | 0.0107 | 581 | 1.9 | 17.8
548.6 | -1.1022e-05 | 4.276e-06 | -1.0955 | 4.0827e-05 | 0.0114 | 579 | 1.9 | 17.8
649 | -2.6274e-05 | 4.1653e-06 | -1.0954 | 4.0048e-05 | 0.0675 | 552 | 1.9 | 17.8
749.9 | -2.3941e-05 | 4.1175e-06 | -1.0954 | 3.8317e-05 | 0.0553 | 579 | 1.9 | 17.8
899.4 | -3.0395e-05 | 2.7851e-06 | -1.0954 | 2.6112e-05 | 0.0931 | 1162 | 1.9 | 17.8
1099.2 | -3.8933e-05 | 2.2162e-06 | -1.0954 | 2.0685e-05 | 0.2122 | 1148 | 1.9 | 17.8
1299.5 | -4.3984e-05 | 1.7655e-06 | -1.0956 | 1.6503e-05 | 0.3515 | 1147 | 1.9 | 17.8
1499.4 | -4.8295e-05 | 2.356e-06 | -1.0957 | 2.1783e-05 | 0.273 | 1121 | 1.9 | 17.8
1698.6 | -5.1799e-05 | 2.5124e-06 | -1.0955 | 2.321e-05 | 0.2793 | 1099 | 1.9 | 15.8
1899 | -4.3029e-05 | 2.7885e-06 | -1.0952 | 2.6008e-05 | 0.1841 | 1057 | 2 | 15.8
2058.9 | -2.6381e-05 | 2.4099e-06 | -1.095 | 2.2606e-05 | 0.0692 | 1615 | 2 | 17.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 9.6318e-08
f(P) R*R: | 0.6011
f(P) c0 | -1.0946
f(P) c1 | -1.9063e-06
f(P) c2 | -3.4089e-09
f(P) c3 | 1.5447e-11
f(P) c4 | -1.8135e-14
f(P) c5 | 8.655e-18
f(P) c6 | -1.4582e-21
NO KO DATA FILE LOADED INTO CALIBRATION GUI!!!:
Has the k0 calibration be completed yet?:
Is the MSC # correct? Check MSC / SN match up:

## GDF16_Calibration_20220428
PH CALIBRATION FILE GENERATED ON: 04/28/2022 15:07
GDF 16
MSC 1019
APEX ID 21984
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21984, APEX# 16
8, number of calibration coefficients
-1.38489, =k0 Pon; -1.38343 =k0 Poff; -1.37242 =k0 HCl
6.86793e-12,-5.97821e-08,-1.04978e-03, =k2(fP)*T
-2.63893e-06, =k3*P
-1.19042e-08, =k4*P^3
2.99830e-11, =k5*P^4
-3.04942e-14, =k6*P^5
1.37777e-17, =k7*P^6
-2.26082e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
04/06-04/09/2022 | A |  | -0.33 |  |  |  |  | 6.8679e-12 | -5.9782e-08 | -0.0010498 | -1.3724 | 0.9998 |  |  | -2.2608e-21 | 1.3778e-17 | -3.0494e-14 | 2.9983e-11 | -1.1904e-08 | -2.6389e-06 | -1.3723 | 0.9928 | GDF |  | 16 | 04/25-04/27/2022 |  | -1.3834 | 0.00038109 | -1.3849 | 0.00040847
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 16
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF#16_newchip_chanA_220405.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-04-06 17:39 | 2022-04-09 04:20
Data subset: | ASCENDING
Vrs drift rate: | -0.33029 | mV/day
Vk drift rate: | -6.5191 | mV/day
Avg std Vk | 168.4373 | uVolts
Average k2: | -0.0010836 | V/degC
Std dev k2: | 3.2128e-05 | V/degC
k2 f(P) c0 | -0.0010498
k2 f(P) c1 | -5.9782e-08
k2 f(P) c2 | 6.8679e-12
k0 HCL: | -1.3724 | Volts
Avg Ik: | -5.54 | nano amps
Std dev Ik: | 3.55 | nano amps
Avg Ib: | 11.49 | nano amps
Std dev Ib: | 3.69 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.001041 | 3.4242e-06 | -1.3724 | 4.5925e-05 | 0.9998 | 16 | 2.3 | 19.8
14.6 | -0.0010458 | 3.1053e-06 | -1.3724 | 4.3241e-05 | 0.9998 | 25 | 5.1 | 19.8
24.1 | -0.0010472 | 3.3554e-06 | -1.3724 | 4.3558e-05 | 0.9998 | 24 | 2.3 | 19.8
40 | -0.0010534 | 3.4293e-06 | -1.3724 | 4.6722e-05 | 0.9998 | 22 | 2.3 | 19.8
64.2 | -0.001056 | 3.1325e-06 | -1.3725 | 4.5625e-05 | 0.9998 | 22 | 2.3 | 19.8
93.7 | -0.001059 | 3.3629e-06 | -1.3726 | 4.6789e-05 | 0.9998 | 18 | 5.1 | 19.8
127.7 | -0.0010574 | 4.3004e-06 | -1.3728 | 6.5031e-05 | 0.9997 | 22 | 2.1 | 19.8
172.3 | -0.0010666 | 3.8171e-06 | -1.3729 | 5.2485e-05 | 0.9997 | 23 | 5.1 | 19.8
245.7 | -0.0010644 | 4.0727e-06 | -1.3733 | 5.6895e-05 | 0.9993 | 50 | 5.1 | 19.8
349.4 | -0.0010754 | 4.1487e-06 | -1.3737 | 6.093e-05 | 0.9992 | 54 | 2.1 | 19.8
454.4 | -0.0010862 | 4.0226e-06 | -1.374 | 5.7528e-05 | 0.9995 | 42 | 5.1 | 19.8
552.3 | -0.0010866 | 3.9373e-06 | -1.3744 | 5.5738e-05 | 0.9995 | 43 | 5.1 | 19.8
653.9 | -0.0010833 | 3.6214e-06 | -1.3749 | 5.2337e-05 | 0.9996 | 41 | 5.1 | 19.8
749.8 | -0.0010847 | 4.5223e-06 | -1.3752 | 6.3856e-05 | 0.9993 | 43 | 2.2 | 19.8
894.6 | -0.0010921 | 4.3338e-06 | -1.3757 | 6.0278e-05 | 0.9986 | 90 | 2.2 | 19.8
1100.3 | -0.0010963 | 4.5984e-06 | -1.3763 | 6.8279e-05 | 0.9987 | 78 | 2.3 | 19.8
1297.4 | -0.0011023 | 5.6472e-06 | -1.377 | 8.5263e-05 | 0.9981 | 74 | 2.3 | 19.8
1499.5 | -0.001131 | 3.296e-06 | -1.3773 | 4.6675e-05 | 0.9993 | 83 | 2.3 | 19.8
1695.4 | -0.0011518 | 3.0687e-06 | -1.3772 | 4.4699e-05 | 0.9997 | 48 | 2.3 | 19.8
1904.8 | -0.0011463 | 9.8275e-06 | -1.3772 | 8.9562e-05 | 0.9989 | 17 | 2.3 | 19.8
2037.3 | -0.0011281 | 3.2025e-05 | -1.3771 | 0.00017145 | 0.996 | 7 | 2.3 | 10.1
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.1993e-08
f(P) R*R: | 0.9928
f(P) c0 | -1.3723
f(P) c1 | -2.6389e-06
f(P) c2 | -1.1904e-08
f(P) c3 | 2.9983e-11
f(P) c4 | -3.0494e-14
f(P) c5 | 1.3778e-17
f(P) c6 | -2.2608e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1019
k0 calibration window | 2022-04-25 18:30 | 2022-04-27 19:34
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 4
Ik outlier % | 0.27
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1019_DURA_2.dat | MSC1019_DURA.dat
median k0 | -1.3849 | -1.3835 | Volts
std k0 | 0.00040847 | 0.00038109 | Volts
Avg Vk | -0.8996 | -0.8996 | Volts
Avg std Vk | 60.1 | 60.2 | uVolts
median Ik: | -5.03 | -5.62 | nano amps
std Ik: | 4.86 | 4.99 | nano amps
median Ib: | 0.21 | 0.62 | nano amps
std Ib: | 5.95 | 5.95 | nano amps

## GDF19_Calibration_20220428
PH CALIBRATION FILE GENERATED ON: 04/28/2022 15:09
GDF 19
MSC 1018
APEX ID 21844
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21844, APEX# 19
8, number of calibration coefficients
-1.42549, =k0 Pon; -1.42511 =k0 Poff; -1.41687 =k0 HCl
-1.30827e-11,-3.73109e-08,-9.75665e-04, =k2(fP)*T
-5.42259e-06, =k3*P
3.99797e-09, =k4*P^3
-2.81865e-12, =k5*P^4
-1.95381e-15, =k6*P^5
2.49903e-18, =k7*P^6
-5.83543e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
03/24-03/27/2022 | A |  | -0.116 |  |  |  |  | -1.3083e-11 | -3.7311e-08 | -0.00097567 | -1.4169 | 0.9997 |  |  | -5.8354e-22 | 2.499e-18 | -1.9538e-15 | -2.8186e-12 | 3.998e-09 | -5.4226e-06 | -1.4167 | 0.9959 | GDF |  | 19 | 04/25-04/27/2022 |  | -1.4251 | 8.2989e-05 | -1.4255 | 7.2773e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 19
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | chanA_GDF19_ts_g_newchip_032322.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-03-24 18:36 | 2022-03-27 05:16
Data subset: | ASCENDING
Vrs drift rate: | -0.11584 | mV/day
Vk drift rate: | -7.6462 | mV/day
Avg std Vk | 166.6897 | uVolts
Average k2: | -0.0010119 | V/degC
Std dev k2: | 4.1334e-05 | V/degC
k2 f(P) c0 | -0.00097566
k2 f(P) c1 | -3.7311e-08
k2 f(P) c2 | -1.3083e-11
k0 HCL: | -1.4169 | Volts
Avg Ik: | -2.38 | nano amps
Std dev Ik: | 4.22 | nano amps
Avg Ib: | 12.7 | nano amps
Std dev Ib: | 4.09 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.00096209 | 4.6522e-06 | -1.4169 | 6.2353e-05 | 0.9997 | 16 | 2.3 | 19.8
14.4 | -0.00097817 | 4.8255e-06 | -1.4167 | 7.286e-05 | 0.9997 | 16 | 5.1 | 19.8
23.5 | -0.00097011 | 3.5576e-06 | -1.4168 | 4.9314e-05 | 0.9997 | 22 | 2.3 | 19.8
39.2 | -0.00097892 | 4.2184e-06 | -1.4169 | 5.7236e-05 | 0.9997 | 18 | 2.3 | 19.8
61.1 | -0.0009806 | 3.647e-06 | -1.417 | 5.0306e-05 | 0.9998 | 20 | 2.3 | 19.8
88.5 | -0.00098095 | 3.668e-06 | -1.4171 | 5.1262e-05 | 0.9997 | 22 | 2.3 | 19.8
125.5 | -0.00098832 | 3.9968e-06 | -1.4172 | 5.5893e-05 | 0.9996 | 25 | 2.3 | 19.8
172.7 | -0.00098492 | 5.0577e-06 | -1.4175 | 7.1574e-05 | 0.9994 | 26 | 2.3 | 19.8
250.2 | -0.00099017 | 3.5097e-06 | -1.4178 | 4.8686e-05 | 0.9993 | 57 | 2.3 | 19.8
351.6 | -0.00099638 | 3.6743e-06 | -1.4181 | 5.3684e-05 | 0.9995 | 41 | 2.3 | 19.8
448.3 | -0.00099944 | 4.1708e-06 | -1.4185 | 6.0503e-05 | 0.9993 | 41 | 5.2 | 19.8
549.1 | -0.0010023 | 4.6401e-06 | -1.4189 | 7.025e-05 | 0.9992 | 39 | 5.2 | 19.8
651.5 | -0.0010042 | 5.0647e-06 | -1.4194 | 7.287e-05 | 0.999 | 43 | 5.2 | 19.8
750.5 | -0.00099854 | 4.7268e-06 | -1.42 | 6.3448e-05 | 0.999 | 47 | 2.2 | 19.8
894.6 | -0.0010136 | 5.4817e-06 | -1.4206 | 7.5622e-05 | 0.9975 | 86 | 2.2 | 19.8
1092.8 | -0.0010272 | 5.6701e-06 | -1.4214 | 8.0983e-05 | 0.9976 | 82 | 2.2 | 19.8
1288.5 | -0.0010453 | 6.2313e-06 | -1.4222 | 9.3079e-05 | 0.9976 | 70 | 5.1 | 19.8
1497.2 | -0.0010655 | 4.4493e-06 | -1.4228 | 6.8506e-05 | 0.9989 | 67 | 2.2 | 19.8
1688 | -0.001083 | 3.3848e-06 | -1.423 | 4.9531e-05 | 0.9996 | 46 | 2.3 | 19.8
1924.3 | -0.0011108 | 8.7455e-06 | -1.4227 | 8.6385e-05 | 0.9992 | 15 | 2.3 | 19.8
2036.6 | -0.0010891 | 1.1986e-05 | -1.4228 | 6.4381e-05 | 0.9996 | 5 | 2.3 | 10.1
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.8813e-08
f(P) R*R: | 0.9959
f(P) c0 | -1.4167
f(P) c1 | -5.4226e-06
f(P) c2 | 3.998e-09
f(P) c3 | -2.8186e-12
f(P) c4 | -1.9538e-15
f(P) c5 | 2.499e-18
f(P) c6 | -5.8354e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1018
k0 calibration window | 2022-04-25 18:41 | 2022-04-27 19:34
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 5
Ik outlier % | 0.34
Ib outlier count | 1
Ib outlier % | 0.07
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1018_DURA_2.dat | MSC1018_DURA.dat
median k0 | -1.4255 | -1.4251 | Volts
std k0 | 7.2773e-05 | 8.2989e-05 | Volts
Avg Vk | -0.9406 | -0.9406 | Volts
Avg std Vk | 43.5 | 43.6 | uVolts
median Ik: | -7.34 | -7.49 | nano amps
std Ik: | 5.21 | 5.2 | nano amps
median Ib: | -0.83 | -0.23 | nano amps
std Ib: | 6.62 | 6.83 | nano amps

## GDF21_Calibration_20220428
PH CALIBRATION FILE GENERATED ON: 04/28/2022 15:11
GDF 21
MSC 1020
APEX ID 21142
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21142, APEX# 21
8, number of calibration coefficients
-1.43165, =k0 Pon; -1.43075 =k0 Poff; -1.42591 =k0 HCl
-2.26690e-11,-1.97716e-08,-1.01277e-03, =k2(fP)*T
-7.66925e-06, =k3*P
1.61288e-08, =k4*P^3
-3.37346e-11, =k5*P^4
3.41662e-14, =k6*P^5
-1.58710e-17, =k7*P^6
2.74234e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
03/17-03/20/2022 | B |  | -0.021 |  |  |  |  | -2.2669e-11 | -1.9772e-08 | -0.0010128 | -1.4259 | 0.9998 |  |  | 2.7423e-21 | -1.5871e-17 | 3.4166e-14 | -3.3735e-11 | 1.6129e-08 | -7.6692e-06 | -1.4257 | 0.9954 | GDF |  | 21 | 04/25-04/27/2022 |  | -1.4307 | 0.00013175 | -1.4317 | 0.00012796
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 21
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | chanB_GDF21_031622.csv
PT test channel : | B
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-03-17 17:51 | 2022-03-20 04:32
Data subset: | ASCENDING
Vrs drift rate: | -0.02051 | mV/day
Vk drift rate: | 3.8081 | mV/day
Avg std Vk | 51.2148 | uVolts
Average k2: | -0.0010458 | V/degC
Std dev k2: | 4.2526e-05 | V/degC
k2 f(P) c0 | -0.0010128
k2 f(P) c1 | -1.9772e-08
k2 f(P) c2 | -2.2669e-11
k0 HCL: | -1.4259 | Volts
Avg Ik: | -28.65 | nano amps
Std dev Ik: | 12.09 | nano amps
Avg Ib: | -1.57 | nano amps
Std dev Ib: | 11.97 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.001003 | 3.6437e-06 | -1.4259 | 5.0411e-05 | 0.9998 | 15 | 2.4 | 19.8
15.3 | -0.001003 | 3.2857e-06 | -1.4259 | 4.9262e-05 | 0.9998 | 18 | 2.1 | 19.8
24.7 | -0.0010064 | 2.7414e-06 | -1.4259 | 3.8621e-05 | 0.9998 | 23 | 5.1 | 19.8
42 | -0.0010101 | 3.8415e-06 | -1.426 | 5.3143e-05 | 0.9997 | 23 | 2.4 | 19.8
66.5 | -0.0010138 | 3.646e-06 | -1.4261 | 5.2672e-05 | 0.9997 | 22 | 2.4 | 19.8
91.7 | -0.0010137 | 4.8353e-06 | -1.4262 | 4.9668e-05 | 0.9998 | 9 | 2.4 | 19.8
121.2 | -0.0010168 | 4.1144e-06 | -1.4263 | 6.1721e-05 | 0.9995 | 33 | 2.4 | 19.8
172.7 | -0.0010273 | 4.2426e-06 | -1.4265 | 6.3167e-05 | 0.9997 | 18 | 5.2 | 19.8
246.4 | -0.0010253 | 3.7115e-06 | -1.4269 | 5.126e-05 | 0.9994 | 50 | 2.1 | 19.8
354 | -0.0010302 | 4.1123e-06 | -1.4273 | 6.2472e-05 | 0.9993 | 49 | 2.3 | 19.8
452.4 | -0.001039 | 4.2825e-06 | -1.4277 | 6.2271e-05 | 0.9994 | 40 | 5.2 | 19.8
547.5 | -0.0010393 | 3.802e-06 | -1.4281 | 5.0853e-05 | 0.9994 | 46 | 2.3 | 19.8
649.1 | -0.0010388 | 3.4189e-06 | -1.4286 | 4.5798e-05 | 0.9995 | 48 | 2.3 | 19.8
746.4 | -0.0010417 | 3.7038e-06 | -1.429 | 4.836e-05 | 0.9994 | 50 | 2.3 | 19.8
896.1 | -0.0010465 | 4.299e-06 | -1.4296 | 5.7176e-05 | 0.9984 | 95 | 2.3 | 19.8
1100.4 | -0.0010582 | 3.9235e-06 | -1.4303 | 5.6161e-05 | 0.9989 | 84 | 2.3 | 19.8
1301.5 | -0.0010703 | 4.1297e-06 | -1.4307 | 6.2145e-05 | 0.999 | 68 | 2.3 | 19.8
1503.5 | -0.0010846 | 4.2801e-06 | -1.4311 | 6.373e-05 | 0.9989 | 75 | 2.3 | 19.8
1701.9 | -0.0011007 | 5.7081e-06 | -1.4318 | 8.2974e-05 | 0.9984 | 61 | 2.3 | 19.8
1906.3 | -0.0011156 | 4.4594e-06 | -1.4322 | 4.3716e-05 | 0.9997 | 22 | 2.4 | 19.8
2038.5 | -0.0011773 | 4.1698e-05 | -1.4318 | 0.00016932 | 0.9975 | 4 | 2.4 | 5.2
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.6693e-08
f(P) R*R: | 0.9954
f(P) c0 | -1.4257
f(P) c1 | -7.6692e-06
f(P) c2 | 1.6129e-08
f(P) c3 | -3.3735e-11
f(P) c4 | 3.4166e-14
f(P) c5 | -1.5871e-17
f(P) c6 | 2.7423e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1020
k0 calibration window | 2022-04-25 18:43 | 2022-04-27 19:36
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 10
Ik outlier % | 0.68
Ib outlier count | 17
Ib outlier % | 1.16
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1020_DURA_2.dat | MSC1020_DURA.dat
median k0 | -1.4317 | -1.4308 | Volts
std k0 | 0.00012796 | 0.00013175 | Volts
Avg Vk | -0.8588 | -0.8588 | Volts
Avg std Vk | 50.7 | 55.8 | uVolts
median Ik: | -2.02 | -1.87 | nano amps
std Ik: | 38.22 | 39.74 | nano amps
median Ib: | -0.74 | 1.06 | nano amps
std Ib: | 38.8 | 39.02 | nano amps

## GDF18_Calibration_20220504
PH CALIBRATION FILE GENERATED ON: 05/04/2022 15:50
GDF 18
MSC 1021
APEX ID 21811
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21811, APEX# 18
8, number of calibration coefficients
-1.40272, =k0 Pon; -1.40049 =k0 Poff; -1.39113 =k0 HCl
-4.66454e-14,1.10090e-10,-1.10978e-07,-1.04074e-03, =k2(fP)*T
-3.41611e-06, =k3*P
-3.70574e-09, =k4*P^3
1.82963e-11, =k5*P^4
-2.73159e-14, =k6*P^5
1.52006e-17, =k7*P^6
-2.82355e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
04/06-04/09/2022 | B |  | -0.527 |  |  |  | -4.6645e-14 | 1.1009e-10 | -1.1098e-07 | -0.0010407 | -1.3911 | 0.9998 |  |  | -2.8235e-21 | 1.5201e-17 | -2.7316e-14 | 1.8296e-11 | -3.7057e-09 | -3.4161e-06 | -1.391 | 0.9973 | GDF |  | 18 | 05/02-05/04/2022 |  | -1.4005 | 6.5575e-05 | -1.4027 | 4.0872e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 18
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF#18_newchip_chanB_220405.csv
PT test channel : | B
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-04-06 17:39 | 2022-04-09 04:20
Data subset: | ASCENDING
Vrs drift rate: | -0.52696 | mV/day
Vk drift rate: | 15.2755 | mV/day
Avg std Vk | 51.7378 | uVolts
Average k2: | -0.0010827 | V/degC
Std dev k2: | 4.3668e-05 | V/degC
k2 f(P) c0 | -0.0010407
k2 f(P) c1 | -1.1098e-07
k2 f(P) c2 | 1.1009e-10
k2 f(P) c3 | -4.6645e-14
k0 HCL: | -1.3911 | Volts
Avg Ik: | -27.85 | nano amps
Std dev Ik: | 10.7 | nano amps
Avg Ib: | -2.21 | nano amps
Std dev Ib: | 11.27 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.0010349 | 4.1033e-06 | -1.3911 | 5.5042e-05 | 0.9998 | 16 | 2.4 | 19.8
14.7 | -0.0010399 | 3.132e-06 | -1.3911 | 4.1909e-05 | 0.9998 | 22 | 2.4 | 19.8
23.8 | -0.0010441 | 3.6381e-06 | -1.3911 | 5.0485e-05 | 0.9998 | 18 | 2.4 | 19.8
39.7 | -0.001051 | 2.7014e-06 | -1.391 | 3.897e-05 | 0.9999 | 21 | 5.2 | 19.8
64.7 | -0.001046 | 3.875e-06 | -1.3913 | 5.777e-05 | 0.9998 | 20 | 2.4 | 19.8
93.5 | -0.0010533 | 2.5719e-06 | -1.3913 | 3.6285e-05 | 0.9999 | 17 | 5.2 | 19.8
124.7 | -0.0010499 | 2.6505e-06 | -1.3915 | 3.667e-05 | 0.9999 | 25 | 2.4 | 19.8
170.7 | -0.001057 | 2.9969e-06 | -1.3916 | 4.1178e-05 | 0.9998 | 22 | 2.4 | 19.8
243.1 | -0.001062 | 3.8161e-06 | -1.3919 | 5.8162e-05 | 0.9995 | 37 | 5.2 | 19.8
352.7 | -0.0010692 | 3.0898e-06 | -1.3922 | 4.5842e-05 | 0.9996 | 51 | 2.3 | 19.8
454.8 | -0.0010811 | 2.7722e-06 | -1.3924 | 3.9881e-05 | 0.9997 | 41 | 5.2 | 19.8
552.3 | -0.0010805 | 3.5802e-06 | -1.3927 | 4.88e-05 | 0.9995 | 46 | 5.2 | 19.8
653 | -0.0010798 | 4.8768e-06 | -1.3931 | 7.0609e-05 | 0.9992 | 40 | 5.2 | 19.8
748.3 | -0.0010766 | 4.4886e-06 | -1.3936 | 5.8635e-05 | 0.9992 | 50 | 2.3 | 19.8
897.6 | -0.0010787 | 5.5236e-06 | -1.3943 | 7.6313e-05 | 0.9977 | 88 | 2.3 | 19.8
1095.4 | -0.0010885 | 6.5996e-06 | -1.3954 | 9.5036e-05 | 0.9972 | 79 | 2.3 | 19.8
1297.3 | -0.0010977 | 6.8103e-06 | -1.3967 | 9.4583e-05 | 0.9968 | 86 | 2.3 | 19.8
1502.9 | -0.0011205 | 4.075e-06 | -1.3977 | 5.3697e-05 | 0.9988 | 94 | 2.3 | 19.8
1698.6 | -0.0011495 | 2.4444e-06 | -1.398 | 3.1536e-05 | 0.9997 | 64 | 2.3 | 19.8
1916.2 | -0.0011762 | 4.8688e-06 | -1.3978 | 4.947e-05 | 0.9996 | 23 | 2.3 | 19.8
2037.5 | -0.0012006 | Inf | -1.3975 | Inf | 1 | 2 | 2.3 | 10.1
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.6152e-08
f(P) R*R: | 0.9973
f(P) c0 | -1.391
f(P) c1 | -3.4161e-06
f(P) c2 | -3.7057e-09
f(P) c3 | 1.8296e-11
f(P) c4 | -2.7316e-14
f(P) c5 | 1.5201e-17
f(P) c6 | -2.8235e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1021
k0 calibration window | 2022-05-02 18:51 | 2022-05-04 19:47
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 7
Ik outlier % | 0.48
Ib outlier count | 3
Ib outlier % | 0.2
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1021_DURA_2.dat | MSC1021_DURA.dat
median k0 | -1.4027 | -1.4005 | Volts
std k0 | 4.0871e-05 | 6.5575e-05 | Volts
Avg Vk | -0.8347 | -0.8348 | Volts
Avg std Vk | 50.5 | 51.1 | uVolts
median Ik: | -13.96 | -13.44 | nano amps
std Ik: | 4.77 | 4.86 | nano amps
median Ib: | -0.66 | -1.17 | nano amps
std Ib: | 6.94 | 7.06 | nano amps

## GDF17_Calibration_20220504
PH CALIBRATION FILE GENERATED ON: 05/04/2022 15:59
GDF 17
MSC 1022
APEX ID 21057
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21057, APEX# 17
8, number of calibration coefficients
-1.38274, =k0 Pon; -1.38181 =k0 Poff; -1.37540 =k0 HCl
9.34710e-17,-3.93616e-13,4.81263e-10,-1.97568e-07,-1.00860e-03, =k2(fP)*T
-8.53675e-07, =k3*P
-1.06454e-08, =k4*P^3
2.15332e-11, =k5*P^4
-2.49496e-14, =k6*P^5
1.31454e-17, =k7*P^6
-2.44275e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
04/16-04/19/2022 | A |  | -0.191 |  |  | 9.3471e-17 | -3.9362e-13 | 4.8126e-10 | -1.9757e-07 | -0.0010086 | -1.3754 | 0.9997 |  |  | -2.4427e-21 | 1.3145e-17 | -2.495e-14 | 2.1533e-11 | -1.0645e-08 | -8.5368e-07 | -1.3754 | 0.9984 | GDF |  | 17 | 05/03-05/04/2022 |  | -1.3818 | 6.0515e-05 | -1.3827 | 7.7683e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 17
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF#17_041522_chanA_newchip_peek.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-04-16 15:23 | 2022-04-19 02:06
Data subset: | ASCENDING
Vrs drift rate: | -0.19078 | mV/day
Vk drift rate: | -7.2636 | mV/day
Avg std Vk | 177.2047 | uVolts
Average k2: | -0.0010415 | V/degC
Std dev k2: | 3.6199e-05 | V/degC
k2 f(P) c0 | -0.0010086
k2 f(P) c1 | -1.9757e-07
k2 f(P) c2 | 4.8126e-10
k2 f(P) c3 | -3.9362e-13
k2 f(P) c4 | 9.3471e-17
k0 HCL: | -1.3754 | Volts
Avg Ik: | -5.26 | nano amps
Std dev Ik: | 3.14 | nano amps
Avg Ib: | 9.77 | nano amps
Std dev Ib: | 3.92 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.0010099 | 4.8384e-06 | -1.3754 | 6.4874e-05 | 0.9997 | 16 | 2.3 | 19.8
15.2 | -0.0010126 | 4.1476e-06 | -1.3754 | 6.4938e-05 | 0.9997 | 18 | 2.3 | 19.8
24.7 | -0.0010168 | 3.5536e-06 | -1.3754 | 5.1962e-05 | 0.9997 | 26 | 2.3 | 19.8
42.1 | -0.0010177 | 3.2702e-06 | -1.3754 | 4.2686e-05 | 0.9998 | 23 | 2.3 | 19.8
65.7 | -0.0010231 | 3.5285e-06 | -1.3754 | 4.8655e-05 | 0.9998 | 22 | 2.3 | 19.8
91.4 | -0.0010207 | 3.0009e-06 | -1.3756 | 3.6893e-05 | 0.9999 | 12 | 2.3 | 19.8
121.2 | -0.0010249 | 3.5076e-06 | -1.3757 | 5.1972e-05 | 0.9997 | 29 | 2.3 | 19.8
174.4 | -0.0010215 | 3.5364e-06 | -1.3759 | 4.6927e-05 | 0.9997 | 24 | 5.1 | 19.8
251.5 | -0.0010304 | 3.378e-06 | -1.3761 | 5.0617e-05 | 0.9994 | 55 | 2.3 | 19.8
354.5 | -0.0010331 | 2.6692e-06 | -1.3764 | 3.5436e-05 | 0.9997 | 49 | 2.3 | 19.8
448.8 | -0.0010317 | 3.6526e-06 | -1.3767 | 5.0849e-05 | 0.9994 | 46 | 5.1 | 19.8
548.2 | -0.0010303 | 3.7517e-06 | -1.3772 | 5.7819e-05 | 0.9994 | 51 | 5.1 | 19.8
647.2 | -0.0010343 | 3.5945e-06 | -1.3775 | 4.9136e-05 | 0.9995 | 46 | 2.2 | 19.8
747.4 | -0.0010249 | 4.0846e-06 | -1.3781 | 5.6103e-05 | 0.9993 | 48 | 2.2 | 19.8
899.1 | -0.0010248 | 5.2316e-06 | -1.379 | 7.6744e-05 | 0.9976 | 95 | 2.2 | 19.8
1102 | -0.0010236 | 6.5038e-06 | -1.3803 | 9.0941e-05 | 0.9965 | 89 | 2.2 | 19.8
1299.6 | -0.0010441 | 6.6148e-06 | -1.3815 | 9.599e-05 | 0.9962 | 96 | 2.2 | 19.8
1503 | -0.0010788 | 4.5132e-06 | -1.3823 | 6.3909e-05 | 0.9985 | 85 | 2.3 | 19.8
1690 | -0.0011078 | 2.1519e-06 | -1.3824 | 2.9489e-05 | 0.9998 | 65 | 2.3 | 19.8
1900.6 | -0.0011342 | 3.7839e-06 | -1.3823 | 3.4205e-05 | 0.9996 | 34 | 2.3 | 19.8
2039.8 | -0.0011272 | 1.2366e-05 | -1.3823 | 7.5764e-05 | 0.9996 | 5 | 2.3 | 10.1
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.0795e-08
f(P) R*R: | 0.9984
f(P) c0 | -1.3754
f(P) c1 | -8.5368e-07
f(P) c2 | -1.0645e-08
f(P) c3 | 2.1533e-11
f(P) c4 | -2.495e-14
f(P) c5 | 1.3145e-17
f(P) c6 | -2.4427e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1022
k0 calibration window | 2022-05-03 13:04 | 2022-05-04 19:44
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 0
Ik outlier % | 0
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1022_DURA_2.dat | MSC1022_DURA.dat
median k0 | -1.3827 | -1.3818 | Volts
std k0 | 7.7683e-05 | 6.0515e-05 | Volts
Avg Vk | -0.858 | -0.858 | Volts
Avg std Vk | 42.8 | 42.5 | uVolts
median Ik: | -1.65 | -1.37 | nano amps
std Ik: | 3.13 | 3.11 | nano amps
median Ib: | -0.52 | -0.47 | nano amps
std Ib: | 5.23 | 5.78 | nano amps

## GDF20_Calibration_20220504
PH CALIBRATION FILE GENERATED ON: 05/04/2022 16:04
GDF 20
MSC 1024
APEX ID 20040
** START CONFIG FILE COPY BLOCK **
Durafet SN = 20040, APEX# 20
8, number of calibration coefficients
-1.40570, =k0 Pon; -1.40537 =k0 Poff; -1.39904 =k0 HCl
6.11460e-17,-2.83066e-13,3.97122e-10,-1.84091e-07,-1.02599e-03, =k2(fP)*T
-4.14129e-06, =k3*P
1.46090e-09, =k4*P^3
-3.41278e-12, =k5*P^4
4.35541e-16, =k6*P^5
1.55886e-18, =k7*P^6
-5.20596e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
04/16-04/19/2022 | B |  | -0.349 |  |  | 6.1146e-17 | -2.8307e-13 | 3.9712e-10 | -1.8409e-07 | -0.001026 | -1.399 | 0.9994 |  |  | -5.206e-22 | 1.5589e-18 | 4.3554e-16 | -3.4128e-12 | 1.4609e-09 | -4.1413e-06 | -1.3989 | 0.9963 | GDF |  | 20 | 05/02-05/04/2022 |  | -1.4054 | 5.9719e-05 | -1.4057 | 7.1754e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 20
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF#20_041522_chanB_newchip.csv
PT test channel : | B
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-04-16 15:24 | 2022-04-19 02:06
Data subset: | ASCENDING
Vrs drift rate: | -0.34928 | mV/day
Vk drift rate: | -6.4597 | mV/day
Avg std Vk | 49.9561 | uVolts
Average k2: | -0.0010473 | V/degC
Std dev k2: | 1.739e-05 | V/degC
k2 f(P) c0 | -0.001026
k2 f(P) c1 | -1.8409e-07
k2 f(P) c2 | 3.9712e-10
k2 f(P) c3 | -2.8307e-13
k2 f(P) c4 | 6.1146e-17
k0 HCL: | -1.399 | Volts
Avg Ik: | -28.75 | nano amps
Std dev Ik: | 10.49 | nano amps
Avg Ib: | -2.16 | nano amps
Std dev Ib: | 11.22 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.0010257 | 6.6264e-06 | -1.399 | 8.8818e-05 | 0.9994 | 16 | 2.3 | 19.8
14.8 | -0.0010268 | 4.7142e-06 | -1.399 | 6.8011e-05 | 0.9996 | 19 | 2.3 | 19.8
24.8 | -0.0010328 | 5.1726e-06 | -1.399 | 7.5106e-05 | 0.9994 | 26 | 2.3 | 19.8
41.1 | -0.0010336 | 4.0653e-06 | -1.399 | 5.5549e-05 | 0.9997 | 22 | 2.3 | 19.8
65.7 | -0.0010385 | 4.8294e-06 | -1.3991 | 6.8097e-05 | 0.9996 | 21 | 2.3 | 19.8
94.4 | -0.0010396 | 7.1738e-06 | -1.3992 | 9.2469e-05 | 0.9996 | 11 | 5.1 | 19.8
122.8 | -0.0010447 | 5.4383e-06 | -1.3993 | 8.3894e-05 | 0.9993 | 27 | 5.1 | 19.8
173.2 | -0.001045 | 5.4555e-06 | -1.3996 | 7.2343e-05 | 0.9994 | 25 | 5.1 | 19.8
251.5 | -0.0010481 | 4.1696e-06 | -1.4 | 6.055e-05 | 0.9991 | 60 | 2.3 | 19.8
353.6 | -0.001053 | 4.1586e-06 | -1.4004 | 5.7372e-05 | 0.9993 | 45 | 2.3 | 19.8
444.4 | -0.0010514 | 4.2956e-06 | -1.4007 | 6.2428e-05 | 0.9993 | 41 | 2.3 | 19.8
547.7 | -0.0010544 | 4.5481e-06 | -1.4011 | 7.2447e-05 | 0.9992 | 47 | 5.1 | 19.8
649.9 | -0.0010511 | 4.641e-06 | -1.4016 | 6.619e-05 | 0.9993 | 40 | 2.3 | 19.8
746.7 | -0.0010365 | 4.1911e-06 | -1.4022 | 5.6809e-05 | 0.9993 | 47 | 2.2 | 19.8
898.2 | -0.0010277 | 4.7151e-06 | -1.4031 | 6.4933e-05 | 0.9978 | 107 | 2.3 | 19.8
1099.6 | -0.0010359 | 3.981e-06 | -1.404 | 5.3502e-05 | 0.9986 | 96 | 2.3 | 19.8
1296.8 | -0.0010429 | 3.9026e-06 | -1.4047 | 5.7172e-05 | 0.9987 | 94 | 2.3 | 19.8
1503.1 | -0.0010596 | 2.5151e-06 | -1.4052 | 3.3654e-05 | 0.9995 | 94 | 2.3 | 19.8
1691 | -0.0010709 | 1.9495e-06 | -1.4053 | 2.8228e-05 | 0.9998 | 64 | 2.3 | 19.8
1889.9 | -0.0010788 | 5.2066e-06 | -1.4052 | 4.72e-05 | 0.9993 | 31 | 2.3 | 19.8
2040 | -0.0010968 | Inf | -1.4052 | Inf | 1 | 2 | 2.3 | 5.2
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.9463e-08
f(P) R*R: | 0.9963
f(P) c0 | -1.3989
f(P) c1 | -4.1413e-06
f(P) c2 | 1.4609e-09
f(P) c3 | -3.4128e-12
f(P) c4 | 4.3554e-16
f(P) c5 | 1.5589e-18
f(P) c6 | -5.206e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1024
k0 calibration window | 2022-05-02 19:06 | 2022-05-04 19:46
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 0
Ik outlier % | 0
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1024_DURA_2.dat | MSC1024_DURA.dat
median k0 | -1.4057 | -1.4054 | Volts
std k0 | 7.1754e-05 | 5.9719e-05 | Volts
Avg Vk | -0.9051 | -0.9051 | Volts
Avg std Vk | 65.8 | 60.3 | uVolts
median Ik: | -9.05 | -9.04 | nano amps
std Ik: | 4.04 | 4.2 | nano amps
median Ib: | 0.23 | -0.2 | nano amps
std Ib: | 6.39 | 6.35 | nano amps

## GDF11_Calibration_20220524
PH CALIBRATION FILE GENERATED ON: 05/24/2022 15:28
GDF 11
MSC 0980
APEX ID 20148
** START CONFIG FILE COPY BLOCK **
Durafet SN = 20148, APEX# 11
8, number of calibration coefficients
-1.43498, =k0 Pon; -1.43473 =k0 Poff; -1.43163 =k0 HCl
-1.91087e-11,-2.20036e-08,-1.00974e-03, =k2(fP)*T
-4.00026e-06, =k3*P
-1.60827e-08, =k4*P^3
4.52130e-11, =k5*P^4
-4.81131e-14, =k6*P^5
2.22488e-17, =k7*P^6
-3.71675e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
10/19-10/25/2021 |  |  | 0.087 |  |  |  |  | -1.9109e-11 | -2.2004e-08 | -0.0010097 | -1.4316 | 0.9993 |  |  | -3.7167e-21 | 2.2249e-17 | -4.8113e-14 | 4.5213e-11 | -1.6083e-08 | -4.0003e-06 | -1.4315 | 0.9957 | GDF |  | 11 | 11/09-11/11/2021 |  | -1.4347 | 0.00018007 | -1.435 | 0.00017683
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 11
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF11_fullcal.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2021-10-19 17:43 | 2021-10-25 22:24
Data subset: | ASCENDING
Vrs drift rate: | 0.08686 | mV/day
Vk drift rate: | -3.4621 | mV/day
Avg std Vk | 201.7045 | uVolts
Average k2: | -0.0010443 | V/degC
Std dev k2: | 3.8737e-05 | V/degC
k2 f(P) c0 | -0.0010097
k2 f(P) c1 | -2.2004e-08
k2 f(P) c2 | -1.9109e-11
k0 HCL: | -1.4316 | Volts
Avg Ik: | 7.2 | nano amps
Std dev Ik: | 0.55 | nano amps
Avg Ib: | -0.96 | nano amps
Std dev Ib: | 7.6 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
22.8 | -0.0010057 | 2.9962e-06 | -1.4316 | 3.8034e-05 | 0.9993 | 81 | 1.9 | 19.7
38.4 | -0.0010097 | 2.6822e-06 | -1.4317 | 3.3199e-05 | 0.9993 | 99 | 1.9 | 19.7
60.8 | -0.0010095 | 2.5775e-06 | -1.4318 | 3.2766e-05 | 0.9993 | 115 | 2 | 19.7
85.8 | -0.0010116 | 2.2856e-06 | -1.432 | 2.8915e-05 | 0.9995 | 109 | 1.9 | 19.7
123.6 | -0.001014 | 2.1035e-06 | -1.4322 | 2.6779e-05 | 0.9991 | 208 | 1.9 | 19.7
174.1 | -0.0010173 | 1.9953e-06 | -1.4325 | 2.4898e-05 | 0.9992 | 202 | 1.9 | 19.7
249.3 | -0.001018 | 2.0176e-06 | -1.4329 | 2.5364e-05 | 0.9985 | 395 | 2 | 19.7
348.7 | -0.0010211 | 1.9763e-06 | -1.4335 | 2.4902e-05 | 0.9986 | 389 | 2 | 19.7
448.8 | -0.0010245 | 2.0229e-06 | -1.434 | 2.5754e-05 | 0.9985 | 375 | 1.9 | 19.7
549.4 | -0.0010276 | 1.8971e-06 | -1.4345 | 2.4228e-05 | 0.9988 | 363 | 1.9 | 19.7
649.2 | -0.0010352 | 1.5364e-06 | -1.4348 | 1.9279e-05 | 0.9992 | 377 | 1.9 | 19.7
749.6 | -0.00104 | 1.3585e-06 | -1.4351 | 1.6906e-05 | 0.9994 | 381 | 1.9 | 19.7
899.9 | -0.0010435 | 1.6341e-06 | -1.4356 | 2.034e-05 | 0.9982 | 751 | 1.9 | 19.7
1099.1 | -0.0010518 | 2.4651e-06 | -1.4364 | 3.0611e-05 | 0.996 | 739 | 1.9 | 19.7
1299.2 | -0.0010658 | 2.2622e-06 | -1.4374 | 2.824e-05 | 0.9967 | 743 | 1.9 | 19.7
1499.3 | -0.0010851 | 1.4145e-06 | -1.438 | 1.7581e-05 | 0.9988 | 724 | 1.9 | 19.7
1699.2 | -0.0011068 | 9.7726e-07 | -1.4382 | 1.2241e-05 | 0.9994 | 724 | 1.9 | 19.7
1900.4 | -0.0011232 | 8.693e-07 | -1.438 | 1.0758e-05 | 0.9996 | 718 | 1.9 | 19.7
2033.9 | -0.0011313 | 1.5347e-06 | -1.4378 | 1.9208e-05 | 0.9996 | 227 | 2 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.8829e-08
f(P) R*R: | 0.9957
f(P) c0 | -1.4315
f(P) c1 | -4.0003e-06
f(P) c2 | -1.6083e-08
f(P) c3 | 4.5213e-11
f(P) c4 | -4.8113e-14
f(P) c5 | 2.2249e-17
f(P) c6 | -3.7167e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 980
k0 calibration window | 2021-11-09 20:57 | 2021-11-11 22:35
Color pH time offset | 8 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 5
Ik outlier % | 0.17
Ib outlier count | 1
Ib outlier % | 0.03
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC0980_DURA_2.dat | MSC0980_DURA.dat
median k0 | -1.435 | -1.4348 | Volts
std k0 | 0.00017683 | 0.00018007 | Volts
Avg Vk | -0.9902 | -0.9902 | Volts
Avg std Vk | 79.4 | 81 | uVolts
median Ik: | -4.12 | -4.34 | nano amps
std Ik: | 4.56 | 4.47 | nano amps
median Ib: | -0.02 | 0.5 | nano amps
std Ib: | 8.14 | 8.17 | nano amps

## GDF12_Calibration_20220524
PH CALIBRATION FILE GENERATED ON: 05/24/2022 15:31
GDF 12
MSC 0979
APEX ID 20675
** START CONFIG FILE COPY BLOCK **
Durafet SN = 20675, APEX# 12
8, number of calibration coefficients
-1.43168, =k0 Pon; -1.43140 =k0 Poff; -1.43057 =k0 HCl
-3.83738e-14,9.52623e-11,-1.00326e-07,-1.01419e-03, =k2(fP)*T
2.31509e-06, =k3*P
-3.55665e-08, =k4*P^3
7.55022e-11, =k5*P^4
-7.03724e-14, =k6*P^5
2.93994e-17, =k7*P^6
-4.48976e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
10/11-10/17/2021 |  |  | 0.039 |  |  |  | -3.8374e-14 | 9.5262e-11 | -1.0033e-07 | -0.0010142 | -1.4306 | 0.9995 |  |  | -4.4898e-21 | 2.9399e-17 | -7.0372e-14 | 7.5502e-11 | -3.5567e-08 | 2.3151e-06 | -1.4308 | 0.9947 | GDF |  | 12 | 11/09-11/11/2021 |  | -1.4314 | 0.00012886 | -1.4317 | 0.00014215
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 12
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF12_cal_every2C.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2021-10-11 11:25 | 2021-10-17 16:06
Data subset: | ASCENDING
Vrs drift rate: | 0.03859 | mV/day
Vk drift rate: | -1.2606 | mV/day
Avg std Vk | 215.4753 | uVolts
Average k2: | -0.0010541 | V/degC
Std dev k2: | 3.5568e-05 | V/degC
k2 f(P) c0 | -0.0010142
k2 f(P) c1 | -1.0033e-07
k2 f(P) c2 | 9.5262e-11
k2 f(P) c3 | -3.8374e-14
k0 HCL: | -1.4306 | Volts
Avg Ik: | 7.35 | nano amps
Std dev Ik: | 0.5 | nano amps
Avg Ib: | -0.4 | nano amps
Std dev Ib: | 13.11 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
22.7 | -0.0010155 | 2.4009e-06 | -1.4306 | 2.9272e-05 | 0.9995 | 83 | 1.9 | 19.7
39.1 | -0.0010207 | 3.1472e-06 | -1.4306 | 4.0958e-05 | 0.9991 | 98 | 1.9 | 19.7
62.2 | -0.0010203 | 2.3495e-06 | -1.4307 | 2.921e-05 | 0.9994 | 117 | 1.9 | 19.7
87.2 | -0.0010238 | 2.3711e-06 | -1.4308 | 2.9105e-05 | 0.9994 | 106 | 1.9 | 19.7
123.7 | -0.0010225 | 1.339e-06 | -1.4311 | 1.6478e-05 | 0.9996 | 211 | 1.9 | 19.7
174 | -0.0010273 | 1.412e-06 | -1.4313 | 1.7645e-05 | 0.9996 | 211 | 1.9 | 19.7
249.1 | -0.0010324 | 1.4426e-06 | -1.4316 | 1.8004e-05 | 0.9992 | 403 | 1.9 | 19.7
348.9 | -0.0010394 | 1.6576e-06 | -1.4319 | 2.0627e-05 | 0.999 | 399 | 1.9 | 19.7
448.5 | -0.001045 | 1.9066e-06 | -1.4323 | 2.4038e-05 | 0.9987 | 379 | 1.9 | 19.7
550.1 | -0.0010445 | 1.5643e-06 | -1.4327 | 1.9747e-05 | 0.9991 | 382 | 1.9 | 19.7
650 | -0.0010482 | 1.4748e-06 | -1.4331 | 1.8334e-05 | 0.9992 | 386 | 1.9 | 19.7
749.7 | -0.0010559 | 1.2902e-06 | -1.4334 | 1.6074e-05 | 0.9994 | 385 | 1.9 | 19.7
899.4 | -0.0010637 | 1.2748e-06 | -1.4337 | 1.5859e-05 | 0.9989 | 759 | 1.9 | 19.7
1099.6 | -0.0010603 | 1.7334e-06 | -1.4344 | 2.1561e-05 | 0.998 | 748 | 1.9 | 19.7
1299.7 | -0.0010556 | 2.8952e-06 | -1.4355 | 3.5913e-05 | 0.9945 | 743 | 1.9 | 19.7
1499.6 | -0.0010728 | 2.9021e-06 | -1.4369 | 3.6089e-05 | 0.9947 | 731 | 1.9 | 19.7
1699.8 | -0.0011102 | 1.1865e-06 | -1.4372 | 1.4841e-05 | 0.9992 | 726 | 1.9 | 19.7
1899.8 | -0.0011308 | 6.8479e-07 | -1.4371 | 8.4671e-06 | 0.9997 | 717 | 1.9 | 19.7
2032 | -0.0011393 | 1.1837e-06 | -1.4369 | 1.4753e-05 | 0.9998 | 226 | 1.9 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.5898e-08
f(P) R*R: | 0.9947
f(P) c0 | -1.4308
f(P) c1 | 2.3151e-06
f(P) c2 | -3.5567e-08
f(P) c3 | 7.5502e-11
f(P) c4 | -7.0372e-14
f(P) c5 | 2.9399e-17
f(P) c6 | -4.4898e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 979
k0 calibration window | 2021-11-09 21:00 | 2021-11-11 22:34
Color pH time offset | 8 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 16
Ik outlier % | 0.54
Ib outlier count | 3
Ib outlier % | 0.1
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC0979_DURA_2.dat | MSC0979_DURA.dat
median k0 | -1.4316 | -1.4314 | Volts
std k0 | 0.00014215 | 0.00012886 | Volts
Avg Vk | -0.9658 | -0.9657 | Volts
Avg std Vk | 78.3 | 79 | uVolts
median Ik: | -13.84 | -13.48 | nano amps
std Ik: | 6.35 | 6.22 | nano amps
median Ib: | -1.3 | -0.49 | nano amps
std Ib: | 9.4 | 9.3 | nano amps

## GDF13_Calibration_20220524
PH CALIBRATION FILE GENERATED ON: 05/24/2022 15:35
GDF 13
MSC 0998
APEX ID 21489
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21489, APEX# 13
8, number of calibration coefficients
-1.40332, =k0 Pon; -1.40294 =k0 Poff; -1.40197 =k0 HCl
-3.03378e-14,7.40985e-11,-7.89069e-08,-1.02746e-03, =k2(fP)*T
1.91939e-06, =k3*P
-2.47090e-08, =k4*P^3
4.27436e-11, =k5*P^4
-3.57104e-14, =k6*P^5
1.37425e-17, =k7*P^6
-1.93134e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
12/01-12/07/2021 |  |  | 0.002 |  |  |  | -3.0338e-14 | 7.4098e-11 | -7.8907e-08 | -0.0010275 | -1.402 | 0.9998 |  |  | -1.9313e-21 | 1.3743e-17 | -3.571e-14 | 4.2744e-11 | -2.4709e-08 | 1.9194e-06 | -1.4021 | 0.9959 | GDF |  | 13 | 01/10-01/13/2022 |  | -1.4029 | 8.4454e-05 | -1.4033 | 8.0669e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 13
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF13_cal_every2C.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2021-12-01 17:25 | 2021-12-07 21:14
Data subset: | ASCENDING
Vrs drift rate: | 0.00202 | mV/day
Vk drift rate: | 6.401 | mV/day
Avg std Vk | 189.383 | uVolts
Average k2: | -0.0010599 | V/degC
Std dev k2: | 2.9445e-05 | V/degC
k2 f(P) c0 | -0.0010275
k2 f(P) c1 | -7.8907e-08
k2 f(P) c2 | 7.4098e-11
k2 f(P) c3 | -3.0338e-14
k0 HCL: | -1.402 | Volts
Avg Ik: | 20.12 | nano amps
Std dev Ik: | 0.71 | nano amps
Avg Ib: | 10.11 | nano amps
Std dev Ib: | 54.46 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.4 | -0.0010298 | 1.5561e-06 | -1.402 | 1.9298e-05 | 0.9998 | 106 | 2 | 19.7
38.7 | -0.0010287 | 1.4372e-06 | -1.402 | 1.8165e-05 | 0.9997 | 138 | 2 | 19.7
61.2 | -0.0010327 | 1.3998e-06 | -1.402 | 1.7383e-05 | 0.9997 | 143 | 2 | 19.7
86.7 | -0.0010346 | 1.4111e-06 | -1.4021 | 1.7983e-05 | 0.9998 | 131 | 2 | 19.7
124.4 | -0.0010386 | 1.2529e-06 | -1.4022 | 1.5627e-05 | 0.9996 | 252 | 2 | 19.7
174.2 | -0.0010388 | 1.1877e-06 | -1.4024 | 1.4835e-05 | 0.9997 | 239 | 2 | 19.7
248.6 | -0.0010427 | 1.2883e-06 | -1.4027 | 1.6241e-05 | 0.9993 | 467 | 2 | 19.7
349.7 | -0.0010456 | 1.3604e-06 | -1.4031 | 1.729e-05 | 0.9992 | 454 | 2 | 19.7
449.2 | -0.0010497 | 1.4369e-06 | -1.4035 | 1.8022e-05 | 0.9992 | 455 | 2 | 19.7
550 | -0.0010513 | 1.5195e-06 | -1.404 | 1.8875e-05 | 0.999 | 461 | 1.9 | 19.7
649.5 | -0.0010537 | 1.597e-06 | -1.4045 | 1.9906e-05 | 0.999 | 443 | 1.9 | 19.7
748.6 | -0.0010565 | 1.6759e-06 | -1.405 | 2.0981e-05 | 0.9989 | 452 | 1.9 | 19.7
898.9 | -0.001066 | 1.9039e-06 | -1.4056 | 2.3731e-05 | 0.9972 | 895 | 1.9 | 19.7
1099 | -0.0010741 | 1.9667e-06 | -1.4065 | 2.4512e-05 | 0.9971 | 883 | 1.9 | 19.7
1298.9 | -0.0010675 | 2.5884e-06 | -1.4077 | 3.2353e-05 | 0.9949 | 878 | 1.9 | 19.7
1499.3 | -0.0010715 | 3.0751e-06 | -1.4091 | 3.846e-05 | 0.9929 | 871 | 2 | 19.7
1698.8 | -0.0010978 | 2.4595e-06 | -1.41 | 3.0799e-05 | 0.9957 | 854 | 2 | 19.7
1896.6 | -0.0011251 | 1.4605e-06 | -1.4103 | 1.8186e-05 | 0.9986 | 840 | 2 | 19.7
2038.9 | -0.0011338 | 1.5463e-06 | -1.4103 | 2.0612e-05 | 0.9994 | 312 | 2 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 3.3581e-08
f(P) R*R: | 0.9959
f(P) c0 | -1.4021
f(P) c1 | 1.9194e-06
f(P) c2 | -2.4709e-08
f(P) c3 | 4.2744e-11
f(P) c4 | -3.571e-14
f(P) c5 | 1.3742e-17
f(P) c6 | -1.9313e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 998
k0 calibration window | 2022-01-10 13:24 | 2022-01-13 01:36
Color pH time offset | 8 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 13
Ik outlier % | 0.54
Ib outlier count | 8
Ib outlier % | 0.33
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC0998_DURA_2.dat | MSC0998_DURA.dat
median k0 | -1.4033 | -1.4029 | Volts
std k0 | 8.0668e-05 | 8.4454e-05 | Volts
Avg Vk | -0.9394 | -0.9394 | Volts
Avg std Vk | 56.1 | 53.8 | uVolts
median Ik: | -4.44 | -4.49 | nano amps
std Ik: | 40.06 | 38.94 | nano amps
median Ib: | -1.35 | -0.85 | nano amps
std Ib: | 41.65 | 43.71 | nano amps

## GDF14_Calibration_20220524
PH CALIBRATION FILE GENERATED ON: 05/24/2022 15:40
GDF 14
MSC 0997
APEX ID 21141
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21141, APEX# 14
8, number of calibration coefficients
-1.42032, =k0 Pon; -1.41968 =k0 Poff; -1.41490 =k0 HCl
-1.50103e-11,-3.60248e-08,-1.09013e-03, =k2(fP)*T
-5.13652e-06, =k3*P
-2.65875e-09, =k4*P^3
6.27730e-12, =k5*P^4
-1.93945e-15, =k6*P^5
-8.74549e-19, =k7*P^6
3.96582e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
12/10-12/15/2021 |  |  | -0.024 |  |  |  |  | -1.501e-11 | -3.6025e-08 | -0.0010901 | -1.4149 | 0.9997 |  |  | 3.9658e-22 | -8.7455e-19 | -1.9394e-15 | 6.2773e-12 | -2.6588e-09 | -5.1365e-06 | -1.4148 | 0.9915 | GDF |  | 14 | 01/10-01/13/2022 |  | -1.4197 | 0.00011783 | -1.4203 | 0.00011463
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 14
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF14_cal_every2C.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2021-12-10 02:15 | 2021-12-15 23:04
Data subset: | ASCENDING
Vrs drift rate: | -0.02406 | mV/day
Vk drift rate: | 4.7252 | mV/day
Avg std Vk | 189.4812 | uVolts
Average k2: | -0.001131 | V/degC
Std dev k2: | 4.2501e-05 | V/degC
k2 f(P) c0 | -0.0010901
k2 f(P) c1 | -3.6025e-08
k2 f(P) c2 | -1.501e-11
k0 HCL: | -1.4149 | Volts
Avg Ik: | 20.02 | nano amps
Std dev Ik: | 0.67 | nano amps
Avg Ib: | 7.86 | nano amps
Std dev Ib: | 49.43 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.2 | -0.0010896 | 2.0694e-06 | -1.4149 | 2.4935e-05 | 0.9997 | 91 | 2 | 19.7
39.3 | -0.0010891 | 2.0882e-06 | -1.415 | 2.4757e-05 | 0.9996 | 116 | 2 | 19.7
62.3 | -0.0010919 | 1.8817e-06 | -1.4151 | 2.2522e-05 | 0.9996 | 128 | 2 | 19.7
87.1 | -0.001091 | 2.0628e-06 | -1.4153 | 2.4858e-05 | 0.9996 | 119 | 2 | 19.7
123.8 | -0.0010957 | 1.7783e-06 | -1.4154 | 2.1321e-05 | 0.9994 | 237 | 1.9 | 19.7
173.9 | -0.0010981 | 1.8016e-06 | -1.4157 | 2.1495e-05 | 0.9994 | 234 | 1.9 | 19.7
248.8 | -0.0011023 | 1.7447e-06 | -1.4161 | 2.0861e-05 | 0.9989 | 449 | 1.9 | 19.7
348.9 | -0.0011081 | 1.6979e-06 | -1.4166 | 2.0268e-05 | 0.999 | 445 | 1.9 | 19.7
449.1 | -0.001113 | 1.6071e-06 | -1.4171 | 1.9237e-05 | 0.9991 | 434 | 1.9 | 19.7
549.7 | -0.0011161 | 1.497e-06 | -1.4175 | 1.786e-05 | 0.9992 | 433 | 1.9 | 19.7
649.8 | -0.0011188 | 1.418e-06 | -1.4179 | 1.6992e-05 | 0.9993 | 430 | 1.9 | 19.7
749.7 | -0.0011235 | 1.332e-06 | -1.4183 | 1.5865e-05 | 0.9994 | 423 | 1.9 | 19.7
899.1 | -0.0011341 | 1.2114e-06 | -1.4186 | 1.448e-05 | 0.999 | 844 | 1.9 | 19.7
1098.8 | -0.0011434 | 9.0498e-07 | -1.4189 | 1.0806e-05 | 0.9995 | 837 | 1.9 | 19.7
1299.3 | -0.0011575 | 9.4617e-07 | -1.419 | 1.1307e-05 | 0.9994 | 835 | 1.9 | 19.7
1500 | -0.0011782 | 1.3529e-06 | -1.4192 | 1.6189e-05 | 0.9989 | 816 | 2 | 19.7
1700 | -0.0012033 | 1.0579e-06 | -1.4193 | 1.2683e-05 | 0.9994 | 808 | 2 | 19.7
1897.2 | -0.0012149 | 1.0148e-06 | -1.4192 | 1.2057e-05 | 0.9995 | 784 | 2 | 19.7
2029.1 | -0.0012201 | 1.7942e-06 | -1.4191 | 2.1589e-05 | 0.9994 | 270 | 2 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.5309e-08
f(P) R*R: | 0.9915
f(P) c0 | -1.4148
f(P) c1 | -5.1365e-06
f(P) c2 | -2.6587e-09
f(P) c3 | 6.2773e-12
f(P) c4 | -1.9395e-15
f(P) c5 | -8.7455e-19
f(P) c6 | 3.9658e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 997
k0 calibration window | 2022-01-10 16:06 | 2022-01-13 01:00
Color pH time offset | 8 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 1
Ik outlier % | 0.04
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC0997_DURA_2.dat | MSC0997_DURA.dat
median k0 | -1.4203 | -1.4197 | Volts
std k0 | 0.00011463 | 0.00011783 | Volts
Avg Vk | -0.9266 | -0.9266 | Volts
Avg std Vk | 61.5 | 62.7 | uVolts
median Ik: | -1.19 | -1.29 | nano amps
std Ik: | 4.3 | 4.35 | nano amps
median Ib: | 0.45 | 0.16 | nano amps
std Ib: | 6.65 | 6.58 | nano amps

## GDF_Calibration_20220623
PH CALIBRATION FILE GENERATED ON: 06/23/2022 16:36
GDF
MSC 1031
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX#
8, number of calibration coefficients
-1.41566, =k0 Pon; -1.41530 =k0 Poff; -1.40816 =k0 HCl
-4.38471e-14,1.12304e-10,-1.02946e-07,-1.06369e-03, =k2(fP)*T
-2.99379e-06, =k3*P
-7.13216e-10, =k4*P^3
2.30616e-12, =k5*P^4
-5.62324e-15, =k6*P^5
3.95117e-18, =k7*P^6
-8.05273e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
05/20-05/25/2022 |  |  | -0.102 |  |  |  | -4.3847e-14 | 1.123e-10 | -1.0295e-07 | -0.0010637 | -1.4082 | 0.9994 |  |  | -8.0527e-22 | 3.9512e-18 | -5.6232e-15 | 2.3062e-12 | -7.1322e-10 | -2.9938e-06 | -1.408 | 0.9959 | GDF |  |  | 06/20-06/22/2022 |  | -1.4153 | 0.00011816 | -1.4157 | 0.00011437
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN:
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF 22_18May2022_ChD.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-05-20 03:05 | 2022-05-25 06:07
Data subset: | ASCENDING
Vrs drift rate: | -0.10214 | mV/day
Vk drift rate: | 1.4045 | mV/day
Avg std Vk | 132.6922 | uVolts
Average k2: | -0.0010974 | V/degC
Std dev k2: | 2.9629e-05 | V/degC
k2 f(P) c0 | -0.0010637
k2 f(P) c1 | -1.0295e-07
k2 f(P) c2 | 1.123e-10
k2 f(P) c3 | -4.3847e-14
k0 HCL: | -1.4082 | Volts
Avg Ik: | -34.85 | nano amps
Std dev Ik: | 12.12 | nano amps
Avg Ib: | -0.5 | nano amps
Std dev Ib: | 13.95 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
24.8 | -0.0010561 | 4.7434e-06 | -1.4082 | 6.7972e-05 | 0.9994 | 31 | 2.2 | 19.7
37.7 | -0.0010577 | 3.0996e-06 | -1.4082 | 4.4221e-05 | 0.9996 | 51 | 2.2 | 19.7
60.6 | -0.0010794 | 4.5093e-06 | -1.4079 | 6.9236e-05 | 0.9994 | 36 | 9.9 | 19.7
87 | -0.0010729 | 5.453e-06 | -1.4082 | 8.5469e-05 | 0.9994 | 27 | 2.1 | 19.7
123 | -0.0010772 | 4.0949e-06 | -1.4083 | 6.2244e-05 | 0.9993 | 49 | 2.1 | 19.7
172.6 | -0.0010778 | 4.0379e-06 | -1.4085 | 6.0453e-05 | 0.9994 | 44 | 2.1 | 19.7
247.8 | -0.0010841 | 3.2787e-06 | -1.4087 | 4.9856e-05 | 0.9993 | 84 | 2 | 19.7
350.6 | -0.0010964 | 3.6724e-06 | -1.4089 | 5.6568e-05 | 0.9992 | 74 | 1.9 | 19.7
449.9 | -0.0011013 | 3.5703e-06 | -1.4092 | 5.4877e-05 | 0.9993 | 71 | 2 | 19.7
551.3 | -0.0010937 | 3.7233e-06 | -1.4097 | 5.592e-05 | 0.9992 | 74 | 2.1 | 19.7
648.3 | -0.0010991 | 3.6118e-06 | -1.4101 | 5.4456e-05 | 0.9992 | 73 | 2.1 | 19.7
752.4 | -0.0010897 | 3.3341e-06 | -1.4107 | 4.7337e-05 | 0.9992 | 83 | 2.1 | 19.7
900 | -0.0010875 | 3.5915e-06 | -1.4115 | 4.9956e-05 | 0.9982 | 165 | 2 | 19.7
1099.7 | -0.0010881 | 3.6348e-06 | -1.4125 | 5.0863e-05 | 0.9982 | 167 | 2 | 19.7
1298.4 | -0.0011029 | 3.5287e-06 | -1.4132 | 5.0138e-05 | 0.9984 | 157 | 2.1 | 19.7
1502.7 | -0.0011171 | 2.7392e-06 | -1.4139 | 3.857e-05 | 0.9991 | 158 | 2.1 | 19.7
1698.6 | -0.0011439 | 1.2613e-06 | -1.4139 | 1.754e-05 | 0.9998 | 156 | 2.1 | 19.7
1899.5 | -0.0011591 | 1.3185e-06 | -1.4136 | 1.8343e-05 | 0.9998 | 162 | 2.2 | 19.7
2034.7 | -0.001167 | 2.433e-06 | -1.4134 | 3.4064e-05 | 0.9998 | 50 | 2.2 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.9039e-08
f(P) R*R: | 0.9959
f(P) c0 | -1.4079
f(P) c1 | -2.9938e-06
f(P) c2 | -7.1322e-10
f(P) c3 | 2.3062e-12
f(P) c4 | -5.6232e-15
f(P) c5 | 3.9512e-18
f(P) c6 | -8.0527e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1031
k0 calibration window | 2022-06-20 21:57 | 2022-06-22 09:36
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 1
Ik outlier % | 0.07
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1031_DURA_2.dat | MSC1031_DURA.dat
median k0 | -1.4156 | -1.4153 | Volts
std k0 | 0.00011437 | 0.00011816 | Volts
Avg Vk | -0.9367 | -0.9367 | Volts
Avg std Vk | 75.8 | 78.7 | uVolts
median Ik: | -12.9 | -13.18 | nano amps
std Ik: | 49.02 | 53.06 | nano amps
median Ib: | 0.69 | -1.59 | nano amps
std Ib: | 56.31 | 55.47 | nano amps

## GDF23_Calibration_20220706
PH CALIBRATION FILE GENERATED ON: 07/06/2022 15:00
GDF 23
MSC XXXX
APEX ID XXXX
** START CONFIG FILE COPY BLOCK **
Durafet SN = XXXX, APEX# 23
8, number of calibration coefficients
NO K0 DATA WAS PEOCESSED YET OR LOADED INTO GUI!!!!
-4.57190e-14,1.43555e-10,-1.42007e-07,-1.03739e-03, =k2(fP)*T
-5.76316e-06, =k3*P
2.04877e-08, =k4*P^3
-3.21301e-11, =k5*P^4
1.85986e-14, =k6*P^5
-3.07186e-18, =k7*P^6
-2.05729e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
06/29-07/04/2022 |  |  | -0.009 |  |  |  | -4.5719e-14 | 1.4356e-10 | -1.4201e-07 | -0.0010374 | -1.3842 | 0.9999 |  |  | -2.0573e-22 | -3.0719e-18 | 1.8599e-14 | -3.213e-11 | 2.0488e-08 | -5.7632e-06 | -1.384 | 0.8071 | GDF |  | 23
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 23
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF23_062722_Ch E.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-06-29 04:13 | 2022-07-04 07:18
Data subset: | ASCENDING
Vrs drift rate: | -0.00925 | mV/day
Vk drift rate: | -3.8196 | mV/day
Avg std Vk | 101.9374 | uVolts
Average k2: | -0.0010725 | V/degC
Std dev k2: | 2.2709e-05 | V/degC
k2 f(P) c0 | -0.0010374
k2 f(P) c1 | -1.4201e-07
k2 f(P) c2 | 1.4355e-10
k2 f(P) c3 | -4.5719e-14
k0 HCL: | -1.3842 | Volts
Avg Ik: | 7.58 | nano amps
Std dev Ik: | 1.44 | nano amps
Avg Ib: | 1.15 | nano amps
Std dev Ib: | 2.22 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
24.2 | -0.0010444 | 2.6038e-06 | -1.3842 | 3.7596e-05 | 0.9999 | 23 | 1.9 | 19.8
38.9 | -0.0010468 | 1.9763e-06 | -1.3843 | 2.7248e-05 | 0.9999 | 34 | 1.9 | 19.8
61.3 | -0.0010475 | 1.6303e-06 | -1.3843 | 2.3313e-05 | 0.9999 | 34 | 1.9 | 19.8
88.1 | -0.001049 | 1.9386e-06 | -1.3843 | 2.7603e-05 | 0.9999 | 28 | 1.9 | 19.8
124 | -0.0010517 | 1.5221e-06 | -1.3843 | 2.1609e-05 | 0.9999 | 50 | 1.9 | 19.8
174.5 | -0.0010527 | 1.5305e-06 | -1.3844 | 2.1522e-05 | 0.9999 | 49 | 1.9 | 19.8
247.4 | -0.0010539 | 1.4602e-06 | -1.3846 | 2.0106e-05 | 0.9998 | 94 | 1.9 | 19.8
348.7 | -0.0010616 | 1.4133e-06 | -1.3847 | 1.9822e-05 | 0.9998 | 97 | 1.9 | 19.8
450.7 | -0.0010724 | 1.4926e-06 | -1.3849 | 2.0995e-05 | 0.9998 | 91 | 1.9 | 19.8
551.3 | -0.0010899 | 1.4779e-06 | -1.3848 | 2.0401e-05 | 0.9998 | 93 | 1.9 | 19.8
648.2 | -0.0011014 | 2.1985e-06 | -1.3848 | 3.3245e-05 | 0.9997 | 70 | 1.9 | 19.8
748.7 | -0.0011007 | 1.7065e-06 | -1.3849 | 2.3667e-05 | 0.9998 | 85 | 1.9 | 19.8
899.8 | -0.0010708 | 2.8051e-06 | -1.3856 | 3.9137e-05 | 0.9988 | 179 | 1.9 | 19.8
1098.7 | -0.0010508 | 5.293e-06 | -1.3869 | 7.5139e-05 | 0.9958 | 167 | 2 | 19.8
1300.6 | -0.0010841 | 5.5532e-06 | -1.3869 | 7.7354e-05 | 0.9956 | 169 | 2 | 19.8
1500.6 | -0.0010889 | 5.3966e-06 | -1.3862 | 7.4416e-05 | 0.9959 | 169 | 2 | 19.8
1700.1 | -0.0010938 | 6.2764e-06 | -1.3855 | 8.6786e-05 | 0.9946 | 166 | 2.1 | 19.8
1899.9 | -0.0011041 | 7.3057e-06 | -1.3849 | 0.00010288 | 0.9929 | 166 | 2.1 | 19.8
2033 | -0.0011126 | 1.3901e-05 | -1.3844 | 0.00018766 | 0.9916 | 56 | 2.1 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.6891e-07
f(P) R*R: | 0.8071
f(P) c0 | -1.384
f(P) c1 | -5.7632e-06
f(P) c2 | 2.0488e-08
f(P) c3 | -3.213e-11
f(P) c4 | 1.8599e-14
f(P) c5 | -3.0719e-18
f(P) c6 | -2.0573e-22
NO KO DATA FILE LOADED INTO CALIBRATION GUI!!!:
Has the k0 calibration be completed yet?:
Is the MSC # correct? Check MSC / SN match up:

## GDF023_Calibration_20220714_all
PH CALIBRATION FILE GENERATED ON: 07/14/2022 13:39
GDF 023
MSC XXXX
APEX ID XXXX
** START CONFIG FILE COPY BLOCK **
Durafet SN = XXXX, APEX# 023
8, number of calibration coefficients
NO K0 DATA WAS PEOCESSED YET OR LOADED INTO GUI!!!!
-3.73204e-11,6.21584e-09,-1.09112e-03, =k2(fP)*T
-2.36126e-06, =k3*P
6.50722e-10, =k4*P^3
-1.45437e-12, =k5*P^4
-1.10200e-15, =k6*P^5
1.67402e-18, =k7*P^6
-4.10807e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
07/09-07/14/2022 | B |  | -0.134 |  |  |  |  | -3.732e-11 | 6.2158e-09 | -0.0010911 | -1.4021 | 0.9993 |  |  | -4.1081e-22 | 1.674e-18 | -1.102e-15 | -1.4544e-12 | 6.5072e-10 | -2.3613e-06 | -1.4019 | 0.9715 | GDF |  | 23
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 23
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF023_070722_chanB.csv
PT test channel : | B
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-07-09 00:22 | 2022-07-14 10:16
Data subset: | ALL
Vrs drift rate: | -0.13442 | mV/day
Vk drift rate: | 2.6441 | mV/day
Avg std Vk | 62.6979 | uVolts
Average k2: | -0.0011228 | V/degC
Std dev k2: | 4.546e-05 | V/degC
k2 f(P) c0 | -0.0010911
k2 f(P) c1 | 6.2158e-09
k2 f(P) c2 | -3.732e-11
k0 HCL: | -1.4021 | Volts
Avg Ik: | -27.02 | nano amps
Std dev Ik: | 9.59 | nano amps
Avg Ib: | -1.37 | nano amps
Std dev Ib: | 10.16 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
21 | -0.0010769 | 4.7771e-07 | -1.4021 | 6.4079e-06 | 0.9993 | 3468 | 2.2 | 19.9
38.8 | -0.0010851 | 5.2874e-06 | -1.4019 | 7.8722e-05 | 0.999 | 45 | 2.3 | 19.9
61.3 | -0.0010912 | 5.81e-06 | -1.4019 | 8.3102e-05 | 0.9988 | 43 | 2.3 | 19.9
86.2 | -0.0010899 | 7.4332e-06 | -1.402 | 0.00010513 | 0.9982 | 41 | 2.3 | 19.9
122.9 | -0.0010922 | 6.3814e-06 | -1.4021 | 9.3901e-05 | 0.9976 | 71 | 2.3 | 19.9
174.2 | -0.0010901 | 8.5676e-06 | -1.4023 | 0.00012566 | 0.9958 | 71 | 2.3 | 19.9
249.6 | -0.001102 | 7.4371e-06 | -1.4023 | 0.00010684 | 0.9938 | 140 | 2.2 | 19.9
349 | -0.0010963 | 1.0105e-05 | -1.4028 | 0.00014633 | 0.9899 | 122 | 2.2 | 19.9
449.3 | -0.0011081 | 1.1846e-05 | -1.4028 | 0.00017574 | 0.9881 | 107 | 2.2 | 19.9
545.3 | -0.0011169 | 1.0462e-05 | -1.403 | 0.00015674 | 0.9904 | 112 | 2.2 | 19.9
648 | -0.0011007 | 8.0925e-06 | -1.4036 | 0.00011733 | 0.9934 | 125 | 2.2 | 19.9
749.4 | -0.001104 | 6.5588e-06 | -1.404 | 9.401e-05 | 0.9956 | 127 | 2.2 | 19.9
899.5 | -0.0011094 | 4.2284e-06 | -1.4046 | 6.0762e-05 | 0.9965 | 244 | 2.2 | 19.9
1098.3 | -0.0011254 | 3.911e-06 | -1.4053 | 5.7425e-05 | 0.9972 | 233 | 2.2 | 19.9
1301.7 | -0.001134 | 3.1001e-06 | -1.4061 | 4.5058e-05 | 0.9982 | 238 | 2.3 | 19.9
1502.3 | -0.0011615 | 2.9103e-06 | -1.4064 | 4.2577e-05 | 0.9986 | 233 | 2.3 | 19.9
1699.8 | -0.00119 | 2.7234e-06 | -1.4065 | 4.0578e-05 | 0.9989 | 215 | 2.3 | 19.9
1902.4 | -0.0012216 | 2.4707e-06 | -1.4061 | 3.5863e-05 | 0.9991 | 214 | 2.3 | 19.9
2064.3 | -0.0012383 | 1.1946e-06 | -1.4058 | 1.7562e-05 | 0.9992 | 861 | 2.4 | 19.9
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 9.5167e-08
f(P) R*R: | 0.9715
f(P) c0 | -1.4019
f(P) c1 | -2.3613e-06
f(P) c2 | 6.5072e-10
f(P) c3 | -1.4544e-12
f(P) c4 | -1.102e-15
f(P) c5 | 1.674e-18
f(P) c6 | -4.1081e-22
NO KO DATA FILE LOADED INTO CALIBRATION GUI!!!:
Has the k0 calibration be completed yet?:
Is the MSC # correct? Check MSC / SN match up:

## GDF010_Calibration_20220824
PH CALIBRATION FILE GENERATED ON: 08/24/2022 10:49
GDF 010
MSC 1032
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 010
8, number of calibration coefficients
-1.42250, =k0 Pon; -1.42231 =k0 Poff; -1.41949 =k0 HCl
-2.05184e-14,4.19768e-11,-4.74771e-08,-1.01086e-03, =k2(fP)*T
7.13956e-07, =k3*P
-9.62916e-09, =k4*P^3
1.34868e-11, =k5*P^4
-1.33484e-14, =k6*P^5
6.85171e-18, =k7*P^6
-1.28259e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
05/18-05/21/2022 |  |  | -0.011 |  |  |  | -2.0518e-14 | 4.1977e-11 | -4.7477e-08 | -0.0010109 | -1.4195 | 0.9998 |  |  | -1.2826e-21 | 6.8517e-18 | -1.3348e-14 | 1.3487e-11 | -9.6292e-09 | 7.1396e-07 | -1.4194 | 0.9975 | GDF |  | 10 | 08/10-08/16/2022 |  | -1.4223 | 0.00011209 | -1.4225 | 0.00010864
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 10
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF010_051722_Ch E.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-05-18 19:20 | 2022-05-21 06:03
Data subset: | ASCENDING
Vrs drift rate: | -0.01093 | mV/day
Vk drift rate: | 6.0873 | mV/day
Avg std Vk | 99.7254 | uVolts
Average k2: | -0.001034 | V/degC
Std dev k2: | 2.6389e-05 | V/degC
k2 f(P) c0 | -0.0010109
k2 f(P) c1 | -4.7477e-08
k2 f(P) c2 | 4.1977e-11
k2 f(P) c3 | -2.0518e-14
k0 HCL: | -1.4195 | Volts
Avg Ik: | 8.46 | nano amps
Std dev Ik: | 1.4 | nano amps
Avg Ib: | 1.18 | nano amps
Std dev Ib: | 2.3 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.0010055 | 3.4517e-06 | -1.4195 | 4.7291e-05 | 0.9998 | 16 | 2.6 | 20.2
14.2 | -0.0010111 | 2.5081e-06 | -1.4194 | 3.6301e-05 | 0.9999 | 24 | 5.4 | 20.2
23.3 | -0.0010083 | 3.413e-06 | -1.4195 | 4.0466e-05 | 0.9998 | 19 | 2.6 | 20.2
38.4 | -0.0010136 | 2.3561e-06 | -1.4194 | 3.3303e-05 | 0.9999 | 23 | 2.6 | 20.2
61.8 | -0.0010162 | 2.299e-06 | -1.4194 | 3.3577e-05 | 0.9999 | 21 | 5.4 | 20.2
85.1 | -0.0010185 | 3.047e-06 | -1.4194 | 3.769e-05 | 0.9999 | 16 | 5.4 | 20.2
126.4 | -0.0010171 | 1.7114e-06 | -1.4195 | 2.6628e-05 | 0.9999 | 34 | 2.6 | 20.2
176.5 | -0.0010225 | 2.591e-06 | -1.4195 | 3.5656e-05 | 0.9999 | 24 | 2.5 | 20.2
246.5 | -0.0010221 | 2.5332e-06 | -1.4197 | 3.5636e-05 | 0.9997 | 48 | 2.6 | 20.2
351.9 | -0.0010224 | 2.8286e-06 | -1.42 | 4.1908e-05 | 0.9996 | 52 | 2.5 | 20.2
453 | -0.0010249 | 2.8416e-06 | -1.4203 | 4.0648e-05 | 0.9996 | 50 | 2.5 | 20.2
550.7 | -0.0010259 | 2.811e-06 | -1.4207 | 3.8049e-05 | 0.9997 | 48 | 2.5 | 20.2
646.8 | -0.001027 | 3.1984e-06 | -1.4211 | 4.3232e-05 | 0.9995 | 51 | 2.5 | 20.2
745.2 | -0.0010303 | 2.5351e-06 | -1.4215 | 3.5026e-05 | 0.9997 | 46 | 2.5 | 20.2
889.9 | -0.0010339 | 4.4234e-06 | -1.4221 | 6.2623e-05 | 0.9984 | 89 | 2.5 | 20.2
1099 | -0.0010393 | 5.7416e-06 | -1.4231 | 8.7453e-05 | 0.9974 | 89 | 2.5 | 20.2
1301.3 | -0.0010516 | 4.2547e-06 | -1.424 | 6.2462e-05 | 0.9987 | 80 | 2.5 | 20.2
1498.1 | -0.0010585 | 3.1341e-06 | -1.4246 | 4.4996e-05 | 0.9993 | 80 | 2.5 | 20.2
1687.8 | -0.0010651 | 3.0615e-06 | -1.425 | 4.3339e-05 | 0.9996 | 53 | 2.6 | 20.2
1913.8 | -0.001093 | 9.3583e-06 | -1.4251 | 7.4989e-05 | 0.9992 | 13 | 2.6 | 10.4
2039.1 | -0.0011073 | 9.2511e-06 | -1.425 | 6.0184e-05 | 0.9997 | 7 | 2.6 | 10.4
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 9.9629e-09
f(P) R*R: | 0.9975
f(P) c0 | -1.4194
f(P) c1 | 7.1396e-07
f(P) c2 | -9.6292e-09
f(P) c3 | 1.3487e-11
f(P) c4 | -1.3348e-14
f(P) c5 | 6.8517e-18
f(P) c6 | -1.2826e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1032
k0 calibration window | 2022-08-10 19:30 | 2022-08-16 21:59
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 0
Ik outlier % | 0
Ib outlier count | 2
Ib outlier % | 0.02
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1032_DURA_2 (3).dat | MSC1032_DURA (3).dat
median k0 | -1.4225 | -1.4223 | Volts
std k0 | 0.00010864 | 0.00011209 | Volts
Avg Vk | -0.9216 | -0.9216 | Volts
Avg std Vk | 69.6 | 69.7 | uVolts
median Ik: | -1.59 | -1.49 | nano amps
std Ik: | 3.97 | 3.9 | nano amps
median Ib: | 0.86 | 0.85 | nano amps
std Ib: | 8.28 | 8.26 | nano amps

## GDF24_Calibration_20220824
PH CALIBRATION FILE GENERATED ON: 08/24/2022 10:53
GDF 24
MSC 1036
APEX ID 21302
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21302, APEX# 24
8, number of calibration coefficients
-1.38010, =k0 Pon; -1.37981 =k0 Poff; -1.37489 =k0 HCl
-4.63517e-14,1.24483e-10,-9.45510e-08,-1.00613e-03, =k2(fP)*T
-3.39205e-06, =k3*P
-8.09397e-09, =k4*P^3
1.73954e-11, =k5*P^4
-1.73720e-14, =k6*P^5
7.65214e-18, =k7*P^6
-1.17643e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
06/29-07/04/2022 |  |  | -0.167 |  |  |  | -4.6352e-14 | 1.2448e-10 | -9.4551e-08 | -0.0010061 | -1.3749 | 0.9993 |  |  | -1.1764e-21 | 7.6521e-18 | -1.7372e-14 | 1.7395e-11 | -8.094e-09 | -3.392e-06 | -1.3748 | 0.9947 | GDF |  | 24 | 08/22-08/23/2022 |  | -1.3798 | 6.517e-05 | -1.3801 | 6.4891e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 24
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF24_062722_Ch F.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-06-29 04:14 | 2022-07-04 07:19
Data subset: | ASCENDING
Vrs drift rate: | -0.167 | mV/day
Vk drift rate: | -1.8255 | mV/day
Avg std Vk | 213.4276 | uVolts
Average k2: | -0.0010258 | V/degC
Std dev k2: | 1.6568e-05 | V/degC
k2 f(P) c0 | -0.0010061
k2 f(P) c1 | -9.4551e-08
k2 f(P) c2 | 1.2448e-10
k2 f(P) c3 | -4.6352e-14
k0 HCL: | -1.3749 | Volts
Avg Ik: | -18.92 | nano amps
Std dev Ik: | 11.17 | nano amps
Avg Ib: | 1.22 | nano amps
Std dev Ib: | 9.03 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
24.2 | -0.001003 | 5.8803e-06 | -1.3749 | 8.5086e-05 | 0.9993 | 23 | 2 | 19.8
38.9 | -0.0010078 | 4.4373e-06 | -1.3749 | 6.1322e-05 | 0.9994 | 34 | 2 | 19.8
61.6 | -0.0010103 | 4.4896e-06 | -1.375 | 6.3405e-05 | 0.9993 | 36 | 2 | 19.8
87.7 | -0.0010119 | 5.3219e-06 | -1.3752 | 7.4636e-05 | 0.9992 | 30 | 2 | 19.8
125 | -0.0010158 | 4.2087e-06 | -1.3754 | 6.014e-05 | 0.9992 | 51 | 2 | 19.8
173.6 | -0.001023 | 4.1094e-06 | -1.3756 | 5.9402e-05 | 0.9993 | 47 | 2.1 | 19.8
246.9 | -0.0010278 | 3.1883e-06 | -1.3759 | 4.4238e-05 | 0.9991 | 93 | 2.1 | 19.8
348.7 | -0.0010317 | 3.1147e-06 | -1.3763 | 4.3572e-05 | 0.9991 | 98 | 2.1 | 19.8
448.5 | -0.001034 | 4.0399e-06 | -1.3768 | 6.0576e-05 | 0.9989 | 75 | 2.1 | 19.8
551.4 | -0.0010365 | 4.4045e-06 | -1.3772 | 6.6853e-05 | 0.9986 | 77 | 2.1 | 19.8
650.3 | -0.0010233 | 7.8052e-06 | -1.378 | 0.00012052 | 0.9961 | 69 | 9.9 | 19.8
751.2 | -0.0010041 | 4.5621e-06 | -1.3789 | 6.6851e-05 | 0.9984 | 80 | 2 | 19.8
899.8 | -0.0010251 | 3.7063e-06 | -1.3794 | 5.1832e-05 | 0.9977 | 179 | 2 | 19.8
1100.9 | -0.0010275 | 3.4532e-06 | -1.3802 | 4.8291e-05 | 0.9981 | 173 | 2 | 19.8
1301.2 | -0.0010176 | 5.3191e-06 | -1.3815 | 7.3382e-05 | 0.9955 | 169 | 2.1 | 19.8
1499.6 | -0.0010224 | 4.3791e-06 | -1.3828 | 6.0386e-05 | 0.9971 | 161 | 2.1 | 19.8
1698.8 | -0.0010414 | 2.5645e-06 | -1.3832 | 3.5917e-05 | 0.999 | 162 | 2.1 | 19.8
1899.2 | -0.00106 | 2.1076e-06 | -1.3833 | 2.9879e-05 | 0.9994 | 163 | 2.1 | 19.8
2033 | -0.0010662 | 4.3353e-06 | -1.3832 | 5.8662e-05 | 0.9991 | 56 | 2.2 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 4.5968e-08
f(P) R*R: | 0.9947
f(P) c0 | -1.3748
f(P) c1 | -3.3921e-06
f(P) c2 | -8.094e-09
f(P) c3 | 1.7395e-11
f(P) c4 | -1.7372e-14
f(P) c5 | 7.6521e-18
f(P) c6 | -1.1764e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1036
k0 calibration window | 2022-08-22 20:12 | 2022-08-23 23:36
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 3
Ik outlier % | 0.09
Ib outlier count | 1
Ib outlier % | 0.03
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1036_DURA_2 (2).dat | MSC1036_DURA (2).dat
median k0 | -1.3801 | -1.3798 | Volts
std k0 | 6.4891e-05 | 6.517e-05 | Volts
Avg Vk | -0.8383 | -0.8383 | Volts
Avg std Vk | 170.3 | 173.5 | uVolts
median Ik: | -8.85 | -8.54 | nano amps
std Ik: | 7.86 | 7.93 | nano amps
median Ib: | 1.08 | 0.16 | nano amps
std Ib: | 18.09 | 17.56 | nano amps

## GDF026_Calibration_20220824
PH CALIBRATION FILE GENERATED ON: 08/24/2022 10:56
GDF 026
MSC 1033
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 026
8, number of calibration coefficients
-1.38899, =k0 Pon; -1.38883 =k0 Poff; -1.38318 =k0 HCl
-2.92045e-14,8.08649e-11,-9.27878e-08,-9.96990e-04, =k2(fP)*T
-6.14565e-06, =k3*P
1.75376e-08, =k4*P^3
-3.54504e-11, =k5*P^4
3.02832e-14, =k6*P^5
-1.15112e-17, =k7*P^6
1.64992e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
06/23-06/28/2022 | A |  | -0.085 |  |  |  | -2.9205e-14 | 8.0865e-11 | -9.2788e-08 | -0.00099699 | -1.3832 | 0.9995 |  |  | 1.6499e-21 | -1.1511e-17 | 3.0283e-14 | -3.545e-11 | 1.7538e-08 | -6.1457e-06 | -1.3829 | 0.9876 | GDF |  | 26 | 07/20-07/25/2022 |  | -1.3888 | 0.0001882 | -1.389 | 0.00018724
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 26
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF026_062122_chanA.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-06-23 05:03 | 2022-06-28 07:25
Data subset: | ASCENDING
Vrs drift rate: | -0.08478 | mV/day
Vk drift rate: | -4.6948 | mV/day
Avg std Vk | 224.401 | uVolts
Average k2: | -0.0010314 | V/degC
Std dev k2: | 2.7279e-05 | V/degC
k2 f(P) c0 | -0.00099699
k2 f(P) c1 | -9.2788e-08
k2 f(P) c2 | 8.0865e-11
k2 f(P) c3 | -2.9204e-14
k0 HCL: | -1.3832 | Volts
Avg Ik: | 0.95 | nano amps
Std dev Ik: | 7.38 | nano amps
Avg Ib: | 15.29 | nano amps
Std dev Ib: | 6 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.3 | -0.00099565 | 5.4932e-06 | -1.3832 | 7.7445e-05 | 0.9995 | 17 | 2.1 | 19.9
39.4 | -0.0010009 | 5.7344e-06 | -1.3831 | 8.4348e-05 | 0.9993 | 22 | 2.1 | 19.9
62 | -0.00099675 | 3.7924e-06 | -1.3833 | 5.6989e-05 | 0.9998 | 18 | 2.1 | 19.9
86.9 | -0.0010053 | 3.7036e-06 | -1.3833 | 5.4205e-05 | 0.9998 | 19 | 2.1 | 19.9
124 | -0.0010109 | 3.1959e-06 | -1.3834 | 4.5046e-05 | 0.9997 | 37 | 2.1 | 19.9
172.3 | -0.0010145 | 3.3712e-06 | -1.3835 | 4.7203e-05 | 0.9996 | 38 | 2.1 | 19.9
249.9 | -0.0010169 | 2.937e-06 | -1.3838 | 4.1817e-05 | 0.9994 | 72 | 2.1 | 19.9
348 | -0.0010194 | 3.0437e-06 | -1.3841 | 4.4996e-05 | 0.9995 | 59 | 2.1 | 19.9
445.4 | -0.0010248 | 3.1205e-06 | -1.3844 | 4.7091e-05 | 0.9995 | 57 | 2.1 | 19.9
549.4 | -0.0010327 | 4.3298e-06 | -1.3846 | 6.5868e-05 | 0.9991 | 56 | 2.2 | 19.9
642.2 | -0.00104 | 4.7503e-06 | -1.3848 | 7.5343e-05 | 0.9991 | 44 | 2.1 | 19.9
748.2 | -0.0010295 | 3.5087e-06 | -1.3855 | 4.8572e-05 | 0.9994 | 56 | 2.1 | 19.9
903.1 | -0.001028 | 3.6911e-06 | -1.3862 | 5.3131e-05 | 0.9985 | 115 | 2.1 | 19.9
1103.6 | -0.0010322 | 4.6418e-06 | -1.3871 | 6.7271e-05 | 0.9976 | 120 | 2.1 | 19.9
1299.6 | -0.0010475 | 3.008e-06 | -1.3874 | 4.0727e-05 | 0.9991 | 114 | 2.1 | 19.9
1497.8 | -0.0010535 | 2.5548e-06 | -1.3875 | 3.5875e-05 | 0.9993 | 128 | 2.2 | 19.9
1701.2 | -0.001072 | 2.5682e-06 | -1.3875 | 3.7019e-05 | 0.9993 | 119 | 2.2 | 19.9
1903.8 | -0.0010845 | 2.277e-06 | -1.3871 | 3.2554e-05 | 0.9995 | 122 | 2.2 | 19.9
2035.6 | -0.0010917 | 4.7309e-06 | -1.3868 | 7.2595e-05 | 0.9994 | 32 | 2.2 | 19.9
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 3.1147e-08
f(P) R*R: | 0.9876
f(P) c0 | -1.3829
f(P) c1 | -6.1457e-06
f(P) c2 | 1.7538e-08
f(P) c3 | -3.545e-11
f(P) c4 | 3.0283e-14
f(P) c5 | -1.1511e-17
f(P) c6 | 1.6499e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1033
k0 calibration window | 2022-07-20 18:19 | 2022-07-25 19:00
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 2
Ik outlier % | 0.06
Ib outlier count | 2
Ib outlier % | 0.06
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1033_DURA_2.dat | MSC1033_DURA.dat
median k0 | -1.3889 | -1.3888 | Volts
std k0 | 0.00018724 | 0.0001882 | Volts
Avg Vk | -0.843 | -0.843 | Volts
Avg std Vk | 53.1 | 52.8 | uVolts
median Ik: | -10.6 | -10.51 | nano amps
std Ik: | 2.86 | 2.88 | nano amps
median Ib: | 0.05 | -0.33 | nano amps
std Ib: | 5.38 | 5.64 | nano amps

## GDF027_Calibration_20220824
PH CALIBRATION FILE GENERATED ON: 08/24/2022 10:57
GDF 027
MSC 1035
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 027
8, number of calibration coefficients
-1.40223, =k0 Pon; -1.40206 =k0 Poff; -1.39545 =k0 HCl
-5.19168e-14,1.61500e-10,-1.63383e-07,-9.73364e-04, =k2(fP)*T
-7.80997e-07, =k3*P
-1.50938e-08, =k4*P^3
3.54895e-11, =k5*P^4
-3.80767e-14, =k6*P^5
1.83700e-17, =k7*P^6
-3.19695e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
06/12-06/17/2022 | A |  | -0.157 |  |  |  | -5.1917e-14 | 1.615e-10 | -1.6338e-07 | -0.00097336 | -1.3955 | 0.9995 |  |  | -3.197e-21 | 1.837e-17 | -3.8077e-14 | 3.549e-11 | -1.5094e-08 | -7.81e-07 | -1.3954 | 0.9886 | GDF |  | 27 | 07/20-07/25/2022 |  | -1.4021 | 0.00016529 | -1.4022 | 0.00016465
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 27
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF027_061022_chanA.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-06-12 03:16 | 2022-06-17 05:38
Data subset: | ASCENDING
Vrs drift rate: | -0.15704 | mV/day
Vk drift rate: | 12.2913 | mV/day
Avg std Vk | 216.5869 | uVolts
Average k2: | -0.0010163 | V/degC
Std dev k2: | 2.6307e-05 | V/degC
k2 f(P) c0 | -0.00097336
k2 f(P) c1 | -1.6338e-07
k2 f(P) c2 | 1.615e-10
k2 f(P) c3 | -5.1917e-14
k0 HCL: | -1.3954 | Volts
Avg Ik: | 1.92 | nano amps
Std dev Ik: | 6.28 | nano amps
Avg Ib: | 15.09 | nano amps
Std dev Ib: | 5.21 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.2 | -0.00097453 | 6.0185e-06 | -1.3954 | 9.1126e-05 | 0.9995 | 16 | 2.1 | 19.9
38.2 | -0.0009806 | 4.986e-06 | -1.3954 | 7.1593e-05 | 0.9994 | 26 | 2.1 | 19.8
61.6 | -0.00098569 | 5.698e-06 | -1.3955 | 8.6364e-05 | 0.9993 | 22 | 2.1 | 19.8
86.8 | -0.0009835 | 5.3225e-06 | -1.3956 | 7.4193e-05 | 0.9994 | 22 | 2.1 | 19.8
125.2 | -0.00099081 | 4.7823e-06 | -1.3958 | 7.3338e-05 | 0.9993 | 32 | 2.1 | 19.8
175.9 | -0.00099752 | 5.2854e-06 | -1.3959 | 7.8554e-05 | 0.9991 | 35 | 2.1 | 19.8
250 | -0.0010067 | 4.3073e-06 | -1.3961 | 6.414e-05 | 0.9989 | 63 | 2.1 | 19.8
350.6 | -0.0010128 | 4.8312e-06 | -1.3965 | 7.5429e-05 | 0.9988 | 56 | 2 | 19.8
450.3 | -0.0010135 | 4.3092e-06 | -1.3969 | 6.422e-05 | 0.9989 | 62 | 2 | 19.8
550.8 | -0.0010317 | 5.4212e-06 | -1.397 | 8.5235e-05 | 0.9985 | 56 | 10 | 19.8
651 | -0.0010252 | 4.3278e-06 | -1.3975 | 6.7811e-05 | 0.9991 | 54 | 2.1 | 19.8
753.3 | -0.0010249 | 2.7286e-06 | -1.3979 | 3.8763e-05 | 0.9996 | 64 | 2.1 | 19.9
895.6 | -0.0010339 | 3.0517e-06 | -1.3983 | 4.5214e-05 | 0.999 | 117 | 2.2 | 19.9
1099.6 | -0.0010264 | 4.4028e-06 | -1.399 | 6.9719e-05 | 0.9982 | 101 | 2.2 | 19.9
1300.7 | -0.0010089 | 5.0303e-06 | -1.4002 | 7.4516e-05 | 0.9971 | 118 | 2.2 | 19.9
1501.1 | -0.0010332 | 2.2052e-06 | -1.4005 | 3.1402e-05 | 0.9994 | 125 | 2.3 | 19.9
1697.6 | -0.0010502 | 2.6078e-06 | -1.4001 | 3.746e-05 | 0.9993 | 119 | 2.3 | 19.9
1904.6 | -0.0010623 | 2.5262e-06 | -1.3997 | 3.5634e-05 | 0.9993 | 128 | 2.3 | 19.9
2034.7 | -0.0010672 | 3.7636e-06 | -1.3994 | 5.2434e-05 | 0.9995 | 39 | 2.3 | 19.9
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 3.3092e-08
f(P) R*R: | 0.9886
f(P) c0 | -1.3954
f(P) c1 | -7.81e-07
f(P) c2 | -1.5094e-08
f(P) c3 | 3.5489e-11
f(P) c4 | -3.8077e-14
f(P) c5 | 1.837e-17
f(P) c6 | -3.197e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1035
k0 calibration window | 2022-07-20 14:58 | 2022-07-25 19:03
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 12
Ik outlier % | 0.32
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1035_DURA_2.dat | MSC1035_DURA.dat
median k0 | -1.4022 | -1.402 | Volts
std k0 | 0.00016465 | 0.00016529 | Volts
Avg Vk | -0.9396 | -0.9396 | Volts
Avg std Vk | 41.9 | 42 | uVolts
median Ik: | -3.91 | -3.88 | nano amps
std Ik: | 3 | 2.99 | nano amps
median Ib: | -0.07 | -0.04 | nano amps
std Ib: | 4.24 | 4.49 | nano amps

## GDF028_Calibration_20220824
PH CALIBRATION FILE GENERATED ON: 08/24/2022 11:00
GDF 028
MSC 1037
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 028
8, number of calibration coefficients
-1.39528, =k0 Pon; -1.39521 =k0 Poff; -1.38875 =k0 HCl
-4.06938e-14,1.13631e-10,-1.16104e-07,-9.86186e-04, =k2(fP)*T
-1.92483e-06, =k3*P
-5.74388e-09, =k4*P^3
7.75562e-12, =k5*P^4
-5.31577e-15, =k6*P^5
1.75704e-18, =k7*P^6
-1.83620e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
07/09-07/14/2022 | A |  | -0.086 |  |  |  | -4.0694e-14 | 1.1363e-10 | -1.161e-07 | -0.00098619 | -1.3887 | 0.9994 |  |  | -1.8362e-22 | 1.757e-18 | -5.3158e-15 | 7.7556e-12 | -5.7439e-09 | -1.9248e-06 | -1.3888 | 0.9919 | GDF |  | 28 | 08/10-08/16/2022 |  | -1.3952 | 0.00016323 | -1.3953 | 0.00016073
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 28
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF028_070722_chanA.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-07-09 06:30 | 2022-07-14 08:48
Data subset: | ASCENDING
Vrs drift rate: | -0.08627 | mV/day
Vk drift rate: | 0.28582 | mV/day
Avg std Vk | 346.4954 | uVolts
Average k2: | -0.0010235 | V/degC
Std dev k2: | 2.8318e-05 | V/degC
k2 f(P) c0 | -0.00098619
k2 f(P) c1 | -1.161e-07
k2 f(P) c2 | 1.1363e-10
k2 f(P) c3 | -4.0694e-14
k0 HCL: | -1.3887 | Volts
Avg Ik: | 9.34 | nano amps
Std dev Ik: | 7.8 | nano amps
Avg Ib: | 20.62 | nano amps
Std dev Ib: | 7.12 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23 | -0.00099202 | 6.2764e-06 | -1.3887 | 8.4634e-05 | 0.9994 | 17 | 2.3 | 19.9
38.9 | -0.00098903 | 6.9185e-06 | -1.3888 | 0.00010529 | 0.9992 | 18 | 2.3 | 19.9
60.8 | -0.00098789 | 5.8512e-06 | -1.389 | 8.362e-05 | 0.9994 | 19 | 2.3 | 19.9
85.1 | -0.00099588 | 6.5208e-06 | -1.389 | 9.0878e-05 | 0.9992 | 20 | 2.2 | 19.9
123.1 | -0.00099416 | 4.9327e-06 | -1.3892 | 7.3499e-05 | 0.9992 | 34 | 2.2 | 19.9
173.7 | -0.0010049 | 4.8663e-06 | -1.3893 | 7.2274e-05 | 0.9992 | 36 | 2.2 | 19.9
249.5 | -0.0010109 | 4.792e-06 | -1.3895 | 7.4086e-05 | 0.9988 | 56 | 2.2 | 19.9
346.5 | -0.0010214 | 6.9908e-06 | -1.3898 | 0.00010583 | 0.9978 | 48 | 10 | 19.9
452.2 | -0.0010164 | 6.6077e-06 | -1.3903 | 9.8803e-05 | 0.9979 | 51 | 2 | 19.9
548 | -0.0010299 | 6.378e-06 | -1.3905 | 9.9742e-05 | 0.9982 | 49 | 10 | 19.9
647.4 | -0.0010201 | 4.0405e-06 | -1.3911 | 5.8508e-05 | 0.999 | 64 | 2.2 | 19.9
752.1 | -0.0010242 | 3.8588e-06 | -1.3916 | 5.4871e-05 | 0.9991 | 62 | 2.2 | 19.9
895.5 | -0.0010303 | 3.5763e-06 | -1.3921 | 5.236e-05 | 0.9986 | 115 | 2.2 | 19.9
1102.3 | -0.0010283 | 3.4184e-06 | -1.3928 | 5.0285e-05 | 0.9987 | 118 | 2.2 | 19.9
1300.6 | -0.0010266 | 3.6836e-06 | -1.3935 | 5.3186e-05 | 0.9985 | 115 | 2.2 | 19.9
1505.1 | -0.0010411 | 3.7564e-06 | -1.3941 | 5.3478e-05 | 0.9985 | 114 | 2.3 | 19.9
1699 | -0.0010641 | 2.3609e-06 | -1.3944 | 3.2774e-05 | 0.9994 | 120 | 2.3 | 19.9
1900.8 | -0.0010832 | 2.1916e-06 | -1.3942 | 3.1126e-05 | 0.9995 | 125 | 2.3 | 19.9
2035 | -0.0010864 | 3.7818e-06 | -1.3942 | 5.2342e-05 | 0.9996 | 37 | 2.3 | 19.9
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.9898e-08
f(P) R*R: | 0.9919
f(P) c0 | -1.3888
f(P) c1 | -1.9248e-06
f(P) c2 | -5.7439e-09
f(P) c3 | 7.7556e-12
f(P) c4 | -5.3158e-15
f(P) c5 | 1.757e-18
f(P) c6 | -1.8362e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1037
k0 calibration window | 2022-08-10 14:54 | 2022-08-16 21:59
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 6
Ik outlier % | 0.07
Ib outlier count | 1
Ib outlier % | 0.01
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1037_DURA_2.dat | MSC1037_DURA.dat
median k0 | -1.3953 | -1.3952 | Volts
std k0 | 0.00016073 | 0.00016323 | Volts
Avg Vk | -0.8704 | -0.8704 | Volts
Avg std Vk | 56.9 | 57.6 | uVolts
median Ik: | -5.25 | -5.38 | nano amps
std Ik: | 2.86 | 2.83 | nano amps
median Ib: | 0.14 | 0.02 | nano amps
std Ib: | 5.53 | 5.62 | nano amps

## GDF9_Calibration_20220824
PH CALIBRATION FILE GENERATED ON: 08/24/2022 15:24
GDF 9
MSC 1034
APEX ID 17465
** START CONFIG FILE COPY BLOCK **
Durafet SN = 17465, APEX# 9
8, number of calibration coefficients
-1.40625, =k0 Pon; -1.40605 =k0 Poff; -1.40077 =k0 HCl
-3.95502e-14,4.76247e-11,-5.90569e-09,-1.03936e-03, =k2(fP)*T
2.02476e-06, =k3*P
-2.26849e-08, =k4*P^3
5.32741e-11, =k5*P^4
-6.41658e-14, =k6*P^5
3.35486e-17, =k7*P^6
-6.16815e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
05/18-05/21/2022 |  |  | -0.033 |  |  |  | -3.955e-14 | 4.7625e-11 | -5.9057e-09 | -0.0010394 | -1.4008 | 0.9998 |  |  | -6.1682e-21 | 3.3549e-17 | -6.4166e-14 | 5.3274e-11 | -2.2685e-08 | 2.0248e-06 | -1.4007 | 0.9978 | GDF |  | 9 | 07/20-07/25/2022 |  | -1.4061 | 0.0001994 | -1.4062 | 0.00019955
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 9
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF9_051722_Ch F.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-05-18 19:22 | 2022-05-21 06:04
Data subset: | ASCENDING
Vrs drift rate: | -0.03308 | mV/day
Vk drift rate: | 16.1406 | mV/day
Avg std Vk | 301.0769 | uVolts
Average k2: | -0.0010563 | V/degC
Std dev k2: | 4.1768e-05 | V/degC
k2 f(P) c0 | -0.0010394
k2 f(P) c1 | -5.9057e-09
k2 f(P) c2 | 4.7625e-11
k2 f(P) c3 | -3.955e-14
k0 HCL: | -1.4008 | Volts
Avg Ik: | -9.08 | nano amps
Std dev Ik: | 15.7 | nano amps
Avg Ib: | 5.88 | nano amps
Std dev Ib: | 10.86 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
7 | -0.0010375 | 3.7459e-06 | -1.4008 | 5.1446e-05 | 0.9998 | 16 | 2.7 | 20.2
14.5 | -0.0010323 | 3.9655e-06 | -1.4007 | 5.0682e-05 | 0.9997 | 21 | 2.7 | 20.2
23.3 | -0.0010396 | 5.4278e-06 | -1.4007 | 6.4623e-05 | 0.9995 | 19 | 2.7 | 20.2
40.2 | -0.0010388 | 4.0422e-06 | -1.4006 | 5.6108e-05 | 0.9997 | 24 | 2.7 | 20.2
61.8 | -0.0010401 | 4.0448e-06 | -1.4006 | 5.9205e-05 | 0.9997 | 21 | 5.6 | 20.2
85.1 | -0.00104 | 5.9619e-06 | -1.4007 | 7.4028e-05 | 0.9995 | 16 | 5.6 | 20.2
126.1 | -0.0010401 | 3.0597e-06 | -1.4008 | 4.7753e-05 | 0.9997 | 34 | 5.6 | 20.2
176.5 | -0.001043 | 4.7927e-06 | -1.4008 | 6.6123e-05 | 0.9995 | 24 | 2.7 | 20.2
247.6 | -0.001039 | 3.5123e-06 | -1.401 | 4.9793e-05 | 0.9995 | 48 | 5.6 | 20.2
351.2 | -0.0010421 | 3.3686e-06 | -1.4012 | 4.982e-05 | 0.9995 | 53 | 5.6 | 20.2
455.5 | -0.0010446 | 3.8378e-06 | -1.4015 | 5.5016e-05 | 0.9994 | 50 | 2.6 | 20.2
550.5 | -0.0010355 | 3.8079e-06 | -1.4019 | 5.1667e-05 | 0.9994 | 49 | 2.6 | 20.2
648.6 | -0.0010376 | 4.5528e-06 | -1.4024 | 6.0985e-05 | 0.999 | 54 | 2.6 | 20.2
746.5 | -0.001034 | 3.8165e-06 | -1.403 | 5.3803e-05 | 0.9994 | 45 | 2.6 | 20.2
895.3 | -0.0010199 | 7.1564e-06 | -1.4042 | 0.00010019 | 0.9955 | 93 | 2.6 | 20.2
1101 | -0.0010274 | 1.0049e-05 | -1.406 | 0.00015041 | 0.9915 | 92 | 2.6 | 20.2
1300.5 | -0.0010541 | 9.5353e-06 | -1.4079 | 0.00013605 | 0.9932 | 86 | 2.6 | 20.2
1499.8 | -0.0010892 | 4.5309e-06 | -1.4091 | 6.2908e-05 | 0.9984 | 95 | 2.7 | 20.2
1689.1 | -0.0011047 | 3.0085e-06 | -1.4094 | 3.7723e-05 | 0.9995 | 72 | 2.7 | 20.2
1924.7 | -0.0011632 | 1.7386e-05 | -1.4092 | 8.9008e-05 | 0.9964 | 18 | 2.7 | 5.7
2039.2 | -0.0011797 | 3.7784e-05 | -1.4091 | 0.00016686 | 0.9959 | 6 | 2.7 | 5.6
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.4748e-08
f(P) R*R: | 0.9978
f(P) c0 | -1.4007
f(P) c1 | 2.0248e-06
f(P) c2 | -2.2685e-08
f(P) c3 | 5.3274e-11
f(P) c4 | -6.4166e-14
f(P) c5 | 3.3549e-17
f(P) c6 | -6.1682e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1034
k0 calibration window | 2022-07-20 17:33 | 2022-07-25 18:30
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 84
Ik outlier % | 2.31
Ib outlier count | 64
Ib outlier % | 1.76
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1034_DURA_2.dat | MSC1034_DURA.dat
median k0 | -1.4062 | -1.406 | Volts
std k0 | 0.00019955 | 0.00019941 | Volts
Avg Vk | -0.9353 | -0.9353 | Volts
Avg std Vk | 45.5 | 45.9 | uVolts
median Ik: | -5.1 | -5.31 | nano amps
std Ik: | 6.04 | 6.28 | nano amps
median Ib: | -0.44 | -0.4 | nano amps
std Ib: | 7.6 | 7.91 | nano amps

## GDF31_Calibration_20220922
PH CALIBRATION FILE GENERATED ON: 09/22/2022 14:03
GDF 31
MSC 1038
APEX ID 21479
** START CONFIG FILE COPY BLOCK **
Durafet SN = 21479, APEX# 31
8, number of calibration coefficients
-1.38017, =k0 Pon; -1.38008 =k0 Poff; -1.37116 =k0 HCl
5.74178e-17,-2.76773e-13,4.04347e-10,-1.98665e-07,-1.05986e-03, =k2(fP)*T
-8.50036e-06, =k3*P
1.44779e-08, =k4*P^3
-2.50179e-11, =k5*P^4
2.05569e-14, =k6*P^5
-7.64489e-18, =k7*P^6
1.07814e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
08/20-08/25/2022 |  |  | -0.1 |  |  | 5.7418e-17 | -2.7677e-13 | 4.0435e-10 | -1.9867e-07 | -0.0010599 | -1.3712 | 0.9998 |  |  | 1.0781e-21 | -7.6449e-18 | 2.0557e-14 | -2.5018e-11 | 1.4478e-08 | -8.5004e-06 | -1.3709 | 0.9977 | GDF |  | 31 | 09/19-09/22/2022 |  | -1.3801 | 0.00032129 | -1.3802 | 0.00032189
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 31
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF31_081822_ChC.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-08-20 06:22 | 2022-08-25 09:14
Data subset: | ASCENDING
Vrs drift rate: | -0.09998 | mV/day
Vk drift rate: | 1.1293 | mV/day
Avg std Vk | 115.5401 | uVolts
Average k2: | -0.001087 | V/degC
Std dev k2: | 1.8873e-05 | V/degC
k2 f(P) c0 | -0.0010599
k2 f(P) c1 | -1.9867e-07
k2 f(P) c2 | 4.0435e-10
k2 f(P) c3 | -2.7677e-13
k2 f(P) c4 | 5.7418e-17
k0 HCL: | -1.3712 | Volts
Avg Ik: | -73.63 | nano amps
Std dev Ik: | 86.65 | nano amps
Avg Ib: | -9.86 | nano amps
Std dev Ib: | 77.59 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
24.7 | -0.0010664 | 3.065e-06 | -1.3712 | 4.2066e-05 | 0.9998 | 29 | 1.9 | 19.7
39 | -0.0010679 | 2.2371e-06 | -1.3712 | 3.2171e-05 | 0.9998 | 41 | 1.9 | 19.7
61.4 | -0.0010714 | 2.2002e-06 | -1.3714 | 3.137e-05 | 0.9999 | 35 | 1.9 | 19.7
86.7 | -0.0010752 | 2.5334e-06 | -1.3715 | 3.5677e-05 | 0.9999 | 29 | 1.9 | 19.7
124.3 | -0.0010779 | 2.8196e-06 | -1.3718 | 3.9378e-05 | 0.9997 | 51 | 1.9 | 19.7
175.2 | -0.00108 | 2.8306e-06 | -1.3721 | 4.0403e-05 | 0.9997 | 44 | 1.9 | 19.7
248.9 | -0.0010868 | 2.7327e-06 | -1.3725 | 3.7613e-05 | 0.9995 | 88 | 1.9 | 19.7
348 | -0.0010877 | 2.4255e-06 | -1.373 | 3.3763e-05 | 0.9996 | 86 | 1.9 | 19.7
448.4 | -0.0010898 | 2.4914e-06 | -1.3734 | 3.4783e-05 | 0.9996 | 85 | 1.9 | 19.7
548.7 | -0.0010886 | 2.7394e-06 | -1.3739 | 3.9457e-05 | 0.9995 | 76 | 1.9 | 19.7
650.5 | -0.0010872 | 2.6944e-06 | -1.3743 | 3.9833e-05 | 0.9996 | 73 | 1.9 | 19.7
747.2 | -0.0010865 | 2.4887e-06 | -1.3747 | 3.4756e-05 | 0.9996 | 79 | 2 | 19.7
899.6 | -0.0010747 | 3.3438e-06 | -1.3755 | 4.6618e-05 | 0.9985 | 157 | 2 | 19.7
1098.8 | -0.0010689 | 2.673e-06 | -1.3765 | 3.7149e-05 | 0.999 | 159 | 2 | 19.7
1297.7 | -0.0010767 | 1.894e-06 | -1.3769 | 2.6236e-05 | 0.9995 | 156 | 2 | 19.7
1494.3 | -0.0010887 | 2.2627e-06 | -1.3773 | 3.269e-05 | 0.9994 | 146 | 2 | 19.7
1701.8 | -0.0011136 | 1.2561e-06 | -1.3773 | 1.7628e-05 | 0.9998 | 149 | 2 | 19.7
1902.2 | -0.0011302 | 1.0615e-06 | -1.377 | 1.5202e-05 | 0.9999 | 143 | 2 | 19.7
2034.2 | -0.0011348 | 1.8958e-06 | -1.3769 | 2.6627e-05 | 0.9999 | 49 | 2.1 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.0031e-08
f(P) R*R: | 0.9977
f(P) c0 | -1.3709
f(P) c1 | -8.5004e-06
f(P) c2 | 1.4478e-08
f(P) c3 | -2.5018e-11
f(P) c4 | 2.0557e-14
f(P) c5 | -7.6449e-18
f(P) c6 | 1.0781e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1038
k0 calibration window | 2022-09-19 11:41 | 2022-09-22 16:05
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 79
Ik outlier % | 3.44
Ib outlier count | 49
Ib outlier % | 2.14
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1038_DURA_2.dat | MSC1038_DURA.dat
median k0 | -1.3802 | -1.3801 | Volts
std k0 | 0.00032189 | 0.00032129 | Volts
Avg Vk | -0.9098 | -0.9099 | Volts
Avg std Vk | 61.7 | 61.3 | uVolts
median Ik: | -1.33 | -1.66 | nano amps
std Ik: | 11.88 | 10.75 | nano amps
median Ib: | -9.11 | -8.47 | nano amps
std Ib: | 13.73 | 13.53 | nano amps

## GDF32_Calibration_20220922
PH CALIBRATION FILE GENERATED ON: 09/22/2022 14:06
GDF 32
MSC 1039
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 32
8, number of calibration coefficients
-1.36621, =k0 Pon; -1.36606 =k0 Poff; -1.35872 =k0 HCl
-1.48602e-11,5.61684e-09,-1.00412e-03, =k2(fP)*T
-1.66191e-05, =k3*P
4.29405e-08, =k4*P^3
-4.27365e-11, =k5*P^4
1.66403e-14, =k6*P^5
-1.23766e-18, =k7*P^6
-4.07717e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
08/20-08/25/2022 |  |  | -0.123 |  |  |  |  | -1.486e-11 | 5.6168e-09 | -0.0010041 | -1.3587 | 0.9995 |  |  | -4.0772e-22 | -1.2377e-18 | 1.664e-14 | -4.2737e-11 | 4.294e-08 | -1.6619e-05 | -1.3581 | 0.9406 | GDF |  | 32 | 09/16-09/21/2022 |  | -1.3661 | 0.00040014 | -1.3662 | 0.00039247
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 32
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF32_081822_ChD.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-08-20 06:22 | 2022-08-25 09:14
Data subset: | ASCENDING
Vrs drift rate: | -0.12277 | mV/day
Vk drift rate: | 0.04896 | mV/day
Avg std Vk | 137.8855 | uVolts
Average k2: | -0.0010143 | V/degC
Std dev k2: | 1.6652e-05 | V/degC
k2 f(P) c0 | -0.0010041
k2 f(P) c1 | 5.6168e-09
k2 f(P) c2 | -1.486e-11
k0 HCL: | -1.3587 | Volts
Avg Ik: | -38.84 | nano amps
Std dev Ik: | 15.02 | nano amps
Avg Ib: | -0.93 | nano amps
Std dev Ib: | 16.73 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
24.4 | -0.00099814 | 4.4639e-06 | -1.3587 | 6.6001e-05 | 0.9995 | 25 | 2.2 | 19.7
38.6 | -0.001007 | 4.0723e-06 | -1.3587 | 6.4083e-05 | 0.9995 | 35 | 9.9 | 19.7
61.2 | -0.00099846 | 3.2026e-06 | -1.359 | 4.5085e-05 | 0.9996 | 37 | 2.1 | 19.7
86.6 | -0.0010019 | 4.1642e-06 | -1.3592 | 5.8237e-05 | 0.9995 | 30 | 2.1 | 19.7
124.5 | -0.0010024 | 3.3071e-06 | -1.3594 | 4.5617e-05 | 0.9994 | 54 | 2.1 | 19.7
175.4 | -0.0010121 | 5.133e-06 | -1.3597 | 7.215e-05 | 0.9988 | 47 | 2 | 19.7
247.6 | -0.0010165 | 2.8847e-06 | -1.36 | 4.0927e-05 | 0.9993 | 83 | 2 | 19.7
348 | -0.00099197 | 2.7012e-06 | -1.3605 | 3.9756e-05 | 0.9994 | 77 | 2 | 19.7
447.1 | -0.00099997 | 2.8018e-06 | -1.3602 | 4.3591e-05 | 0.9995 | 66 | 2 | 19.7
550.3 | -0.001016 | 2.2646e-06 | -1.3598 | 3.4396e-05 | 0.9997 | 68 | 2.1 | 19.7
649 | -0.0010093 | 2.7163e-06 | -1.3596 | 3.9683e-05 | 0.9995 | 75 | 2.1 | 19.7
748.1 | -0.0010062 | 2.222e-06 | -1.3595 | 3.0884e-05 | 0.9996 | 81 | 2.1 | 19.7
898.4 | -0.0010077 | 1.6201e-06 | -1.3595 | 2.2447e-05 | 0.9996 | 162 | 2.1 | 19.7
1097.7 | -0.0010164 | 2.3131e-06 | -1.3597 | 3.3019e-05 | 0.9992 | 151 | 2.1 | 19.7
1299 | -0.0010195 | 2.3598e-06 | -1.3601 | 3.3373e-05 | 0.9992 | 150 | 2.1 | 19.7
1498.6 | -0.0010283 | 1.7087e-06 | -1.3604 | 2.3823e-05 | 0.9996 | 159 | 2.1 | 19.7
1700.4 | -0.001038 | 1.7608e-06 | -1.3606 | 2.435e-05 | 0.9996 | 155 | 2.2 | 19.7
1900.2 | -0.0010457 | 1.6458e-06 | -1.3608 | 2.2834e-05 | 0.9996 | 153 | 2.2 | 19.7
2033.7 | -0.0010565 | 2.0151e-06 | -1.3609 | 2.8194e-05 | 0.9998 | 50 | 2.2 | 19.7
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.0144e-08
f(P) R*R: | 0.9406
f(P) c0 | -1.3581
f(P) c1 | -1.6619e-05
f(P) c2 | 4.294e-08
f(P) c3 | -4.2736e-11
f(P) c4 | 1.664e-14
f(P) c5 | -1.2377e-18
f(P) c6 | -4.0772e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1039
k0 calibration window | 2022-09-16 20:22 | 2022-09-21 21:46
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 0
Ik outlier % | 0
Ib outlier count | 2
Ib outlier % | 0.05
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1039_DURA_2.dat | MSC1039_DURA.dat
median k0 | -1.3661 | -1.3659 | Volts
std k0 | 0.00039247 | 0.00040014 | Volts
Avg Vk | -0.8742 | -0.8742 | Volts
Avg std Vk | 44.2 | 43.9 | uVolts
median Ik: | -10.16 | -9.97 | nano amps
std Ik: | 3.75 | 3.79 | nano amps
median Ib: | 0.24 | 0.16 | nano amps
std Ib: | 4.43 | 4.44 | nano amps

## GDF015_Calibration_20221005
PH CALIBRATION FILE GENERATED ON: 10/05/2022 09:40
GDF 015
MSC XXXX
APEX ID XXXX
** START CONFIG FILE COPY BLOCK **
Durafet SN = XXXX, APEX# 015
8, number of calibration coefficients
NO K0 DATA WAS PEOCESSED YET OR LOADED INTO GUI!!!!
8.09848e-12,-5.43762e-08,-1.69334e-03, =k2(fP)*T
-4.47064e-06, =k3*P
-2.15526e-09, =k4*P^3
1.39995e-11, =k5*P^4
-1.47649e-14, =k6*P^5
6.26038e-18, =k7*P^6
-9.38381e-22, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
08/23-08/28/2022 | A |  | -0.018 |  |  |  |  | 8.0985e-12 | -5.4376e-08 | -0.0016933 | -1.0447 | 0.9998 |  |  | -9.3838e-22 | 6.2604e-18 | -1.4765e-14 | 1.4e-11 | -2.1553e-09 | -4.4706e-06 | -1.0445 | 0.9275 | GDF |  | 15
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 15
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF015_RFD_082222_chanA.csv
PT test channel : | A
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-08-23 20:16 | 2022-08-28 22:38
Data subset: | ASCENDING
Vrs drift rate: | -0.01801 | mV/day
Vk drift rate: | 0.51807 | mV/day
Avg std Vk | 362.6287 | uVolts
Average k2: | -0.0017254 | V/degC
Std dev k2: | 2.5992e-05 | V/degC
k2 f(P) c0 | -0.0016933
k2 f(P) c1 | -5.4376e-08
k2 f(P) c2 | 8.0985e-12
k0 HCL: | -1.0447 | Volts
Avg Ik: | 6.15 | nano amps
Std dev Ik: | 7.06 | nano amps
Avg Ib: | 17.48 | nano amps
Std dev Ib: | 8.27 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
24 | -0.0016879 | 4.7926e-06 | -1.0447 | 7.0633e-05 | 0.9998 | 22 | 2.1 | 19.9
39 | -0.0016945 | 3.2538e-06 | -1.0446 | 4.6717e-05 | 0.9998 | 43 | 2.1 | 19.9
61.8 | -0.0016946 | 3.8222e-06 | -1.0448 | 5.4598e-05 | 0.9998 | 41 | 2.1 | 19.9
86.5 | -0.0016973 | 3.2754e-06 | -1.0449 | 4.6787e-05 | 0.9999 | 41 | 2.1 | 19.9
123.9 | -0.001699 | 2.2429e-06 | -1.0451 | 3.1358e-05 | 0.9999 | 77 | 2.2 | 19.9
174.1 | -0.0017035 | 2.5386e-06 | -1.0453 | 3.5234e-05 | 0.9998 | 74 | 2.2 | 19.9
248.4 | -0.0017097 | 2.4102e-06 | -1.0456 | 3.2392e-05 | 0.9997 | 131 | 2.2 | 19.9
349.8 | -0.0017151 | 2.1387e-06 | -1.0458 | 2.894e-05 | 0.9998 | 133 | 2.2 | 19.9
448.2 | -0.0017171 | 1.8911e-06 | -1.0461 | 2.5775e-05 | 0.9998 | 131 | 2.2 | 19.9
550.8 | -0.0017267 | 2.0133e-06 | -1.0462 | 2.7582e-05 | 0.9998 | 134 | 2.2 | 19.9
650.4 | -0.0017291 | 1.8282e-06 | -1.0464 | 2.5448e-05 | 0.9998 | 138 | 2.2 | 19.9
749.7 | -0.0017353 | 1.6322e-06 | -1.0465 | 2.2923e-05 | 0.9999 | 134 | 2.2 | 19.9
899.5 | -0.0017367 | 1.2656e-06 | -1.0465 | 1.7854e-05 | 0.9999 | 266 | 2.2 | 19.9
1098.8 | -0.0017345 | 1.3372e-06 | -1.0467 | 1.8845e-05 | 0.9998 | 263 | 2.2 | 19.9
1298.9 | -0.0017434 | 1.2678e-06 | -1.0467 | 1.7886e-05 | 0.9999 | 257 | 2.2 | 19.9
1499 | -0.0017492 | 1.7047e-06 | -1.0468 | 2.4091e-05 | 0.9998 | 261 | 2.2 | 19.9
1699.9 | -0.0017663 | 1.5671e-06 | -1.0466 | 2.2459e-05 | 0.9998 | 246 | 2.2 | 19.9
1901.1 | -0.0017709 | 1.8764e-06 | -1.0462 | 2.7062e-05 | 0.9997 | 241 | 2.3 | 19.9
2034.8 | -0.0017722 | 2.8846e-06 | -1.0459 | 4.1377e-05 | 0.9998 | 79 | 2.3 | 19.9
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.3047e-08
f(P) R*R: | 0.9275
f(P) c0 | -1.0445
f(P) c1 | -4.4706e-06
f(P) c2 | -2.1553e-09
f(P) c3 | 1.4e-11
f(P) c4 | -1.4765e-14
f(P) c5 | 6.2604e-18
f(P) c6 | -9.3838e-22
NO KO DATA FILE LOADED INTO CALIBRATION GUI!!!:
Has the k0 calibration be completed yet?:
Is the MSC # correct? Check MSC / SN match up:

## GDF029_Calibration_20221221
PH CALIBRATION FILE GENERATED ON: 12/21/2022 13:34
GDF 029
MSC 1062
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 029
8, number of calibration coefficients
-1.40135, =k0 Pon; -1.40058 =k0 Poff; -1.39702 =k0 HCl
-5.61671e-14,1.86142e-10,-2.06218e-07,-8.72717e-04, =k2(fP)*T
-1.03779e-05, =k3*P
3.45448e-08, =k4*P^3
-5.80814e-11, =k5*P^4
4.67334e-14, =k6*P^5
-1.80594e-17, =k7*P^6
2.72335e-21, =k8*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
11/05-11/10/2022 |  |  | -0.039 |  |  |  | -5.6167e-14 | 1.8614e-10 | -2.0622e-07 | -0.00087272 | -1.397 | 0.9995 |  |  | 2.7233e-21 | -1.8059e-17 | 4.6733e-14 | -5.8081e-11 | 3.4545e-08 | -1.0378e-05 | -1.3967 | 0.9425 | GDF |  | 29 | 12/05-12/08/2022 |  | -1.4006 | 0.00013008 | -1.4014 | 6.6983e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 29
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF029_110322_Ch E_new DF board on Ch F.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-11-05 05:48 | 2022-11-10 07:44
Data subset: | ASCENDING
Vrs drift rate: | -0.03868 | mV/day
Vk drift rate: | 0.46308 | mV/day
Avg std Vk | 109.3773 | uVolts
Average k2: | -0.00092971 | V/degC
Std dev k2: | 3.3726e-05 | V/degC
k2 f(P) c0 | -0.00087272
k2 f(P) c1 | -2.0622e-07
k2 f(P) c2 | 1.8614e-10
k2 f(P) c3 | -5.6167e-14
k0 HCL: | -1.397 | Volts
Avg Ik: | 7.38 | nano amps
Std dev Ik: | 1.49 | nano amps
Avg Ib: | 1.13 | nano amps
Std dev Ib: | 2.19 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
22.9 | -0.00088004 | 4.903e-06 | -1.397 | 7.1409e-05 | 0.9995 | 18 | 1.9 | 19.8
39 | -0.00088218 | 5.0816e-06 | -1.3971 | 7.051e-05 | 0.9991 | 28 | 1.9 | 19.8
61.4 | -0.00088458 | 5.4336e-06 | -1.3972 | 7.4483e-05 | 0.9992 | 24 | 1.9 | 19.8
85.9 | -0.00088523 | 4.6544e-06 | -1.3974 | 6.6857e-05 | 0.9993 | 28 | 1.9 | 19.8
125.3 | -0.00089012 | 4.2263e-06 | -1.3976 | 5.8561e-05 | 0.9989 | 49 | 1.9 | 19.8
174.8 | -0.0008931 | 5.9421e-06 | -1.3979 | 8.1772e-05 | 0.9981 | 44 | 1.9 | 19.8
248.4 | -0.00092284 | 3.1906e-06 | -1.3979 | 4.4661e-05 | 0.9989 | 90 | 1.9 | 19.8
348.7 | -0.00093505 | 3.1448e-06 | -1.3979 | 4.4144e-05 | 0.999 | 89 | 1.9 | 19.8
449.3 | -0.00094026 | 3.3037e-06 | -1.3979 | 4.5838e-05 | 0.999 | 84 | 1.9 | 19.8
546.4 | -0.00093788 | 4.2392e-06 | -1.3982 | 6.1849e-05 | 0.9984 | 80 | 1.9 | 19.8
652.6 | -0.00092742 | 4.304e-06 | -1.3987 | 6.5509e-05 | 0.9985 | 73 | 1.9 | 19.8
749.2 | -0.00095282 | 4.504e-06 | -1.3986 | 6.2681e-05 | 0.9983 | 80 | 1.9 | 19.8
902.1 | -0.00094915 | 3.4367e-06 | -1.3991 | 4.9187e-05 | 0.998 | 156 | 1.9 | 19.8
1098.3 | -0.00094519 | 3.3943e-06 | -1.3995 | 4.717e-05 | 0.9979 | 163 | 1.9 | 19.8
1300.7 | -0.00094636 | 3.2469e-06 | -1.3999 | 4.5404e-05 | 0.9981 | 165 | 1.9 | 19.8
1499.7 | -0.00095593 | 3.0216e-06 | -1.4001 | 4.2128e-05 | 0.9984 | 161 | 2 | 19.8
1702.8 | -0.00096339 | 3.1209e-06 | -1.4003 | 4.5202e-05 | 0.9985 | 149 | 2 | 19.8
1901.5 | -0.00098371 | 2.9651e-06 | -1.4001 | 4.154e-05 | 0.9985 | 162 | 2 | 19.8
2034.4 | -0.00098918 | 5.5694e-06 | -1.3999 | 8.0936e-05 | 0.9986 | 46 | 2 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 6.1594e-08
f(P) R*R: | 0.9425
f(P) c0 | -1.3967
f(P) c1 | -1.0378e-05
f(P) c2 | 3.4545e-08
f(P) c3 | -5.8081e-11
f(P) c4 | 4.6733e-14
f(P) c5 | -1.8059e-17
f(P) c6 | 2.7234e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1062
k0 calibration window | 2022-12-05 14:56 | 2022-12-08 00:05
Color pH time offset | 8 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 2
Ik outlier % | 0.12
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1062_DURA_2.dat | MSC1062_DURA.dat
median k0 | -1.4013 | -1.4005 | Volts
std k0 | 6.6983e-05 | 0.00013008 | Volts
Avg Vk | -0.9358 | -0.9357 | Volts
Avg std Vk | 47.9 | 69.8 | uVolts
median Ik: | 1.16 | 1.5 | nano amps
std Ik: | 2.25 | 2.36 | nano amps
median Ib: | 0.17 | 0.43 | nano amps
std Ib: | 4.2 | 4.08 | nano amps

## GDF030_Calibration_20221221
PH CALIBRATION FILE GENERATED ON: 12/21/2022 13:37
GDF 030
MSC 1061
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = NaN, APEX# 030
9, number of calibration coefficients
-1.28273, =k0 Pon; -1.28246 =k0 Poff; -1.27647 =k0 HCl
3.05274e-12,-5.38581e-08,-8.34422e-04, =k2(fP)*T
-3.79192e-06, =k3*P
1.06468e-08, =k4*P^3
-5.35981e-11, =k5*P^4
9.51281e-14, =k6*P^5
-7.54089e-17, =k7*P^6
2.79855e-20, =k8*P^7
-3.97189e-24, =k9*P^8
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
11/05-11/10/2022 |  |  | -0.019 |  |  |  |  | 3.0527e-12 | -5.3858e-08 | -0.00083442 | -1.2765 | 0.9997 |  | -3.9719e-24 | 2.7985e-20 | -7.5409e-17 | 9.5128e-14 | -5.3598e-11 | 1.0647e-08 | -3.7919e-06 | -1.2764 | 0.9777 | GDF |  | 30 | 12/05-12/07/2022 |  | -1.2825 | 9.6274e-05 | -1.2827 | 6.5121e-05
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 30
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF030_110322_Ch E_new DF board on Ch F.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-11-05 05:48 | 2022-11-10 07:44
Data subset: | ASCENDING
Vrs drift rate: | -0.01903 | mV/day
Vk drift rate: | -0.02384 | mV/day
Avg std Vk | 160.4049 | uVolts
Average k2: | -0.000871 | V/degC
Std dev k2: | 3.2109e-05 | V/degC
k2 f(P) c0 | -0.00083442
k2 f(P) c1 | -5.3858e-08
k2 f(P) c2 | 3.0527e-12
k0 HCL: | -1.2765 | Volts
Avg Ik: | -19.75 | nano amps
Std dev Ik: | 2.84 | nano amps
Avg Ib: | 7.14 | nano amps
Std dev Ib: | 3.04 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
22.9 | -0.00083386 | 3.6846e-06 | -1.2765 | 5.3725e-05 | 0.9997 | 18 | 1.9 | 19.8
39 | -0.00083346 | 2.9671e-06 | -1.2766 | 4.122e-05 | 0.9997 | 28 | 1.9 | 19.8
61.4 | -0.00083805 | 3.4578e-06 | -1.2766 | 4.7458e-05 | 0.9996 | 24 | 1.9 | 19.8
85.9 | -0.00083893 | 3.827e-06 | -1.2767 | 5.5034e-05 | 0.9995 | 28 | 1.9 | 19.8
125.3 | -0.00084164 | 2.7105e-06 | -1.2768 | 3.7603e-05 | 0.9995 | 49 | 1.9 | 19.8
174.8 | -0.00084297 | 2.4146e-06 | -1.277 | 3.3271e-05 | 0.9997 | 44 | 1.9 | 19.8
248.4 | -0.00085087 | 2.6779e-06 | -1.2772 | 3.7532e-05 | 0.9991 | 90 | 1.9 | 19.8
348.7 | -0.00085719 | 2.5583e-06 | -1.2776 | 3.5956e-05 | 0.9992 | 89 | 1.9 | 19.8
448.4 | -0.00085943 | 3.0621e-06 | -1.278 | 4.2179e-05 | 0.9989 | 87 | 1.9 | 19.8
548.6 | -0.00084224 | 4.4457e-06 | -1.2789 | 6.2309e-05 | 0.9976 | 88 | 1.9 | 19.8
649.5 | -0.00088333 | 2.2549e-06 | -1.2789 | 3.5323e-05 | 0.9996 | 69 | 1.9 | 19.8
751.8 | -0.00087966 | 2.0794e-06 | -1.2792 | 2.9405e-05 | 0.9996 | 82 | 2 | 19.8
900.3 | -0.00088343 | 1.1931e-06 | -1.2793 | 1.6725e-05 | 0.9997 | 170 | 2 | 19.8
1099.5 | -0.00088848 | 1.102e-06 | -1.2794 | 1.5263e-05 | 0.9997 | 166 | 2 | 19.8
1299.8 | -0.00089578 | 1.2153e-06 | -1.2793 | 1.7015e-05 | 0.9997 | 168 | 2 | 19.8
1500.1 | -0.00090368 | 1.2554e-06 | -1.2793 | 1.7441e-05 | 0.9997 | 161 | 2 | 19.8
1700.5 | -0.00091312 | 1.321e-06 | -1.2791 | 1.877e-05 | 0.9997 | 150 | 2.1 | 19.8
1902.1 | -0.0009268 | 1.5438e-06 | -1.2788 | 2.1583e-05 | 0.9996 | 162 | 2.1 | 19.8
2034.9 | -0.000936 | 2.628e-06 | -1.2785 | 3.6169e-05 | 0.9996 | 52 | 2.1 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.8149e-08
f(P) R*R: | 0.9777
f(P) c0 | -1.2764
f(P) c1 | -3.7919e-06
f(P) c2 | 1.0647e-08
f(P) c3 | -5.3598e-11
f(P) c4 | 9.5128e-14
f(P) c5 | -7.5409e-17
f(P) c6 | 2.7986e-20
f(P) c7 | -3.9719e-24
** KO CALIBRATION SUMMARY: **
MSC # : | 1061
k0 calibration window | 2022-12-05 15:43 | 2022-12-07 23:39
Color pH time offset | 8 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 1
Ik outlier % | 0.06
Ib outlier count | 1
Ib outlier % | 0.06
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1061_DURA_2.dat | MSC1061_DURA.dat
median k0 | -1.2827 | -1.2824 | Volts
std k0 | 6.5121e-05 | 9.6274e-05 | Volts
Avg Vk | -0.7575 | -0.7575 | Volts
Avg std Vk | 49.2 | 48.5 | uVolts
median Ik: | -9.37 | -9.51 | nano amps
std Ik: | 2.7 | 2.74 | nano amps
median Ib: | -0.27 | -0.05 | nano amps
std Ib: | 5.98 | 5.79 | nano amps

## GDF25_Calibration_20230525
PH CALIBRATION FILE GENERATED ON: 05/25/2023 10:55
GDF 25
MSC 1051
APEX ID 21910
** START CONFIG FILE COPY BLOCK **
Durafet SN = 25, APEX# 21910
9, number of calibration coefficients
-1.41788, =k0 Pon; -1.41782 =k0 Poff; -1.41689 =k0 HCl
3.36172e-17,-1.42063e-13,1.66803e-10,-6.27447e-08,-9.68688e-04, =k2(fP)*T
-1.48176e-06, =k3*P^1
-1.32566e-08, =k4*P^2
3.76481e-11, =k5*P^3
-6.14464e-14, =k6*P^4
4.99300e-17, =k7*P^5
-1.89209e-20, =k8*P^6
2.69854e-24, =k9*P^7
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
09/03-09/08/2022 |  |  | 0.136 |  |  | 3.3617e-17 | -1.4206e-13 | 1.668e-10 | -6.2745e-08 | -0.00096869 | -1.4169 | 0.9998 |  | 2.6985e-24 | -1.8921e-20 | 4.993e-17 | -6.1446e-14 | 3.7648e-11 | -1.3257e-08 | -1.4818e-06 | -1.4168 | 0.9975 | GDF |  | 25 | 10/28-11/01/2022 |  | -1.4178 | 0.0001741 | -1.4179 | 0.00015921
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 25
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF25_090122_Ch E.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-09-03 04:17 | 2022-09-08 07:16
Data subset: | ASCENDING
Vrs drift rate: | 0.13574 | mV/day
Vk drift rate: | -1.1321 | mV/day
Avg std Vk | 103.9831 | uVolts
Average k2: | -0.0009824 | V/degC
Std dev k2: | 1.747e-05 | V/degC
k2 f(P) c0 | -0.00096869
k2 f(P) c1 | -6.2745e-08
k2 f(P) c2 | 1.668e-10
k2 f(P) c3 | -1.4206e-13
k2 f(P) c4 | 3.3617e-17
k0 HCL: | -1.4169 | Volts
Avg Ik: | 7.04 | nano amps
Std dev Ik: | 1.43 | nano amps
Avg Ib: | 1.23 | nano amps
Std dev Ib: | 2.33 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.9 | -0.00096488 | 3.179e-06 | -1.4169 | 4.5543e-05 | 0.9998 | 21 | 1.9 | 19.8
39 | -0.00097218 | 2.333e-06 | -1.4168 | 3.3384e-05 | 0.9998 | 33 | 1.9 | 19.8
61.4 | -0.00097429 | 3.0549e-06 | -1.4169 | 4.2418e-05 | 0.9997 | 31 | 1.9 | 19.8
86.3 | -0.00097302 | 2.8771e-06 | -1.417 | 3.9923e-05 | 0.9998 | 28 | 1.9 | 19.8
123.7 | -0.00097546 | 2.3048e-06 | -1.4171 | 3.2105e-05 | 0.9997 | 51 | 1.9 | 19.8
174.2 | -0.00097728 | 2.3119e-06 | -1.4173 | 3.2525e-05 | 0.9997 | 49 | 1.9 | 19.8
249.7 | -0.00097694 | 2.5021e-06 | -1.4176 | 3.5116e-05 | 0.9994 | 90 | 1.9 | 19.8
350 | -0.00097528 | 3.1219e-06 | -1.418 | 4.6014e-05 | 0.9992 | 81 | 1.9 | 19.8
449.9 | -0.00097215 | 3.0252e-06 | -1.4185 | 4.2816e-05 | 0.9992 | 86 | 1.9 | 19.8
550.4 | -0.0009715 | 3.5385e-06 | -1.419 | 5.1749e-05 | 0.999 | 80 | 1.9 | 19.8
648.5 | -0.00097492 | 3.5275e-06 | -1.4195 | 5.1006e-05 | 0.999 | 80 | 1.9 | 19.7
750.5 | -0.0009695 | 3.3126e-06 | -1.4201 | 4.717e-05 | 0.999 | 84 | 1.9 | 19.7
898.4 | -0.00097087 | 4.0224e-06 | -1.421 | 5.5643e-05 | 0.9971 | 170 | 1.9 | 19.7
1096.8 | -0.00097812 | 4.2242e-06 | -1.4221 | 6.0178e-05 | 0.997 | 164 | 1.9 | 19.7
1299.4 | -0.00098456 | 1.8306e-06 | -1.423 | 2.5507e-05 | 0.9994 | 163 | 2 | 19.7
1499.5 | -0.00099592 | 1.2959e-06 | -1.4231 | 1.8061e-05 | 0.9997 | 167 | 2 | 19.7
1700.8 | -0.0010104 | 1.195e-06 | -1.4229 | 1.6734e-05 | 0.9998 | 164 | 2 | 19.8
1901.5 | -0.0010217 | 1.3252e-06 | -1.4226 | 1.8449e-05 | 0.9997 | 164 | 2 | 19.8
2034.6 | -0.0010267 | 2.1576e-06 | -1.4225 | 2.9349e-05 | 0.9998 | 51 | 2 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.2376e-08
f(P) R*R: | 0.9975
f(P) c0 | -1.4168
f(P) c1 | -1.4818e-06
f(P) c2 | -1.3257e-08
f(P) c3 | 3.7648e-11
f(P) c4 | -6.1446e-14
f(P) c5 | 4.993e-17
f(P) c6 | -1.8921e-20
f(P) c7 | 2.6985e-24
** KO CALIBRATION SUMMARY: **
MSC # : | 1051
k0 calibration window | 2022-10-28 15:52 | 2022-11-01 19:56
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 10
Ik outlier % | 0.33
Ib outlier count | 5
Ib outlier % | 0.17
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1051_DURA_2.dat | MSC1051_DURA.dat
median k0 | -1.4179 | -1.4178 | Volts
std k0 | 0.00015921 | 0.0001741 | Volts
Avg Vk | -0.8994 | -0.8994 | Volts
Avg std Vk | 53.7 | 53.3 | uVolts
median Ik: | -5.18 | -5.2 | nano amps
std Ik: | 3.21 | 3.18 | nano amps
median Ib: | -0.36 | 0.03 | nano amps
std Ib: | 5.6 | 5.77 | nano amps

## GDF2001_Calibration_20230627
PH CALIBRATION FILE GENERATED ON: 06/27/2023 11:18
GDF 2001
MSC 1077
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = 2001, APEX# NaN
8, number of calibration coefficients
-1.31856, =k0 Pon; -1.31769 =k0 Poff; -1.31430 =k0 HCl
3.57824e-17,-1.75680e-13,2.78001e-10,-2.07878e-07,-8.71520e-04, =k2(fP)*T
-1.03090e-05, =k3*P^1
-4.36340e-09, =k4*P^2
1.47533e-11, =k5*P^3
-1.28044e-14, =k6*P^4
4.77830e-18, =k7*P^5
-6.30310e-22, =k8*P^6
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
12/29-01/03/2023 |  |  | 0.023 |  |  | 3.5782e-17 | -1.7568e-13 | 2.78e-10 | -2.0788e-07 | -0.00087152 | -1.3143 | 0.9992 |  |  | -6.3031e-22 | 4.7783e-18 | -1.2804e-14 | 1.4753e-11 | -4.3634e-09 | -1.0309e-05 | -1.3142 | 0.9991 | GDF |  | 2001 | 05/01-05/03/2023 |  | -1.3177 | 8.5987e-05 | -1.3186 | 0.00025765
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 2001
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | AE_GDF2001_122822_Ch F.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-12-29 17:37 | 2023-01-03 20:31
Data subset: | ASCENDING
Vrs drift rate: | 0.02283 | mV/day
Vk drift rate: | 0.69059 | mV/day
Avg std Vk | 195.2671 | uVolts
Average k2: | -0.00092989 | V/degC
Std dev k2: | 3.9642e-05 | V/degC
k2 f(P) c0 | -0.00087152
k2 f(P) c1 | -2.0788e-07
k2 f(P) c2 | 2.78e-10
k2 f(P) c3 | -1.7568e-13
k2 f(P) c4 | 3.5782e-17
k0 HCL: | -1.3143 | Volts
Avg Ik: | -19.3 | nano amps
Std dev Ik: | 3.54 | nano amps
Avg Ib: | 4.99 | nano amps
Std dev Ib: | 3.3 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.3 | -0.00088116 | 6.3938e-06 | -1.3143 | 9.0954e-05 | 0.9992 | 17 | 2 | 19.8
38.8 | -0.00088182 | 4.222e-06 | -1.3145 | 6.0331e-05 | 0.9995 | 26 | 2 | 19.8
62.4 | -0.00088151 | 4.9086e-06 | -1.3148 | 6.9434e-05 | 0.9992 | 29 | 2 | 19.8
88.1 | -0.0008851 | 4.8709e-06 | -1.3151 | 6.758e-05 | 0.9994 | 22 | 2 | 19.8
124.1 | -0.00088939 | 5.0854e-06 | -1.3155 | 7.1331e-05 | 0.9985 | 47 | 2 | 19.8
174 | -0.00089999 | 4.7282e-06 | -1.3161 | 6.7659e-05 | 0.9987 | 48 | 2 | 19.8
249.1 | -0.00090938 | 5.4824e-06 | -1.3168 | 7.6473e-05 | 0.9968 | 89 | 2 | 19.8
350.3 | -0.00091236 | 5.2996e-06 | -1.3179 | 7.3913e-05 | 0.9972 | 86 | 2 | 19.8
451.2 | -0.00091872 | 4.8687e-06 | -1.3188 | 6.7916e-05 | 0.9976 | 89 | 2 | 19.8
547.1 | -0.00093652 | 5.5366e-06 | -1.3195 | 8.3355e-05 | 0.9976 | 71 | 2 | 19.8
647 | -0.00094033 | 5.2615e-06 | -1.3203 | 8.0217e-05 | 0.9978 | 73 | 2.1 | 19.8
751.3 | -0.00092955 | 4.7137e-06 | -1.3212 | 6.9833e-05 | 0.9982 | 72 | 2 | 19.8
897.8 | -0.00093618 | 4.736e-06 | -1.3222 | 6.6335e-05 | 0.9958 | 168 | 2 | 19.8
1098.2 | -0.00094276 | 4.1162e-06 | -1.3234 | 5.7384e-05 | 0.997 | 159 | 2 | 19.8
1298.7 | -0.00095339 | 4.1017e-06 | -1.3244 | 5.7275e-05 | 0.9971 | 158 | 2 | 19.8
1498.4 | -0.00097013 | 4.2368e-06 | -1.3254 | 5.9167e-05 | 0.997 | 161 | 2 | 19.8
1700.4 | -0.00098829 | 2.8251e-06 | -1.3261 | 3.9578e-05 | 0.9987 | 161 | 2.1 | 19.8
1903.2 | -0.0010023 | 1.6472e-06 | -1.3265 | 2.3081e-05 | 0.9996 | 159 | 2.1 | 19.8
2035.6 | -0.001009 | 2.1378e-06 | -1.3265 | 3.0162e-05 | 0.9998 | 49 | 2.1 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 1.3395e-08
f(P) R*R: | 0.9991
f(P) c0 | -1.3142
f(P) c1 | -1.0309e-05
f(P) c2 | -4.3634e-09
f(P) c3 | 1.4753e-11
f(P) c4 | -1.2804e-14
f(P) c5 | 4.7783e-18
f(P) c6 | -6.3031e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1077
k0 calibration window | 2023-05-01 14:26 | 2023-05-03 18:38
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 1
Ik outlier % | 0.06
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1077_DURA_2 (3).dat | MSC1077_DURA (3).dat
median k0 | -1.3185 | -1.3177 | Volts
std k0 | 0.00025765 | 8.5987e-05 | Volts
Avg Vk | -0.7464 | -0.7463 | Volts
Avg std Vk | 45.9 | 46.1 | uVolts
median Ik: | -6.3 | -6.01 | nano amps
std Ik: | 3.39 | 3.3 | nano amps
median Ib: | -0.2 | 0.53 | nano amps
std Ib: | 6.49 | 6.39 | nano amps

## GDF2009_Calibration_20230627
PH CALIBRATION FILE GENERATED ON: 06/27/2023 11:24
GDF 2009
MSC 1076
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = 2009, APEX# NaN
8, number of calibration coefficients
-1.32653, =k0 Pon; -1.32579 =k0 Poff; -1.32168 =k0 HCl
-3.54488e-14,1.18899e-10,-1.31134e-07,-7.69328e-04, =k2(fP)*T
-1.83443e-05, =k3*P^1
3.43344e-08, =k4*P^2
-4.78310e-11, =k5*P^3
3.89486e-14, =k6*P^4
-1.61422e-17, =k7*P^5
2.60672e-21, =k8*P^6
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
12/11-12/16/2022 |  |  | -0.116 |  |  |  | -3.5449e-14 | 1.189e-10 | -1.3113e-07 | -0.00076933 | -1.3217 | 0.9997 |  |  | 2.6067e-21 | -1.6142e-17 | 3.8949e-14 | -4.7831e-11 | 3.4334e-08 | -1.8344e-05 | -1.3215 | 0.9953 | GDF |  | 2009 | 05/01-05/03/2023 |  | -1.3258 | 0.00017116 | -1.3265 | 0.00034575
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 2009
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | AE_GDF2009_120922_Ch E.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2022-12-11 00:58 | 2022-12-16 02:54
Data subset: | ASCENDING
Vrs drift rate: | -0.11602 | mV/day
Vk drift rate: | 2.3177 | mV/day
Avg std Vk | 103.5954 | uVolts
Average k2: | -0.00080466 | V/degC
Std dev k2: | 2.1193e-05 | V/degC
k2 f(P) c0 | -0.00076933
k2 f(P) c1 | -1.3113e-07
k2 f(P) c2 | 1.189e-10
k2 f(P) c3 | -3.5449e-14
k0 HCL: | -1.3217 | Volts
Avg Ik: | 7.6 | nano amps
Std dev Ik: | 1.48 | nano amps
Avg Ib: | 1.01 | nano amps
Std dev Ib: | 2.38 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
23.7 | -0.00078849 | 3.2675e-06 | -1.3217 | 4.5138e-05 | 0.9997 | 18 | 1.9 | 19.8
39.3 | -0.00077999 | 1.0381e-05 | -1.3221 | 0.00015875 | 0.9972 | 18 | 1.9 | 19.8
63.1 | -0.00078109 | 1.0116e-05 | -1.3224 | 0.00014617 | 0.9965 | 23 | 1.9 | 19.8
86.4 | -0.00076812 | 1.6368e-05 | -1.323 | 0.00024388 | 0.9941 | 15 | 1.9 | 19.8
122.5 | -0.00076592 | 9.8061e-06 | -1.3236 | 0.00013907 | 0.994 | 39 | 1.9 | 19.8
175 | -0.00077887 | 7.4787e-06 | -1.3241 | 0.00010752 | 0.9968 | 37 | 1.9 | 19.8
249.6 | -0.00079871 | 5.1994e-06 | -1.3245 | 7.5243e-05 | 0.9973 | 66 | 1.9 | 19.8
350.8 | -0.00080876 | 4.0396e-06 | -1.3251 | 5.764e-05 | 0.9984 | 65 | 1.9 | 19.8
448.6 | -0.00080991 | 3.905e-06 | -1.3257 | 5.5541e-05 | 0.9986 | 63 | 1.8 | 19.8
549.5 | -0.00081386 | 5.0604e-06 | -1.3263 | 7.8559e-05 | 0.9979 | 56 | 1.8 | 19.8
653 | -0.00081191 | 4.0075e-06 | -1.3269 | 6.0346e-05 | 0.9987 | 57 | 1.9 | 19.8
749.6 | -0.00081634 | 3.3686e-06 | -1.3272 | 4.8436e-05 | 0.9989 | 64 | 1.9 | 19.8
903.1 | -0.00081858 | 2.5529e-06 | -1.3277 | 3.618e-05 | 0.9988 | 126 | 1.9 | 19.8
1098.3 | -0.00082088 | 2.0807e-06 | -1.3281 | 3.1407e-05 | 0.9993 | 110 | 1.9 | 19.8
1302.4 | -0.00081134 | 2.2349e-06 | -1.3286 | 3.3211e-05 | 0.9991 | 118 | 2 | 19.8
1500 | -0.00081587 | 2.484e-06 | -1.329 | 3.5992e-05 | 0.9989 | 121 | 2 | 19.8
1699.3 | -0.00082409 | 2.4774e-06 | -1.3295 | 3.534e-05 | 0.9989 | 122 | 2 | 19.8
1900.8 | -0.00083354 | 2.3514e-06 | -1.3299 | 3.4142e-05 | 0.9991 | 121 | 2 | 19.8
2033.7 | -0.00084227 | 2.4134e-06 | -1.3301 | 3.2949e-05 | 0.9997 | 38 | 2 | 19.8
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 2.2798e-08
f(P) R*R: | 0.9953
f(P) c0 | -1.3215
f(P) c1 | -1.8344e-05
f(P) c2 | 3.4334e-08
f(P) c3 | -4.7831e-11
f(P) c4 | 3.8949e-14
f(P) c5 | -1.6142e-17
f(P) c6 | 2.6067e-21
** KO CALIBRATION SUMMARY: **
MSC # : | 1076
k0 calibration window | 2023-05-01 14:21 | 2023-05-03 18:41
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 4
Ik outlier % | 0.25
Ib outlier count | 0
Ib outlier % | 0
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1076_DURA_2 (3).dat | MSC1076_DURA (3).dat
median k0 | -1.3265 | -1.3257 | Volts
std k0 | 0.00034575 | 0.00017116 | Volts
Avg Vk | -0.7442 | -0.7441 | Volts
Avg std Vk | 45.2 | 43.7 | uVolts
median Ik: | -3.71 | -3.74 | nano amps
std Ik: | 2.48 | 2.36 | nano amps
median Ib: | -0.45 | -0.01 | nano amps
std Ib: | 5.01 | 5.06 | nano amps

## GDF2000_Calibration_20230627
PH CALIBRATION FILE GENERATED ON: 06/27/2023 12:24
GDF 2000
MSC 1091
APEX ID NaN
** START CONFIG FILE COPY BLOCK **
Durafet SN = 2000, APEX# NaN
8, number of calibration coefficients
-1.28421, =k0 Pon; -1.28406 =k0 Poff; -1.28046 =k0 HCl
2.54966e-20,-1.51701e-16,3.19871e-13,-2.81697e-10,7.84387e-08,-9.01112e-04, =k2(fP)*T
-7.01924e-06, =k3*P^1
5.35221e-09, =k4*P^2
-3.90690e-12, =k5*P^3
-9.66978e-16, =k6*P^4
1.89158e-18, =k7*P^5
-4.64527e-22, =k8*P^6
** END CONFIG FILE COPY BLOCK **
** START pHlog.xlsx COPY BLOCK **
Ptchamber | Ch | PT runs | VrsDrift | k2 f(P) C6 | k2 f(P) C5 | k2 f(P) C4 | k2 f(P) C3 | k2 f(P) C2 | k2 f(P) C1 | k2 f(P) C0 | k0(HCl) | R^2 | f(P) k8 | f(P) k7 | f(P) k6 | f(P) k5 | f(P) k4 | f(P) k3 | f(P) k2 | f(P) k1 | f(P) k0 | R^2 | pH Design | Die | DF# | k0CalDates | k0Soln | k0_pOFF | Stdk0_pOFF | k0_pON | Stdk0_pON
04/13-04/18/2023 |  |  | 0.108 |  | 2.5497e-20 | -1.517e-16 | 3.1987e-13 | -2.817e-10 | 7.8439e-08 | -0.00090111 | -1.2805 | 0.9999 |  |  | -4.6453e-22 | 1.8916e-18 | -9.6698e-16 | -3.9069e-12 | 5.3522e-09 | -7.0192e-06 | -1.2805 | 0.9994 | GDF |  | 2000 | 05/08-05/10/2023 |  | -1.2841 | 0.0001699 | -1.2842 | 0.00013293
** END pHlog.xlsx COPY BLOCK **
** SUMMARY K2 DATA: **
Durafet SN: | 2000
pH stem mfg: | MBARI | GDF
PT Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K2_PTChamber_data\
PT Source file : | GDF2000_041123 Ti cvrplt_Ch B.csv
PT test channel :
HCl concentration: | 0.01 | Molar
k2 Calibration window: | 2023-04-13 01:20 | 2023-04-18 03:57
Data subset: | ASCENDING
Vrs drift rate: | 0.10787 | mV/day
Vk drift rate: | -1.1143 | mV/day
Avg std Vk | 47.4369 | uVolts
Average k2: | -0.00090438 | V/degC
Std dev k2: | 8.6175e-06 | V/degC
k2 f(P) c0 | -0.00090111
k2 f(P) c1 | 7.8439e-08
k2 f(P) c2 | -2.817e-10
k2 f(P) c3 | 3.1987e-13
k2 f(P) c4 | -1.517e-16
k2 f(P) c5 | 2.5497e-20
k0 HCL: | -1.2805 | Volts
Avg Ik: | -25.75 | nano amps
Std dev Ik: | 8.28 | nano amps
Avg Ib: | -1.5 | nano amps
Std dev Ib: | 9.25 | nano amps
** PRESSURE BINNED K2 DATA: **
Avg Press | k2 | std k2 | k0 | std k0 | R*R | N | Tmin | Tmax
6.9 | -0.00089876 | 3.1447e-06 | -1.2805 | 4.4343e-05 | 0.9999 | 10 | 2.1 | 19.9
14.8 | -0.000901 | 9.3163e-07 | -1.2805 | 1.3425e-05 | 0.9999 | 121 | 2.1 | 19.9
24.7 | -0.0009001 | 1.526e-06 | -1.2806 | 2.246e-05 | 0.9998 | 60 | 2.1 | 19.9
39 | -0.00089884 | 1.5402e-06 | -1.2807 | 2.1992e-05 | 0.9998 | 75 | 2.1 | 19.9
61.8 | -0.0008968 | 1.5636e-06 | -1.2809 | 2.2288e-05 | 0.9998 | 65 | 2.1 | 19.9
87.5 | -0.0008965 | 1.6202e-06 | -1.281 | 2.2406e-05 | 0.9998 | 53 | 2.1 | 19.9
125.1 | -0.00089603 | 1.6009e-06 | -1.2813 | 2.2563e-05 | 0.9997 | 90 | 2.1 | 19.9
175.4 | -0.00089436 | 1.5734e-06 | -1.2816 | 2.2041e-05 | 0.9998 | 81 | 2.1 | 19.9
249.5 | -0.00089285 | 1.7825e-06 | -1.282 | 2.5001e-05 | 0.9994 | 152 | 2.1 | 19.9
350.5 | -0.0008962 | 1.8247e-06 | -1.2824 | 2.5673e-05 | 0.9994 | 147 | 2.1 | 19.9
449.1 | -0.00090101 | 1.8748e-06 | -1.2828 | 2.6372e-05 | 0.9994 | 142 | 2.1 | 19.9
550.2 | -0.00090235 | 2.2639e-06 | -1.2833 | 3.2893e-05 | 0.9992 | 132 | 2.1 | 19.9
647.9 | -0.00090602 | 2.1095e-06 | -1.2838 | 2.9678e-05 | 0.9993 | 139 | 2.1 | 19.9
749 | -0.00090722 | 1.8982e-06 | -1.2843 | 2.6658e-05 | 0.9994 | 142 | 2.1 | 19.9
899.1 | -0.00090988 | 2.4918e-06 | -1.285 | 3.4805e-05 | 0.9979 | 279 | 2.1 | 19.9
1098.7 | -0.00091021 | 2.9405e-06 | -1.2861 | 4.1252e-05 | 0.9972 | 273 | 2.1 | 19.9
1299.5 | -0.00091172 | 2.6162e-06 | -1.2871 | 3.6541e-05 | 0.9977 | 276 | 2.1 | 19.9
1500.5 | -0.00091221 | 2.006e-06 | -1.288 | 2.822e-05 | 0.9987 | 269 | 2.2 | 19.9
1700.8 | -0.0009153 | 1.4903e-06 | -1.2885 | 2.0813e-05 | 0.9993 | 269 | 2.2 | 19.9
1900.4 | -0.00092041 | 1.0833e-06 | -1.2889 | 1.5214e-05 | 0.9996 | 262 | 2.2 | 19.9
2032.5 | -0.00092428 | 1.3047e-06 | -1.2891 | 1.7973e-05 | 0.9998 | 86 | 2.2 | 19.9
** PH POLYNOMIAL PRESSURE COEFFICIENTS: **
f(P) NSSR: | 5.287e-09
f(P) R*R: | 0.9994
f(P) c0 | -1.2805
f(P) c1 | -7.0192e-06
f(P) c2 | 5.3522e-09
f(P) c3 | -3.9069e-12
f(P) c4 | -9.6698e-16
f(P) c5 | 1.8916e-18
f(P) c6 | -4.6453e-22
** KO CALIBRATION SUMMARY: **
MSC # : | 1091
k0 calibration window | 2023-05-08 15:17 | 2023-05-10 19:01
Color pH time offset | 7 | hours
k0 poly fit | 1
outlier_threshold | 4
Ik outlier count | 34
Ik outlier % | 2.19
Ib outlier count | 18
Ib outlier % | 1.16
k0 Source directory : | \\atlas\chem\DuraFET\CALIBRATION\K0_seawater_data\
Parameter | Pump ON | Pump OFF | Units
k0 file | MSC1091_DURA_2.dat | MSC1091_DURA.dat
median k0 | -1.2842 | -1.2841 | Volts
std k0 | 0.00013293 | 0.0001699 | Volts
Avg Vk | -0.7994 | -0.7994 | Volts
Avg std Vk | 49.4 | 49.3 | uVolts
median Ik: | -6.91 | -7.04 | nano amps
std Ik: | 9.11 | 9.18 | nano amps
median Ib: | -1.17 | -0.81 | nano amps
std Ib: | 9.47 | 9.76 | nano amps