## TestLog
pTC Experiment Log
Date | SN | NameOfFolder | DescriptionOfExperiment | SummaryResultsComments | Chamber | GDF | Deployed | Fluke | Prunella | Machined Corners | Broached Corners | Test | AE | NG AE
2020-10-26 | GDF004 | GDF4_102620 | Temperature ramps ever 4C between 30 and 2. First test with butterfly board. f(P) done in Ti vesel | Characterized thermal response and steady k2. Ready for SW eq | A | [6 empty] | 1
2021-03-22 | GDF005 | GDF5_032221 | Temperature ramps ever 4C between 30 and 2. Test for spray glider delpoyment | Characterized thermal repsonse. Ready for SW eq | A | [6 empty] | 1
2021-04-22 | GDF006 | GDF6_initialCal | 4 hour pres ramps to 2000 dBar. 20, 15, and 12 C. | Characterized thermal response. fP reproducible, but had clear temperature response. fP magnitude big, clear Pres hysteresis. | A | [6 empty] | 1
2021-04-26 | GDF006 | GDF6_secondCal | 4 hour pres ramps to 2000 dBar. 20, 15, and 12 C. then 1hr ramps at 8 C | Clear hysteresis in presure response and walled Ib. Significant pressure dependence on k2 past 1500 dBar | A
2021-05-02 | GDF007 | GDF7_initialCal | 1hr pres ramps to 2000 dBar, 20, 15, 10, 8, 6, 4, 2, then back up to 20 with same steps | Less hysteresis compared to GDF6, but fP magnitude still large. Pretty large pressure dependence on k2 | A | [6 empty] | 1
2021-05-06 | GDF007 | GDF7_4hr_Ptcycle | 4 hour pres ramps to 2000 dBar at 10C | Chamber leaked in the middle once; lost a cycle. Goal was to see if fP was different between 4 hour and 1 hour cycles | A |  |  |  | 1
2021-05-18 | GDF007 | GDF7_k2_20_1000dBar | T ramp from 2 to 30 C at 20 and 1000 dBar. | Clear pressure dependence on k2. | A |  |  |  | 1
2021-05-16 | GDF007 | GDF7_variable_Prate_20_10_2 | Various pres/flow rates to 2000 dBar, 20, 10, and 2 C. | Had a lot of memory problems, Yui on cruise so couldn't deal with it effectively. Follow up to previous experiment, so we can get a more detailed view on how flow rate affects fP, and across multiple temperatures | A
2021-05-20 | GDF007 | GDF7_variable_Prate_20_10_2_take2 | Various flow rates to 2000 dBar, 10 and 2 C. Finish of the previous experiment | Large f(P), smaller hysteresis but definitive pressure dependence on k2 values. Needs to be pres corrected for repeatable f(P) | A |  |  |  | 1
2021-05-25 | DSD314 | DSD314 | 2 P cycles to 2000 dBar each at 20, 10, and 2 C | Very similar fP as that characterized in other P chamber. No pressure dependence to k2 | A
2021-05-31 | GDF007 | GDF7_2_20_10_2 | 20, 10, 2, 10, 20 C, with 2x P cycles to 2000 dBar each | New cover plate with shorter shoulder, new ISFET chip. Saw large fP still (~20mV) | A
2021-06-08 | GDF008 | GDF8_fP_cal | 20, 10, 2, 10, 20 C, with 2x P cycles to 2000 dBar each | Smallest fP so far (~6 mV). Still see pres influence on k2. | A
2021-06-14 | GDF008 | GDF8_fP_cal_new gasket | 20 C 2x P cycles and 2 C 4x P cycles to 2000 dBar | chip coverplate removed and inspected, new gasket installed to confirm fP response | A
2021-06-16 | GDF008 | GDF8_fP_cal_new gasket slow pres cycle | 20 C 2x P cycles and 2 C 4x P cycles to 2000 dBar | pump rate decreased to 0.02ml/min to test hysterisis at slow pump rates | A
2021-06-18 | GDF008 | GDF8_90A gasket | 2000 dBar pressure cycles at 20 and 2 deg | testing harder durometer gasket. 90A polyurethane, Failed | A
2021-08-04 | GDF007 | GDF7_20_2C_Pcycle | 2000 dBar pressure cycles at 20 and 2 deg | Spun up chamber again. PT cycle to verify everything is working | A
2021-08-11 | GDF007 | GDF7_60uA_Ids | 2000 dBar pressure cycles at 20 and 2 deg, testing impacts of Ids on k2 and fP | Large reduction in fP was observed (~50%); k2 was near 0. | A
2021-08-14 | SBE11046 | SBE11046_60uA_PTstress | Accelerated P-T test. 20C at 30PSI, 2C at 2000 PSI. 1 cycle per 2 hours with 60uA MSC board | No failure after 8 days of cycling | A
2021-08-24 | GDF007 | GDF7_oring | 2 P cycles to 2000 dBar each at 20, 10, and 2 C, back to 188 uA Ids. new o-ring cover plate | Hysteresis maintained but smooth and smaller f(P) than prior iterations. | A
2021-08-31 | GDF007 | GDF7_retest_gasket | 2000 dBar at 20, 10 and 2 deg C. Viton gasket, tall cover plate | Large f(P), small hysteresis. Repeatable at variable k2 values indicating large presure dependence on k2 | A
2021-09-07 | GDF011 | GDF11_machinedcorners | 5 cycles to 2000 dBar. Machined corners, Viton gasket, tall cover plate. ALL GDF S/N ABOVE 10 ARE MACHINED CORNERS | Desirable f(P) with no noticeable hysteresis. Dented Gasket after 5 cycles to 2000 dBar | A
2021-09-08 | GDF011 | GDF11_machinedcorners_2 | 2000 dBar at 20, 10 and 2 deg C. Machined corners, NEW Viton gasket, tall cover plate | Desirable f(P) with slight hysteresis. | A
2021-09-14 | GDF012 | GDF12_machinedcorners | 2001 dBar at 20, 10 and 2 deg C. Machined corners, Viton gasket, tall cover plate | Desirable f(P) after k2 pres correction with slight hysteresis. | A
2021-09-17 | GDF012 | GDF12_stressTest_2000dBar | 2000 dBar test with over 50 cycles and minimal temperature change. | Extremely large hysteresis and used to preliminarily test the durability of machined corners and not as a calibration | A
2021-10-05 | GDF012 | GDF12_2600_dBar | 2600 dBar full calibration at temps 2:2:20 with break in period at beginning | Pump malfunctioned during break in period but showed large amount of hysteresis. Not usable data | A
2021-10-11 | GDF012 | GDF12_cal_every2C | 2000 dBar from 20:-2:2, then 2:2:20. 3x ramps each. To determine cal routines for deployment | Desirable but large f(P) after k2 pres correction with slight hysteresis. | A
2021-10-19 | GDF011 | GDF11_fullcal | 2000 dBar from 20:-2:2, then 2:2:20. 3x ramps each. To determine cal routines for deployment | Desirable but large f(P) after k2 pres correction with slight hysteresis. | A
2021-10-26 | GDF009 | GDF9_broached_1 | Testing the stabilitiy and repeatability of broached cornered GDFs. 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning | Large and undesirable f(P) with slight hysteresis | A
2021-11-10 | GDF009 | GDF9_broached_test | Testing the stabilitiy and repeatability of broached cornered GDFs. 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning | Run cut short due to arrival and testing of machined cornered stems and interest in eliminating noise issues. High variability of Ib | A
2021-11-24 | GDF013 | GDF_1124 | 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning and end (both to 30C) | Connection failed, useless data | A
2021-11-30 | GDF013 | GDF_1130 | 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning and end (both to 30C) | Desirable but large f(P) after k2 pres correction with slight hysteresis. Variable Ib | A
2021-12-08 | GDF014 | GDF14_1208 | 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning and end (both to 30C) | Desirable f(P), 2nd order k2 pres correction with slight hysteresis. Variable Ib | A
2021-12-16 | GDF015 | GDF15_1216 | 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning and end (both to 30C) | Desirable f(P) after k2 pres correction with slight hysteresis. Variable Ib | A
2021-12-29 | GDF015 | GDF15_1229_70Ids | 2000 dBar full calibration at temps 2:2:20 with 2600 break in period at beginning. Switching out resistor on Durafet adap. Board to get 70 Ids | Small f(P) but just as varaible as f(P) with larger magnitude, no difference in characterization. Slight hysteresis, good Ib values | A
2022-01-15 | SBET112 | SBET112_fullcal | 2600 dBar full calibration at temps 2:2:20 with break in period at beginning and end (both to 30C). Test Pcomp. FIRST RUN WITH DUAL pTC (BIVALVE) | Desirable f(P) that matches with cal sheet with minimal hysteresis. Pcomp works. K2 is off | B
2022-01-15 | GDF015 | GDF15_ChanA_2022_01_15 | 2600 dBar full calibration at temps 2:2:20 with break in period at beginning and end (both to 30C). FIRST RUN WITH DUAL pTC (BIVALVE) | Desirable f(P) after k2 pres correction with slight hysteresis. Noise in Ik,Ib associated with channel A | A
2022-01-24 | SBET112 | SBET112_ChanA_2022_01_24 | 2600 dBar full calibration at temps 2:10:30 with break in period at beginning to verify the intermittent temps not as crucial | Desirable f(P) that matches with cal sheet with minimal hysteresis. Noise in Ik,Ib associated with channel A. Verified larger temp steps are ok. K2 is off and noise is suspect | A
2022-01-24 | GDF015 | GDF15_ChanB_2022_01_24 | 2600 dBar full calibration at temps 2:10:30 with break in period at beginning to verify the intermittent temps not as crucial | Desirable f(P) after k2 pres correction with slightly larger hysteresis than before. Verified larger temp steps are ok | B
2022-01-31 | SBE11307 | SBE11307_stress_test | Cycles to 2600 dBar at 20C and 30C to stress test the high k2 sensor | Verified it is operating well and can be deployed despite the "out of family" k2 values | A
2022-01-31 | SBE11311 | SBE11311_stress_test | Cycles to 2600 dBar at 20C and 30C to stress test the high k2 sensors | Verified it is operating well and can be deployed despite the "out of family" k2 values | B
2022-02-01 | SBET110 | SBET110 | Cycles to 2600 dBar at temp steps of 2:10:30 to verify Pcomp ability | Desirable f(P) that matches with cal sheet with minimal hysteresis. Pcomp works. | A
2022-02-01 | SBET113 | SBET113 | Cycles to 2600 dBar at temp steps of 2:10:30 to verify Pcomp ability | Desirable f(P) that matches with cal sheet with minimal hysteresis. K2 is off and noise is suspect | B
2022-02-07 | SBET114 | SBET114 | Cycles to 2600 dBar at temp steps of 2:10:30 to verify Pcomp ability | Desirable f(P) that matches with cal sheet with minimal hysteresis. Pcomp works. | A
2022-02-07 | SBE11296 | SBE11296_stress_test | Cycles to 2600 dBar at temp steps of 2:10:30 to stress test the high k2 sensor | Verified it is operating well and can be deployed despite the "out of family" k2 values | B
2022-02-18 | GDF016 | GDF016_2022_02_18 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar at beginning | Desirable f(P) after k2 pres correction with slight hysteresis. f(P) was no longer drifty in second half of pressure cycle when temp was increasing between cycle segments. (dubbed temp hysteresis) | A
2022-02-18 | GDF017 | GDF017_2022_02_18 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar at beginning | Desirable but large f(P) after k2 pres correction with slight hysteresis. f(P) became less drifty in second half of pressure cycle when temp was increasing between cycle segments. | B
2022-02-23 | GDF018 | GDF018_2022_02_23 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar at beginning | Run lost connection halfway through. Re-ran on 2/28 | A
2022-02-23 | GDF019 | GDF019_2022_02_23 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar at beginning | Run lost connection halfway through. Re-ran on 2/28 | B
2022-02-28 | GDF018 | GDF018_2022_02_28 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar at beginning | Small but not normal f(P) with slight hysteresis. Some variability and interesting mid-pressure (1000dBar) pressure dependence shift indicating non-uniform stress on FET | A
2022-02-28 | GDF019 | GDF019_2022_02_28 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar at beginning | Undesirable and large f(P) that had large variability past 1500 dBar | B
2022-03-03 | SBE11372 | SBE11372_LPH | Verifying and characterizing the low pressure hook of SBE stem. 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar | Desirable f(P) that matches with cal sheet with minimal hysteresis. LPH is visible but characterizable | A
2022-03-03 | SBE11472 | SBE11472_LPH | Verifying and characterizing the low pressure hook of SBE stem. 2000 dBar full calibration at temps 2:10:30 with break in period to 2400 dBar | Desirable f(P) that matches with cal sheet with minimal hysteresis. LPH is visible but characterizable | B
2022-03-08 | GDF019 | GDF019_oring_2022_03_08 | Replaced gasket w/ oring to retest capabilities. 2000 dBar full cal at temps 2:10:30 with break in period to 2400 dBar. No high TP simultaneously | Small but bimodal f(P) that contained a slip in Vrs at decreasing pressure which indicates oring getting stuck and releasing. k2 appears smaller but hysteresis still exists. | B
2022-03-08 | SBE11455 | SBE11455_drifty_2022_03_08 | Verifying and characterizing "drifty" SBE stem. 2000 dBar full cal at temps 2:10:30 with break in period to 2400 dBar. No high TP simultaneously | Desirable f(P) that matches with cal sheet with minimal hysteresis. No noticeable drift | A
2022-03-14 | GDF019 | GDF19_notopplate_031422 | 10 cycles to 2000 dBar and 20C to see if there is insufficient flushing at the reference pellet, kept o-ring and FET | No indication that it helped or hindered. Decreasing pressure slip is still visible and an increasing pressure slip developed. | B
2022-03-14 | GDF020 | GDF20_prelimtest | 10 cycles to 2000 dBar and 20C | So far so good, ready for full call | A
2022-03-16 | GDF020 | GDF020_2022_03_16 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. High presure hook developed and led to slight variability | A
2022-03-16 | GDF021 | GDF021_2022_03_16 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. | B
2022-03-23 | GDF017 | GDF017_TiCoverplate_2022_03_22 | Testing out Titanium coverplate. 2000 dBar full cal at temps 2:10:30 with break in period to 2200 dBar. No high TP simultaneously | Large and slightly drifty f(P) with no noticeable hysteresis. Long equilibration time for Vrs during temperature decreasing time | B
2022-03-23 | GDF019 | GDF019_new_chip_2022_03_22 | Installed new chip, gasket, peek TS cover plate. 2000 dBar full cal at temps 2:10:30 with break in period to 2200 dBar. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. f(P) was no longer drifty in second half of pressure cycle when temp was increasing between cycle segments. Didn't have much to begin with | A
2022-04-05 | GDF016 | GDF016_newchip_220405 | Installed new chip and gasket. 2000 dBar full cal at temps 2:10:30 with break in period to 2200 dBar. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. | A
2022-04-05 | GDF018 | GDF018_newchip_220405 | Installed new chip and gasket. 2000 dBar full cal at temps 2:10:30 with break in period to 2200 dBar. No high TP simultaneously | Desirable f(P) after k2 pres correction with moderate hysteresis. The minimal f(P) drift it had in first half of cycle went away in second half when temp was increasing between cycle segments | B
2022-04-15 | GDF017 | GDF017_newchip_041522 | Installed new chip and gasket. 2000 dBar full cal at temps 2:10:30 with break in period to 2200 dBar. No high TP simultaneously | Desirable f(P) after k2 pres correction with moderate hysteresis. The minimal f(P) drift it had in first half of cycle went away in second half when temp was increasing between cycle segments | A
2022-04-15 | GDF020 | GDF020_newchip_041522 | Installed new chip and gasket. 2000 dBar full cal at temps 2:10:30 with break in period to 2200 dBar. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. The minimal f(P) drift it had in first half of cycle went away in second half when temp was increasing between cycle segments | B
2022-05-02 | DSD140 | DSD140_test_02May2022 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. Questioning why this was sidelined previously, will test in new chamber for inter-chamber calibration | C
2022-05-02 | GDF015 | GDF15_test_02May2022 | 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Small but slightly drifty f(P) after 1500 dBar when the compensation-activity relationship switched that was much larger than previous tests. Moderate hysteresis. Minimal temp hysteresis. | D
2022-05-03 | GDF010 | GDF010_05032022 | Broached corners. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Desirable f(P) after k2 pres correction with noticeable hysteresis. Abrupt f(P) change at 1500 dBar. Test only lasted break in period and confirmed large noise largely due to chamber issues | E
2022-05-03 | GDF009 | GDF009_05032022 | Broached corners. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Desirable f(P) after k2 pres correction with noticeable hysteresis.Test only lasted break in period and confirmed large noise largely due to chamber issues | F
2022-05-06 | SBE10957 | SBE10957_06May2022 | Testing Buna oring for SBE. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Began with railed Ik and very noisy Vk SD, but ok Vrs (odd). Vk then failed and subsequently Vrs. Failed sensor. | C
2022-05-06 | SBE11592 | SBE11592_06May2022 | Testing Buna oring for SBE. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Displayed a repeatable f(P) with no pressure hysteresis, but had a consistent drift that denotes a temperature hysteresis. Further testing will be done to check if reference or FET is culprit | D
2022-05-06 | DSD140 | DSD140_chambercomp_050622 | Interlab pTC comparison. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Desirable f(P) after k2 pres correction with slight hysteresis. Proved to have a good interlab comparison. Changed pump flow rate from 0.15 to 0.2 during 10C on temp ascent | A
2022-05-06 | SBE11557 | SBE115577_050622 | Testing Buna oring for SBE. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Displayed a repeatable f(P) with no pressure hysteresis, but had a consistent drift that denotes a temperature hysteresis. Testing the counterpart, SBE11592, to see what causes drift. Will repeat in chan B. Changed pump flow rate from 0.15 to 0.2 during 10C on temp ascent | B
2022-05-10 | SBE11557 | SBE115577_051022_2 | Re-Testing Buna oring for SBE. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Rerun for SBE11557. Looked better than previous run; temp hysteresis/ratcheting effect seems to have reduced/gone away. At the end fo this run, did a slow PT cycle (6 hour temp hold, 4 hour to 2000 dbar, 4 hour to 20 dBar; 20->10->20, 2x P cycle each. Repeatability improved, drift at new temp significantly reduced. | B
2022-05-10 | SBE11532 | SBE11532_051022_chanA | Testing Buna oring for SBE. 2000 dBar full calibration at temps 2:10:30 with break in period to 2200 dBar at beginning. No high TP simultaneously | Rerun for SBE11532. Looked better than previous run; temp hysteresis/ratcheting effect seems to have reduced/gone away. At the end fo this run, did a slow PT cycle (6 hour temp hold, 4 hour to 2000 dbar, 4 hour to 20 dBar; 20->10->20, 2x P cycle each. Repeatability improved, drift at new temp significantly reduced. | A
2022-05-13 | SBE11592 | SBE 11592_13May2022_ChC | Testing temperature hysteresis with longer temp holds before pressure cycling. No high TP simultaneously | Rerun for SBS11592, in same lab, but swapped chambers.did a slow PT cycle (6 hour temp hold, 4 hour to 2000 dbar, 4 hour to 20 dBar; 20->10->20, 2x P cycle each. Repeatability improved, drift at new temp significantly reduced. Slow negative drift in Vrs was observed for the first 3/4 of the run. Drift was gone for the last 20C cycle. Maybe chem equilibrium effect? | C
2022-05-13 | DSD140 | DSD140_13May2022_ChD | Testing temperature hysteresis with longer temp holds before pressure cycling. No high TP simultaneously. Switching chambers to see temp lag | DSD140 failed mid run | D
2022-05-17 | SBE11532 | SBE 11532_051722_chanA | Going back to PT cycle in SBE11557_050622. Standard cal run, but Pump flow rate back to 0.15 so stay sat 2000 or 30 dBar for only a couple of minutes | Performed shockingly well with minimal racheting despite faster PT cycle. Minimal hysteresis and desirable f(P). EQ time is more dominant than previously thought | A
2022-05-17 | SBE11557 | SBE 11557_051722_chanB | Going back to PT cycle in SBE11557_050622. Standard cal run, but Pump flow rate back to 0.15 so stay sat 2000 or 30 dBar for only a couple of minutes | Performed shockingly well with minimal racheting despite faster PT cycle. Minimal hysteresis and desirable f(P). EQ time is more dominant than previously thought | B
2022-05-17 | SBE11592 | SBE 11592_17May2022_ChC | longer temp holds and slow pump rate - testing temperature hysteresis | stopped after one day to replace GDF15 with GDF22 | C
2022-05-17 | GDF015 | GDF 15_17May2022_ChD | longer temp holds and slow pump rate - testing temperature hysteresis | stopped after one day to replace GDF15 with GDF22 | D
2022-05-17 | GDF010 | GDF010_051722_Ch E | First time using PTC E/F. Ran regular PTC calibration (PTcycle_GDF_nohighTP) 20-10-5-2 deg steps | Desirable f(P) after k2 pres correction with moderate hysteresis. Going to be SW equilibrated | E
2022-05-17 | GDF009 | GDF009_051722_Ch F | First time using PTC E/F. Ran regular PTC calibration (PTcycle_GDF_nohighTP) 20-10-5-2 deg steps | Desirable f(P) after k2 pres correction with moderate hysteresis. Drift rate about 300uV/day. Going to be SW equilibrated | F
2022-05-18 | SBE11592 | SBE 11592_18May2022_ChC | longer temp holds and slow pump rate - testing temperature hysteresis | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can. Extended test to check if drift in other chamber was due to a bubble. Temperature ramps at high pressure. Ok | C
2022-05-18 | GDF022 | GDF 22_18May2022_ChD | longer temp holds and slow pump rate - testing temperature hysteresis |  | D
2022-05-24 | SBE11529 | SBE11529_052422_Ch E | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously | Displayed a repeatable f(P) with no pressure hysteresis, but had a slight drift that denotes a temperature hysteresis. Further testing will be done to check if reference or FET is culprit | E
2022-05-24 | SBE11533 | SBE11533_052422_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously | Displayed a repeatable f(P) with no pressure hysteresis, but got noisy throughout the run. Did not affect f(P) Going into TiP can | F
2022-05-24 | SBE11031 | SBE11031_052422_chanA | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously | Displayed a repeatable f(P) with no pressure hysteresis, but had a slight drift that denotes a temperature hysteresis and noise decreased. Further testing will be done to check if reference or FET is culprit | A
2022-05-24 | SBE11043 | SBE11043_052422_chanB | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | B
2022-06-01 | GDF023 | GDF023_060122_ChC | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles | Very drifty and could not stabilize. Will replace ISFET chip | C
2022-06-01 | GDF024 | GDF024_060122_ChD | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles | Very drifty and could not stabilize. Will replace ISFET chip | D
2022-06-01 | SBE11031 | SBE11031_060122_chanA | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles |  | A
2022-06-01 | GDF025 | GDF025_060122_chanB | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles |  | B
2022-06-01 | SBE11529 | SBE11529_060122_Ch E | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles |  | E
2022-06-01 | GDF026 | GDF026_060122_Ch F | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles |  | F
2022-06-08 | SBE11563 | SBE11563_060822_ChC | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" | Railed Ik, Vrs_std very high (mV's). Vrs looks OK though. Probably a sensor about to fail. Run/computer crashed on 6/11/2022 5AM ish. Restarted run on 6/12/2022 | C
2022-06-08 | SBE11555 | SBE11555_060822_ChD | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" | Kind of drifty Vrs, but could be the zoom. Otherwise looked good. Run/computer crashed on 6/11/2022 5AM ish. Restarted run on 6/12/2022 | D
2022-06-10 | GDF027 | GDF027_061022_chanA | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | A
2022-06-10 | SBE11564 | SBE11564_061022_chanB | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | B
2022-06-10 | SBE11746 | SBE11746_061022_Ch E | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" | poor sensor response and shape of Vrs signal | E
2022-06-10 | SBE11747 | SBE11747_061022_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | F
2022-06-12 | SBE11563 | SBE11563_061222_ChC | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Sample every 30 sec | sensor failed with railed Vrs - returned to SeaBird for diagnosis | C
2022-06-12 | SBE11555 | SBE11555_061222_ChD | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Sample every 30 sec | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | D
2022-06-21 | GDF026 | GDF026_062122_chanA | rerun of GDF26 |  | A
2022-06-21 | SBE11591 | SBE11591_062122_chanB | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Sample every 30 sec | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | B
2022-06-27 | GDF023 | GDF23_062722_Ch E | rerun of GDF23 and GDF24. Ran in C and D before. Longer temp holds and slow pump rate. "SBE test" |  | E
2022-06-27 | GDF024 | GDF24_062722_Ch F | rerun of GDF23 and GDF24. Ran in C and D before. Longer temp holds and slow pump rate. "SBE test" |  | F
2022-07-07 | GDF028 | GDF028_070722_chanA | GDF28 was manufactured during AE training. Longer temp holds and slow pump rate. "SBE test" |  | A
2022-07-07 | GDF023 | GDF023_070722_chanB | rerun of GDF23 with new gasket and new ISFET. Ran in C and D before. Longer temp holds and slow pump rate. "SBE test". Final try, three strikes you're out |  | B
2022-07-12 | SBE11395 | SBE11395_071222_Ch E | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Modified pellet design | Drifted low for the first half of test until 2C. Will re-run | E
2022-07-12 | SBE11762 | SBE11762_071222_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Modified pellet design | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | F
2022-07-15 | SBE11568 | SBE11568_071522_chanA | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Modified pellet design | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | A
2022-07-15 | SBE11597 | SBE11597_071522_chanB | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Modified pellet design | Displayed a repeatable f(P) with no pressure hysteresis, going into TiP can | B
2022-07-20 | SBE11395 | SBE11395_072022_Ch E | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Modified pellet design. ReRun | Passed. Going on APEX Float | E
2022-07-20 | SBE11875 | SBE11875_072022_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Modified pellet design | Passed. Going on APEX Float | F
2022-07-27 | GDF015 | GDF015_072722_RDC_chanA | Measuring the stability of refurbished ISFET chip in used GDF housing. Recyclyed durafet chip (RDC). "SBE test" |  | A
2022-07-27 | GDF023 | GDF023_072722_HB_chanB | Measuring the stability of refurbished ISFET chip in used GDF housing. Honeywell batch (HB). "SBE test" |  | B
2022-08-15 | GDF029 | GDF029_081622_Ch E | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles | f(P) jumped halfway into test at pressures beyond 1500 dBar. Will replace gasket with neoprene one and re run alongside GDF30 | E
2022-08-15 | GDF030 | GDF030_081622_Ch F | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles | f(P) was drift and splayed out at pressures beyond 1500 dBar. Will run in pTC again alongside GDF29 | F
2022-08-18 | GDF031 | GDF31_081822_ChC | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles | Desirable f(P) after k2 pres correction with slight hysteresis. The minimal f(P) drift it had in first half of cycle went away in second half when temp was increasing between cycle segments | C
2022-08-18 | GDF032 | GDF32_081822_ChD | Longer tempholds and slow pump rate. Do temp changes at high pressure, and then decrease pressure for P cycles | Desirable f(P) after k2 pres correction with slight hysteresis. The minimal f(P) drift it had in first half of cycle went away in second half when temp was increasing between cycle segments | D
2022-08-17 | PICO | Pico22290009_081722_chanB | 24 hour test of new PICO pH sensor. Pressure cycling to 2100dBar at 20,10,5C in 80% DIW and 20% Equimolar Tris Batch #3 |  | B
2022-08-18 | PICO | Pico22290009_081822_chanB | 24 hour test of new PICO pH sensor. Pressure cycling to 2100dBar at 35,20,5C in 80% DIW and 20% Equimolar Tris Batch #3. GDF015 shut off from pres |  | B
2022-08-22 | GDF015 | GDF015_RFD_082222_chanA | SBEtest alongside PICO pH sensor. No high TP simultaneously, then many cycles at 20C in 80% DIW and 20% Equimolar Tris Batch #3 |  | A
2022-08-22 | PICO | Pico22290009_082222_chanB | SBEtest with PICO pH sensor. No high TP simultaneously, then many cycles at 20C in 80% DIW and 20% Equimolar Tris Batch #3 |  | B
2022-08-24 | SBE11987 | SBE11987_082422_Ch E | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | E
2022-08-24 | SBE11988 | SBE11988_082422_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | F
2022-08-25 | SBE11991 | SBE11991_082522_ChC | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | C
2022-08-25 | SBE11992 | SBE11992_082522_ChD | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | D
2022-09-01 | GDF025 | GDF025_090122_Ch E | Retest this stem. longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | E
2022-09-01 | GDF029 | GDF029_090122_Ch F | Neoprene gasket installed and retest. longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest" |  | F
2022-09-12 | PICO | Pico22290009_091222_chanA | Sodium carbonate |  | A
2022-09-12 | PICO | Pico22290014_091222_chanB | Sodium carbonate |  | B
2022-09-16 | PICO | Pico22290009_top_091622_chanA | citric acid |  | A
2022-09-16 | PICO | Pico22290014_top_091622_chanB | citric acid |  | B
2022-09-19 | SBE12026 | SBE12026_091922_Ch E | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Viton | Passed. Going on APEX Float. Also known as 9/26/22 test in pHlog | E
2022-09-19 | SBE12027 | SBE12027_091922_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Viton | Passed. Going on APEX Float. Also known as 9/26/22 test in pHlog | F
2022-09-09 | DSD371 | DSD371_090922_ChC |  |  | C
2022-09-09 | DSD372 | DSD372_090922_ChD |  |  | D
2022-09-23 | SBE12028 | SBE12028_092322_ChC | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Viton | Passed. Going on APEX Float | C
2022-09-23 | SBE12114 | SBE12114_092322_ChD | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Viton | Passed. Going on APEX Float | D
2022-09-21 | PICO | Pico22290009_top_092122_chanA | citric acid and sodium citrate buffer. SBE test but starting at lower temps to test the stability of k2 |  | A
2022-09-21 | PICO | Pico22290014_top_092122_chanB | citric acid and sodium citrate buffer. SBE test but starting at lower temps to test the stability of k2 |  | B
2022-09-23 | PICO | Pico22290009_092322_chanA | citric acid and sodium citrate buffer. SBE test but starting at lower temps to test the stability of k2 going from low to higher temperatures |  | A
2022-09-23 | PICO | Pico22290014_092322_chanB | citric acid and sodium citrate buffer. SBE test but starting at lower temps to test the stability of k2 going from low to higher temperatures |  | B
2022-09-13 | PICO | Pico22290009_091322_chanA | sodium carbonate buffer for bottom cal |  | A
2022-09-13 | PICO | Pico22290014_091322_chanB | sodium carbonate buffer for bottom cal |  | B
2022-10-04 | GDF030 | GDF030_100422_Ch E | Re-run of GDF030 to see if now stable w/ no hardware changes. "SBE stest". |  | E
2022-10-04 | SBE12115 | SBE12115_100422_Ch F | longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "SBE stest". Viton o-ring build |  | F
2022-10-11 | GDF029 | GDF029_101122_Ch E_viton | Longer tempholds and slow pump rate. No high TP. "SBE Test". Re-run with Viton gasket |  | E
2022-10-11 | GDF030 | GDF030_101122_Ch F_neoprene | Longer tempholds and slow pump rate. No high TP. "SBE Test". Re-run with Neoprene gasket. Known good cal with Viton |  | F
2022-10-18 | GDF029 | GDF029_101822_Ch E_viton_BATTERY | Repeating previous test with a 12V battery as a power source | Low noise levels and consistent, test stopped shjort due to channel F | E
2022-10-18 | GDF030 | GDF030_101822_Ch F_neoprene_BATTERY | Repeating previous test with a 12V battery as a power source | Began with low noise levels but then std devs and Ik/Ib began to grow in orders of magnitude. Not sure if Neoprene gasket or the battery was culprit. Will replace with viton gasket | F
2022-10-19 | GDF029 | GDF029_101922_Ch E_viton_BATTERY_2 | Repeating previous test with a 12V battery as a power source, 2nd run |  | E
2022-10-19 | GDF030 | GDF030_101922_Ch F_viton_BATTERY_2 | Repeating previous test with a 12V battery as a power source, 2nd run after previous one stopped short. Replaced neoprene with viton gasket |  | F
2022-10-26 | GDF029 | GDF029_101926_Ch E_viton_BATTERY_2 | Repeating previous test with a 12V battery as a power source, 3rd run |  | E
2022-10-26 | GDF030 | GDF030_101926_Ch F_viton_BATTERY_2 | Repeating previous test with a 12V battery as a power source, 3rd run | good calibration run but no significant improvement in std devs of Vrs and Vk | F
2022-11-03 | GDF029 | GDF029_110322_Ch E_new DF board on Ch F | wall power and new DF adapter board on Chan F |  | E
2022-11-03 | GDF030 | GDF030_110322_Ch F_new DF board on Ch F | wall power and new DF adapter board on Chan F |  | F
2022-11-16 | PICO | Pico22360017_111522_chanA | Testing Pico sensor in Tris batch #7 that will be on spray glider. It has been full factory calibrated |  | A
2022-11-28 | PICO | Pico22180003_112822_chanA | Testing Pico sensor in Tris batch #7 that will be on APEX float to 2000 dBar. PEEK adapter. It has been full factory calibrated | IDEX connector wasn't fully secured and a massive leak occurred in the base piece. PICO flooded :( | A
2022-12-01 | PICO | Pico22290009_120122_chanA | Testing Pico sensor in Tris batch #7 that will be on APEX float to 1000 dBar. Delrin adapter. It has been full factory calibrated | Error message froze pTC labview program. Data that got through looked good. Re run to 2000 dBar since Delrin adapter piece was approved for APEX float rating. | A
2022-12-05 | PICO | Pico22290009_120522_chanA | Testing Pico sensor in Tris batch #7 that will be on APEX float to 2000 dBar. Delrin adapter. It has been full factory calibrated |  | A
2022-12-09 | GDF2009 | AE_GDF2009_120922_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | E
2022-12-09 | GDF2010 | AE_GDF2010_120922_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Sensor failed catastrophically with the ISFET cracking into four pieces on the first pressure cycle to 2000dBar | F
2022-12-13 | PICO | Pico22180006_121322_chanA | Testing Pico sensor in Tris batch #7 that will be on APEX float to 2000 dBar. Delrin adapter. It has been full factory calibrated |  | A
2022-12-19 | PICO | Pico22180006_121922_chanA | Testing signal intensity range for PICO after reassembly for float pico 22180006 | Verified that signal intensity changes as it should | A
2022-12-20 | GDF2000 | AE_GDF2000_122022_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test". Adding a long wait time after analyzing data | Odd f(P). will run again due to test not being a full calibration cycle and long drift time | E
2022-12-20 | GDF2001 | AE_GDF2001_122022_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temperature hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test". Adding a long wait time after analyzing data | Passed, will run again due to test not being a full calibration cycle and long drift time | F
2022-12-28 | GDF2000 | AE_GDF2000_122022_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Bad f(P) beyond ~1500 dBar. Will replace coverplate with Ti coverplate to test which components contribute to these results | E
2022-12-28 | GDF2001 | AE_GDF2001_122022_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F
2023-01-04 | GDF2000 | AE_GDF2000_Ti cvrplt_010423_Ch E | Testing AE GDF with Ti COVER PLATE no shoulder. No high TP simultaneously, no 30C. "MBARI GDF test" | Good calibration but will disassemble to put the peek coverplate (tall shoulder) back on to see if the Ti cover plate (no shoulder) actually improved the cal | E
2023-01-04 | GDF2002 | AE_GDF2002_010423_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Good calibration | F
2023-01-11 | GDF2000 | AE_GDF2000_PEEK tall shldr_011123_Ch E | Testing AE GDF build with the original tall shoulder PEEK cover plate after performing well with the Ti no shoulder cover plate |  | E
2023-01-11 | GDF2003 | AE_GDF2003_011123_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F
2023-01-11 | GDF2004 | GDF2004_011123_chanA | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | High noise and bad MSC controller. Will re-Run same test with new electronics | A
2023-01-11 | GDF2005 | GDF2005_011123_chanB | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | High noise and bad MSC controller. Will re-Run same test with new electronics | B
2023-01-17 | GDF2004 | GDF2004_011723_chanA | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | A
2023-01-17 | GDF2005 | GDF2005_011723_chanB | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B
2023-01-18 | GDF2000 | AE_GDF2000_PEEK no shldr_011823_Ch E | Testing AE GDF build with the original PEEK cover plate, but with no shoulder after performing well with the Ti no shoulder cover plate |  | E
2023-01-18 | GDF2003 | AE_GDF2003_PEEK no shldr_011823_Ch F | Testing AE GDF build with the original PEEK cover plate, but with no shoulder after performing well with the Ti no shoulder cover plate |  | F
2023-01-26 | GDF2000 | AE GDF 2000 Ti cvrplt_012623 | Testing AE GDF build with the Ti no shoulder cover plate. Tmep cycles at same pressure. Gene Routine 2 |  | E
2023-01-26 | GDF2003 | AE GDF 2003 Ti cvrplt_012623 | Testing AE GDF build with the Ti no shoulder cover plate. Tmep cycles at same pressure. Gene Routine 2 |  | F
2023-01-26 | GDF2007 | AE GDF2007_012623_Ch A | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Labview timeout after laptop restarted. Will re-run | A
2023-01-26 | GDF2008 | AE GDF2008_012623_Ch B | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Labview timeout after laptop restarted. Will re-run | B
2023-02-03 | GDF2003 | AE GDF2003_020323 | Testing time constant in AE GDF build after undergone several different calibration tests in this configuaration. Ti coverplate. |  | F
2023-02-09 | GDF2007 | GDF2007_020923_Ch A | Re Run after laptop restarted and labview timed out. Testing AE GDF build. longer temp holds and slow pump rate. "MBARI GDF test" | Valves were closed….. Noticed a small leak past the chamber B valve, need to retighten it | A
2023-02-09 | GDF2008 | GDF2008_020923_Ch B | Re Run after laptop restarted and labview timed out. Testing AE GDF build. longer temp holds and slow pump rate. "MBARI GDF test" | Valves were closed….. Noticed a small leak past the chamber B valve, need to retighten it | B
2023-02-13 | GDF2007 | GDF2007_021323_Ch A | Re Run after finding out valves were closed. Testing AE GDF build. longer temp holds and slow pump rate. "MBARI GDF test" |  | A
2023-02-13 | GDF2008 | GDF2008_021323_Ch B | Re Run after finding out valves were closed. Testing AE GDF build. longer temp holds and slow pump rate. "MBARI GDF test" |  | B
2023-03-31 | GDF2008 | GDF2008_033123_Ti cvrplt_Ch E | Testing AE GDF build with the Ti no shoulder cover plate. Tmep cycles at same pressure. Gene Routine 3 | Long equilibration period, but no other issues. Will re-run | E
2023-03-31 | GDF2003 | GDF2003_033123_Ti cvrplt_Ch F | Testing AE GDF build with the Ti no shoulder cover plate. Tmep cycles at same pressure. Gene Routine 3 | Grade A student. | F
2023-03-31 | GDF2000 | GDF2000_033123 Ti cvrplt_Ch B | Testing AE GDF build With Ti Cover Plates. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" | Slow Equilibration time and produced poor calibration. Will re-run | B
2023-03-31 | GDF2007 | GDF2007_033123 Ti cvrplt_Ch A | Testing AE GDF build With Ti Cover Plates. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" | Ok with High temperature variation. Will re-run | A
2023-04-11 | GDF2000 | GDF2000_041123 Ti cvrplt_Ch B | Testing AE GDF build With Ti Cover Plates. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B
2023-04-11 | GDF2007 | GDF2007_041123 Ti cvrplt_Ch A | Testing AE GDF build With Ti Cover Plates. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | A
2023-05-17 | GDF2005 | GDF2005_051723_Ch E | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. | Slow Equilibration time and produced poor calibration. Will re-run | E
2023-05-17 | GDF2002 | GDF2002_051723_Ch F | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. | Slow Equilibration time and produced poor calibration. Will re-run | F
2023-05-21 | GDF2005 | GDF2005_052123_Ch E | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. | Passed. Will compare with full Gene routine | E
2023-05-21 | GDF2002 | GDF2002_052123_Ch F | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. | Passed. Will compare with full Gene routine | F
2023-05-25 | GDF2005 | GDF2005_052523_Ch E | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. | Passed and similar to original calibration from January 2023 | E
2023-05-25 | GDF2002 | GDF2002_052523_Ch F | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. | Passed and similar to original calibration from January 2023 | F
2023-06-06 | GDF2008 | GDF2008_060623_lcgasket_Ch A | Re-run and testing AE GDF build with laser cut gasket. No high TP simultaneously, no 30C. "MBARI GDF test" | Repeatable pressure and temperature responses with minimal hysteresis | A
2023-06-06 | GDF2004 | GDF2004_060623_lcgasket_Ch B | Re-run and testing AE GDF build with laser cut gasket. No high TP simultaneously, no 30C. "MBARI GDF test" | Fairly repeatable but slightly larger hysteresis than 2008. Visible gasket extrusion into the sensing area, will rebuild with slightly larger ID | B
2023-06-23 | GDF2004 | GDF2004_062323_Ch E | Re Run of AE GDF 2004 with a slightly larger ID laser cut gasket (0.112 "). "MBARI GDF test" | Held well at 20C and high pressure, seemed to have leaked at low temperature and high pressure. Run stopped midwaty | E
2023-07-07 | GDF2004 | GDF2004Ti_070723_Ch E | Testing AE GDF build With Ti Cover Plates. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" | probably enumerated incorrectly and reason why data isn't in the performances directory | E
2023-07-21 | GDF2011 | GDF2011_072123_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | E
2023-07-21 | GDF2012 | GDF2012_072123_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F
2023-08-02 | GDF2011 | GDF2011_080223_Ch E | Re-run and testing AE GDF build. "Gene routine" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. |  | E
2023-08-02 | GDF2012 | GDF2012_Ti cvrplt_080223_Ch F | Testing AE GDF build With Ti Cover Plates. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F
2023-08-09 | GDF2013 | GDF2013_080923_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | E
2023-08-09 | GDF2014 | GDF2014_080923_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F
2023-08-17 | GDF2015 | GDF2015_081723_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | These were the ones that were put on seawater and it they need an extra run. Drifty, odd behavior at high pressures and large residuals to its f(P) | F
2023-08-17 | GDF2017 | GDF2017_081723_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | These were the ones that were put on seawater and it they need an extra run. Drifty, odd behavior at high pressures and large residuals to its f(P) | E
2023-08-24 | GDF2015 | GDF2015_082423_Ch F | Re-run and testing AE GDF build. "MBARI GDF Test" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. |  | F
2023-08-24 | GDF2017 | GDF2017_082423_Ch E | Re-run and testing AE GDF build. "MBARI GDF Test" for slow pressurization rate, step changes. No high TP simultaneously, no 30C. |  | E
2023-08-30 | GDF2016 | GDF2016_083023_Ch A | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Odd performance on increasing pressure indicating pinched gasket. Replaced and will re run | A
2023-08-30 | GDF2018 | GDF2018_083023_Ch B | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B
2023-09-01 | GDF2017 | GDF2017_090123_Ch E | Built with a new gasket after it was seen to be extruded. Re-run with *NEW* DIPSE. Small low pressure run to see if DISPE works |  | E
2023-09-01 | DIPSE001 | DIPSE001_090123_Ch F | Testing *NEW* DISPE !!! | PASSED, will run modified normal run to 500 dBar | F
2023-09-01 | GDF2017 | GDF2017_090123_500dBar_Ch E | Re run of AE GDF build. Modified "MBARI GDF test" to only 500dBar maximum so DIPSE can safely pressurize with no parback |  | E
2023-09-01 | DIPSE001 | DIPSE001_090123_500dBar_Ch F | Testing DIPSE. Modified "MBARI GDF test" to only 500dBar maximum so DIPSE can safely pressurize with no parback |  | F
2023-09-01 | GDF2016 | GDF2016_090123_Ch A | Replaced gasket. Testing AE GDF build. longer temp holds and slow pump rate No high TP simultaneously, no 30C. "MBARI GDF test" | Passed | A
2023-09-01 | GDF2018 | GDF2018_090123_Ch B | Testing AE GDF build. longer temp holds and slow pump rateNo high TP simultaneously, no 30C. "MBARI GDF test" | Developed odd hysteresis that indicated pinched gasket. | B
2023-09-07 | DIPSE001 | DIPSE001_090723_Ch E | Testing DIPSE to full pressure. Modified "MBARI GDF test" to have a very slow foirst pressurization and then resuming normal calibration run | ISFET ruptured upon reaching 2000 dBar. Successful before that. It worked until it didn't | E
2023-09-07 | DIPSE002 | DIPSE002_090723_Ch F | Testing DIPSE to full pressure. Modified "MBARI GDF test" to have a very slow foirst pressurization and then resuming normal calibration run | ISFET ruptured upon reaching 2000 dBar. Successful before that. It worked until it didn't | F
2023-09-14 | GDF2017 | GDF2017_091423_Ch E | Re Run of GDF2017 to assess SBE Rev A Lionix chip, paddle board design. Normal "MBARI GDF Test" | Calibrated and passed | E
2023-09-14 | SBEL003 | SBE_RevA_091423_Ch F | Initial test of SBE Rev A Lionix chip, paddle board design. Normal "MBARI GDF Test" | Worked really well. Linear f(P) and k2. | F
2023-09-15 | GDF2018 | GDF2018_09152023_Ch A | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" | Passed | A
2023-09-15 | GDF2019 | GDF2019_09152023_Ch B | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" | Needs a re run | B
2023-09-25 | DIPSEP001 | GDFLPoly001_092523_Ch E | Initial test of GDFL "DISPEP" test with POLYMIDE coating alongside SBE Rev A Lionix paddle baord design. Normal GDF test but only to 1000 dBar | Works well, not taking it to 2000 dBAr however. Linear f(P) and k2 | E
2023-09-25 | SBEL003 | SBE_RevA_092523_Ch F_50uA_Ids | Testing out SBE Rev A Lionix paddle board chip at 50 uA Ids. Normal GDF test but only to 1000 dBar | f(P) shifted but same magnitide. K2 has slightly larger dependency which is weird | F
2023-09-25 | GDF2020 | GDF2020_09252023_Ch A | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | A
2023-09-25 | GDF2021 | GDF2021_09252023_Ch B | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B
2023-10-02 | DIPSEP001 | GDFLPoly001_50uA_Ids100223_Ch E | Testing out GDFL "DIPSEP" POLYMIDE coated chip with 50 uA Ids. Normal GDF test but only to 1000 dBar | Worked well but similar f(P) and k2 values | E
2023-10-02 | SBEL003 | SBE_RevA_100223_Ch F_mfet board | Testing out SBE Rev A Lionix paddle board but on mFET controller | Worked normally as it did with 100uA Ids with MSC board | F
2023-10-04 | GDF2020 | AE2020_Ti100423_Ch E | Re-testing AE GDF build With Ti Cover Plate. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test". A re-run |  | E
2023-10-04 | GDF2021 | AE2021Ti_100423_Ch F | Re-testing AE GDF build With Ti Cover Plate. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test". A re-run |  | F
2023-10-04 | GDF2019 | GDF2019_100423_Ch A | Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test". A re-run |  | A
2023-10-04 | GDF2022 | GDF2022 Ti_100423_Ch B | Testing AE GDF build With Ti Cover Plate. longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B
2023-10-12 | GDF2025 | GDF2025_101223_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | E
2023-10-12 | GDF2026 | GDF2026_101223_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F
2023-10-12 | GDF2023 | GDF2023_101223_Ch A | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Funky shape | A
2023-10-12 | GDF2024 | GDF2024_101223_Ch B | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B | [7 empty] | 1
2023-10-19 | GDF2025 | GDF2025_10192023_Ch A | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | A | [7 empty] | 1
2023-10-19 | GDF2023 | GDF2023_10192023_Ch B | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B | [7 empty] | 1
2023-10-19 | GDF2024 | GDF2024_10192023_Ch E | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | E | [7 empty] | 1
2023-10-19 | GDF2026 | GDF2026_10192023_Ch F | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | F | [8 empty] | 1
2023-10-26 | GDF2027 | GDF2027_10262023_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | funky bump at 800dBar. will rerun | E | [8 empty] | 1
2023-10-26 | GDF2028 | GDF2028_10262023_Ch F | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | passed | F | [8 empty] | 1
2023-10-27 | GDF2019 | GDF2019_10272023_Ch A | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | A | [8 empty] | 1
2023-10-27 | GDF2030 | GDF2030_10272023_Ch B | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B
2023-11-02 | GDF2029 | GDF2029_11022023_Ch E | Testing AE GDF build. longer temp holds and slow pump rate - testing temp hysteresis. No high TP simultaneously, no 30C. "MBARI GDF test" | Looked ok but odd environmental noise on both sensors that skewed some of the data and caused high residuals. Will re-run in same solution. | E
2023-11-02 | GDF2021 | GDF2021_11022023_Ch F | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" | Looked ok but odd environmental noise on both sensors that skewed some of the data and caused high residuals. Will re-run in same solution. | F
2023-11-10 | GDF2029 | GDF2029_11102023_no breakin_Ch E | Re-testing AE GDF build. longer temp holds and slow pump rate - No high TP simultaneously, no 30C. "MBARI GDF test SHORT". No break-in |  | E
2023-11-10 | GDF2021 | GDF2021_11102023_no breakin_Ch F | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test SHORT" |  | F
2023-11-10 | GDF2023 | GDF2023_111023_Ch A | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | A
2023-11-10 | GDF2024 | GDF2024_111023_Ch B | Replaced gasket. Re-testing AE GDF build with longer temp holds and slow pump rate. No high TP simultaneously, no 30C. "MBARI GDF test" |  | B

## O-ring
Using best results for each type of test |  |  |  | f(P) is k2 pressure corrected
| Folder Name | Seal Type | f(P) at 2000 dBar [mV] | ∆f(P) at 2000 dBar [mV] | Hysteresis? | Magnitude Hysteresis (Max.) [mV] | ∆k2(P) [uV/C] | Summary
GDF 007 | GDF7_retest_gasket | Gasket | 13.5 | 0.6 | no | 0 | 150 | Larger f(P) and k2 pressure correction but minimal hysteresis
| GDF7_oring | O-ring | 7 | 0.6 | yes | 3 | 35 | Smaller f(P) but hysteresis remains
GDF 019 | GDF019_new_chip_2022_03_22 | Gasket | 6 | 0.7 | yes | 2 | 120 | Larger f(P) and k2 pressure correction but smaller hysteresis, less of a spread
| GDF019_oring_2022_03_08 | O-ring | 4 | 1 | yes | 3 | 90 | Smaller f(P) but hysteresis remains. The f(P) is more noisy relative to size

## Ti coverplate
Using best results for each type of test |  |  |  |  | f(P) is k2 pressure corrected
| Test Date | Folder Name | Coverplate | f(P) at 2000 dBar [mV] | ∆f(P) at 2000 dBar [mV] | Hysteresis? | Magnitude Hysteresis (Max.) [mV] | ∆k2(P) [uV/C] | Summary
GDF 017 | 2022-02-18 | GDF017_2022_02_18 | Peek | 9 | 1.5 | yes | 1.5 | 160 | Large f(P), spread, and k2 pressure dependence but moderate hysteresis
| 2022-03-22 | GDF017_TiCoverplate_2022_03_22 | Titanium | 11 | 2.5 | no | 0 | 40 | Larger f(P) and spread but no hysteresis and minimal k2 pressure dependence
| 2022-04-15 | GDF017_newchip_041522 | Peek | 6 | 0.8 | yes | 2 | 130 | Smaller f(P) and spread but moderate hysteresis and k2 pressure dependence

## New ISFET
Using most recent test with new chip for "New" |  |  |  | f(P) is k2 pressure corrected
| Folder Name | ISFET Chip Age | f(P) at 2000 dBar [mV] | ∆f(P) at 2000 dBar [mV] | Hysteresis? | Magnitude Hysteresis (Max.) [mV] | ∆k2(P) [uV/C] | Summary
GDF 016 | GDF016_2022_02_18 | Old | 5 | 2.5 | no | 0 | 140 | Large spread of f(P) relative to size but minimal hysteresis. Got better with calibration runtime
| GDF016_newchip_220405 | New | 5 | 0.9 | yes | 2.8 | 120 | Acceptable spread of f(P) but noticeable hysteresis
GDF 017 | GDF017_2022_02_18 | Old | 9 | 1 | yes | 1 | 140 | Larger f(P) and acceptable spread but noticeable hysteresis and k2 pressure dependence
| GDF017_TiCoverplate_2022_03_22 | Old | 11 | 2.5 | no | 0 | 40 | Larger f(P) and spread but no hysteresis and minimal k2 pressure dependence
| GDF017_newchip_041522 | New | 6 | 0.8 | yes | 2 | 130 | Smaller f(P) and spread but moderate hysteresis and k2 pressure dependence
GDF 018 | GDF018_2022_02_23 | Old | 5 | 5 | no | 0 | 190 | Erratic change in Vrs during test so f(P) was unusable. Failed
| GDF018_2022_02_28 | Old | 0 | 1.5 | yes | 2 | 70 | Odd but somewhat repeatable f(P), but large spread relative to size. Small k2 pressure dependence
| GDF018_newchip_220405 | New | 7 | 1 | yes | 3 | 130 | Large(ish) but aceptable f(P) and spread with moderate hysteresis and k2 pressure dependence
GDF 019 | GDF019_2022_02_23 | Old | 8 | 2 | yes | 2 | 60 | Large(ish) f(P) and spread with moderate hysteresis but minimal k2 pressure dependence
| GDF019_2022_02_28 | Old | 5 | 5 | yes | 3 | 110 | Sensor had odd decreasing pressure failure that resulted in erratic f(P)
| GDF019_oring_2022_03_08 | Old | 4 | 1 | yes | 3 | 90 | Smaller f(P) but hysteresis remains. The f(P) is more noisy relative to size
| GDF019_new_chip_2022_03_22 | New | 6 | 0.7 | yes | 2 | 120 | Larger f(P) and k2 pressure correction but smaller hysteresis, less of a spread
GDF 020 | GDF020_2022_03_16 | Old | 5 | 1.2 | yes | 1.5 | 140 | Small f(P) but questionable spread at high pressures with large k2 pressure dependence
| GDF020_newchip_041522 | New | 5 | 1 | yes | 2 | 60 | Small f(P) with consistent (acceptable) spread throughout profile that improved over calibration run. Small k2 pressure dependence

## Durometer change
f(P) is k2 pressure corrected
| Folder Name | Durometer | f(P) at 2000 dBar [mV] | ∆f(P) at 2000 dBar [mV] | Hysteresis? | Magnitude Hysteresis (Max.) [mV] | ∆k2(P) [uV/C] | Summary
GDF 008 | GDF8_fP_cal | 75A | 2.5 | 0.8 | yes | 0.9 | 115 | Smaller (but odd) f(P) and spread and minimal hysteresis but moderate k2 pressure dependence
| GDF8_fP_cal_new gasket | 75A | 4 | 0.5 | yes | 2 | N/A | Small and desirable f(P) with minimal spread but large hysteresis relative to f(P) size
| GDF8_90A gasket | 90A | Sensor failed |  |  |  |  | Sensor failed. Don't do this

## Pressure cycles
All tests for each GDF S/N are displayed in chronological order
Red folder name indicates test with no high temperature and pressure simultaneously |  |  |  | f(P) is k2 pressure corrected
Purple presure indicates tests that went to highest pressure
| Folder Name | Max Presure (dBar) | f(P) at 2000 dBar [mV] | ∆f(P) at 2000 dBar [mV] | Hysteresis? | Magnitude Hysteresis (Max.) [mV] | ∆k2(P) [uV/C] | Summary
GDF 011 | GDF11_machinedcorners | 2000 | 11 | 0.7 | no | 0 | N/A | Not a full calibration but large f(P) with minimal spread
| GDF11_machinedcorners_2 | 2000 | 5 | 0.7 | yes | 1.5 | 145 | Smaller f(P) with desirable spread and small hysteresis. Moderate k2 pressure dependence
| GDF11_fullcal | 2600 | 6 | 1 | yes | 1.5 | 140 | Small(ish) f(P) with acceptable spread with small hysteresis. Moderate k2 pressure dependence
GDF 012 | GDF12_machinedcorners | 2000 | 5 | 0.8 | yes | 1 | 130 | Smaller f(P) with desirable spread and small hysteresis. Moderate k2 pressure dependence
| GDF12_stressTest_2000dBar | 2000 | N/A | N/A | yes | N/A | N/A | Used to stress the sensor for change in performance but minimal altering. The resulting f(P) was hard to caluclate due to rapid changes in pressure and temperature
| GDF12_2600_dBar | 2600 | 5 | 1.5 | yes | 1.8 | 120 | Small(ish) f(P) with questionable spread and moderate hysteresis. Moderate k2 pressure dependence
| GDF12_cal_every2C | 2000 | 6 | 0.8 | yes | 1.8 | 145 | Smaller f(P) with desirable spread and moderate hysteresis. Moderate k2 pressure dependence. Lower pressure improved spread
GDF 015 | GDF15_1216 | 2600 | 5.5 | 0.8 | yes | 1 | 150 | Small(ish) f(P) with desirable spread and small hysteresis. Moderate k2 pressure dependence
| GDF15_ChanA_2022_01_15 | 2600 | 4 | 2 | yes | 1.5 | 90 | Small f(P) with large spread and moderate hysteresis. Small k2 pressure dependence. This cycle permanently strained the sensor performance significantly
| GDF15_ChanB_2022_01_24 | 2600 | 4 | 1.5 | yes | 1.5 | 90 | Small f(P) with large relative spread and moderate hysteresis. Small k2 pressure dependence
| GDF15_test_02May2022 | 2200 | 2 | 1.5 | yes | 1.8 | 70 | Small f(P) with large relative spread and hysteresis. Small k2 pressure dependence. The f(P) changed shape from previous tests
GDF 016 | GDF016_2022_02_18 | 2400 | 5 | 2.5 | no | 0 | 140 | Large spread of f(P) relative to size but minimal hysteresis. Got better with calibration runtime
| GDF016_newchip_220405 | 2200 | 5 | 0.9 | yes | 2.8 | 120 | Acceptable spread of f(P) but noticeable hysteresis
GDF 017 | GDF017_2022_02_18 | 2400 | 9 | 1 | yes | 1.5 | 140 | Large f(P) but with acceptable spread. Moderate k2 pressure dependence and small hysteresis
| GDF017_TiCoverplate_2022_03_22 | 2200 | 11 | 2.5 | no | 0 | 40 | Larger f(P) and spread but no hysteresis and minimal k2 pressure dependence
| GDF017_newchip_041522 | 2200 | 6 | 0.8 | yes | 2 | 130 | Smaller f(P) and spread but moderate hysteresis and k2 pressure dependence
GDF 018 | GDF018_2022_02_23 | 2400 | 5 | 5 | no | 0 | 190 | Erratic change in Vrs during test so f(P) was unusable. Failed
| GDF018_2022_02_28 | 2400 | 0 | 1.5 | yes | 2 | 70 | Odd but somewhat repeatable f(P), but large spread relative to size. Small k2 pressure dependence
| GDF018_newchip_220405 | 2200 | 7 | 1 | yes | 3 | 130 | Large(ish) but aceptable f(P) and spread with moderate hysteresis and k2 pressure dependence. No high TP helps the f(P) spread
GDF 019 | GDF019_2022_02_23 | 2400 | 8 | 2 | yes | 2 | 60 | Large(ish) f(P) and spread with moderate hysteresis but minimal k2 pressure dependence
| GDF019_2022_02_28 | 2400 | 5 | 5 | yes | 3 | 110 | Sensor had odd decreasing pressure failure that resulted in erratic f(P)
| GDF019_oring_2022_03_08 | 2200 | 4 | 1 | yes | 3 | 90 | Smaller f(P) but hysteresis remains. The f(P) is more noisy relative to size
| GDF019_new_chip_2022_03_22 | 2200 | 6 | 0.7 | yes | 2 | 120 | Larger f(P) and k2 pressure correction but smaller hysteresis, less of a spread

## subsetTestFiles
Rows with identifier in them | Selected Folder | Selected File Name | Look for identifer: | 1 | Starting at Column | TestLog!$M$1 | Test Files
6 | GDF004 | GDF4_102620
7 | GDF005 | GDF5_032221
8 | GDF006 | GDF6_initialCal
10 | GDF007 | GDF7_initialCal
#N/A
#N/A
#N/A
#N/A
#N/A
#N/A
#N/A
#N/A
#N/A

## subsetPrunellaFiles
|  |  | Look for identifer: | 1 | Starting at Column | TestLog!$J$1 | Prunella Parts
Rows with identifier in them | Selected Folder | Selected File Name
11 | GDF007 | GDF7_4hr_Ptcycle
12 | GDF007 | GDF7_k2_20_1000dBar
14 | GDF007 | GDF7_variable_Prate_20_10_2_take2
#N/A | #N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A
#N/A | #N/A

## subsetAEFiles
|  |  | Look for identifer: | 1 | Starting at Column | TestLog!$N$1 | AE Parts
Rows with identifier in them | Selected Folder | Selected File Name
248 | GDF2024 | GDF2024_101223_Ch B
249 | GDF2025 | GDF2025_10192023_Ch A
250 | GDF2023 | GDF2023_10192023_Ch B
251 | GDF2024 | GDF2024_10192023_Ch E
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A
#N/A | #N/A | #N/A