## data
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6 |  | Raw Data | [6 empty] | Channel: | B | [19 empty] | Drift | Drift Offset | Chemical |  |  |  |  | Physical | APEX (Up ONLY) |  | Pressure (from Start to Stop)
|  | A: DF #311 and VL #311 B: DF #313 and VL #313 | [28 empty] | 0 | Vrs = Vc + k0 + k2T + f(P), where Vc = Vct + Vcp | [7 empty] | P6= | #VALUE!
[32 empty] | Vct = - (R*T/F)*ln(α[H]*α[Cl]) |  |  |  |  | k2 = | After: | Before: | P5= | #VALUE!
[32 empty] | Vcp = - ΔVo*(1/23061)*0.0242*P |  |  |  |  | -0.00096427 | 2020-10-26 12:50 | 2020-10-27 16:45 | P4= | #VALUE!
[30 empty] | Divisor | Drift Corr (mV/d) |  |  |  |  | Calc: |  |  |  | P3= | #VALUE!
|  |  |  | 1025.3385 | [5 empty] | AVG | [14 empty] | 0.00011248 |  | 0.00016381 | -1.8634e-09 | 1.1468e-09 | 500 | 1 | R = | 8.3143 | J/(mol*K) |  |  | -1.4344 | AVG | -1.4372 | P2= | #VALUE! | #VALUE! | #VALUE!
Soln: 0.01N HCl |  |  |  | 833.4251 | [5 empty] | StDev | [14 empty] | 3.1427e-05 |  | 4.4221e-05 | 5.2168e-09 | 7.9278e-09 | Min | -1.2063 | F = | 96487 | coulomb/mol |  |  | 0.0062857 | StDev | 0.0052317 | P1= | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:50 |  |  |  | 0 | [5 empty] | Min | 0 | 20 | [12 empty] | 3.8e-05 |  | 4.6e-05 | -1.963e-08 | -2.455e-08 | Max | -1.2267 | [5 empty] | -1.4414 | Min | -1.4413 |  | #VALUE! | #VALUE! | #VALUE!
2020-10-27 16:45 |  |  |  | 2089.7924 | [5 empty] | Max | 0 | 20 | [12 empty] | 0.000289 |  | 0.000358 | 1.876e-08 | 2.428e-08 |  |  |  |  | Vct | Vcp |  | -1.4248 | Max | -1.4248 |  |  | #VALUE! | #VALUE!
Labview Date/Time | Thermistor Temp ( C ) | Bath Temp | Bath Set Point | Pressure (dBar) | Max Flow | Flow rate | Pump Vol | CRC | ID | MSC Date/Time | UTC SECS | CTD DEPTH | CTD TEMP | CTD SALINITY | SAMPLE COUNTER | RESET COUNTER | ERROR COUNTER | HOUSING TEMPERATURE | HOUSING HUMIDITY | INPUT VOLTAGE | INPUT CURRENT | FOOBAR PH VALUE | BACKUP BATTERY | VRS (V) | StDev Vrs (V) | Vk (V) | StDev Vk (V) | Ik (V) | Ib (V) | ISCO P (dBar/1k) | Vrs' (V) | Mean γ HCl | ΔVo | -RT/F * ln (α[H]*α[Cl]) | -ΔV0/ F (P–1) | Vrs-Vc (V) | Vrs-Vc-k2T (V) | Vrs-Vc-k2T (V) Up_Only |  | Vrs-Vc-k2T-f(P) (V) [=k0] | Vrs-Vc-k2T-f(P) (V) Up Only | Vrs-Vc-k2T-f(P) (V) | Vrs-Vc-k2T-f(P) (V) Up Only
2020-10-26 12:50 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 110.731 | 0x7642 | pH | 2021-03-09 00:50 | 0 | -1 | -1 | -1 | 1205593 | 180 | 341 | 25.76 | 26.33 | 11.696 | 0.0436 | 7.1234 | 3.616 | -1.2064 | 0.000154 | -1.4362 | 0.000205 | -1.678e-09 | 1.173e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #N/A | #N/A
2020-10-26 12:50 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.002 | 110.7291 | 0x69F6 | pH | 2021-03-09 00:51 | 0 | -1 | -1 | -1 | 1205594 | 180 | 341 | 25.8 | 26.28 | 11.708 | 0.0436 | 7.1234 | 3.616 | -1.2064 | 0.000129 | -1.4349 | 0.000144 | 3.575e-09 | -7.785e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | -1.4248 | -1.4248 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 110.7274 | 0x5017 | pH | 2021-03-09 00:52 | 0 | -1 | -1 | -1 | 1205595 | 180 | 341 | 25.85 | 26.31 | 11.712 | 0.0436 | 7.1234 | 3.616 | -1.2064 | 0.000194 | -1.4336 | 0.000191 | -3.159e-09 | 4.868e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.003 | 110.7256 | 0xA569 | pH | 2021-03-09 00:53 | 0 | -1 | -1 | -1 | 1205596 | 180 | 341 | 25.89 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2063 | 0.000196 | -1.4326 | 0.00025 | -8.166e-09 | 1.405e-08 | 0.0002759 | -1.2063 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 110.724 | 0xF1A7 | pH | 2021-03-09 00:54 | 0 | -1 | -1 | -1 | 1205597 | 180 | 341 | 25.96 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2064 | 0.00012 | -1.4316 | 0.000243 | -3.865e-09 | 2.119e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.7223 | 0x7ED3 | pH | 2021-03-09 00:55 | 0 | -1 | -1 | -1 | 1205598 | 180 | 341 | 26.02 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2064 | 0.000226 | -1.4306 | 0.000316 | 5.318e-09 | 9.499e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 110.7206 | 0x5EA8 | pH | 2021-03-09 00:56 | 0 | -1 | -1 | -1 | 1205599 | 180 | 341 | 26.06 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2064 | 0.000191 | -1.4294 | 0.000235 | 1.113e-09 | -3.631e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.002 | 110.7191 | 0x2BFB | pH | 2021-03-09 00:57 | 0 | -1 | -1 | -1 | 1205600 | 180 | 341 | 26.11 | 25.93 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2064 | 0.000149 | -1.4278 | 0.000183 | 3.767e-09 | -1.228e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.7176 | 0xC729 | pH | 2021-03-09 00:58 | 0 | -1 | -1 | -1 | 1205601 | 180 | 341 | 26.16 | 25.91 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2064 | 0.000203 | -1.4271 | 0.000227 | -5.208e-10 | -3.958e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.7161 | 0xDDE7 | pH | 2021-03-09 00:59 | 0 | -1 | -1 | -1 | 1205602 | 180 | 341 | 26.21 | 25.86 | 11.716 | 0.0434 | 7.1234 | 3.617 | -1.2064 | 0.00021 | -1.4259 | 0.000249 | 4.3e-09 | -8.065e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.239 | 110.5349 | 0x6115 | pH | 2021-03-09 01:00 | 0 | -1 | -1 | -1 | 1205603 | 180 | 341 | 26.25 | 25.88 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000194 | -1.4354 | 0.000319 | -8.066e-09 | 4.418e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:51 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.238 | 110.2921 | 0x40B5 | pH | 2021-03-09 01:01 | 0 | -1 | -1 | -1 | 1205604 | 180 | 341 | 26.29 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000214 | -1.4419 | 0.000287 | -4.613e-09 | 1.45e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.236 | 110.0508 | 0xAD86 | pH | 2021-03-09 01:02 | 0 | -1 | -1 | -1 | 1205605 | 180 | 341 | 26.31 | 25.75 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000133 | -1.4458 | 0.00034 | -2.469e-09 | 1.534e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.236 | 109.8104 | 0x9DA0 | pH | 2021-03-09 01:03 | 0 | -1 | -1 | -1 | 1205606 | 180 | 341 | 26.35 | 25.75 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000179 | -1.4483 | 0.000232 | -3.03e-09 | 1.074e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.007 | 109.6131 | 0xDD42 | pH | 2021-03-09 01:04 | 0 | -1 | -1 | -1 | 1205607 | 180 | 341 | 26.38 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000179 | -1.4486 | 0.000157 | 5.146e-09 | 1.321e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.002 | 109.6096 | 0x473C | pH | 2021-03-09 01:05 | 0 | -1 | -1 | -1 | 1205608 | 180 | 341 | 26.41 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000217 | -1.4457 | 0.000198 | 5.717e-09 | -3.081e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.002 | 109.6073 | 0xF1B3 | pH | 2021-03-09 01:06 | 0 | -1 | -1 | -1 | 1205609 | 180 | 341 | 26.46 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000158 | -1.443 | 0.000201 | -9.844e-09 | -9.169e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 109.6054 | 0x3F75 | pH | 2021-03-09 01:07 | 0 | -1 | -1 | -1 | 1205610 | 180 | 341 | 26.47 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.00016 | -1.4407 | 0.000281 | -4.466e-09 | 1.013e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.002 | 109.6036 | 0x6268 | pH | 2021-03-09 01:08 | 0 | -1 | -1 | -1 | 1205611 | 180 | 341 | 26.51 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000152 | -1.4387 | 0.000309 | -1.073e-08 | -8.973e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 109.6021 | 0x309C | pH | 2021-03-09 01:09 | 0 | -1 | -1 | -1 | 1205612 | 180 | 341 | 26.53 | 25.61 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000232 | -1.4369 | 0.000294 | 5.29e-09 | -5.377e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.778 | 0xF276 | pH | 2021-03-09 01:10 | 0 | -1 | -1 | -1 | 1205613 | 180 | 341 | 26.55 | 25.61 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000289 | -1.4317 | 0.000275 | -9.701e-10 | 2.263e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:52 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.236 | 110.0179 | 0x6503 | pH | 2021-03-09 01:11 | 0 | -1 | -1 | -1 | 1205614 | 180 | 341 | 26.58 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000184 | -1.4276 | 0.000259 | -2.786e-09 | 4.644e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.237 | 110.2589 | 0x6E0D | pH | 2021-03-09 01:12 | 0 | -1 | -1 | -1 | 1205615 | 180 | 341 | 26.59 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000194 | -1.4239 | 0.000173 | 4.228e-09 | -8.994e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.239 | 110.5016 | 0x4514 | pH | 2021-03-09 01:13 | 0 | -1 | -1 | -1 | 1205616 | 180 | 341 | 26.62 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 0.000234 | -1.4191 | 0.000114 | -1.255e-08 | 9.685e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.004 | 110.6669 | 0x248E | pH | 2021-03-09 01:14 | 0 | -1 | -1 | -1 | 1205617 | 180 | 341 | 26.63 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000258 | -1.4157 | 0.000214 | 1.009e-09 | -2.675e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.001 | 110.6694 | 0xE38D | pH | 2021-03-09 01:16 | 0 | -1 | -1 | -1 | 1205618 | 180 | 341 | 26.65 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.618 | -1.2064 | 9.6e-05 | -1.4151 | 8.9e-05 | -5.424e-09 | -3.729e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.6698 | 0xE68B | pH | 2021-03-09 01:17 | 0 | -1 | -1 | -1 | 1205619 | 180 | 341 | 26.66 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.6e-05 | -1.4142 | 0.000202 | -4.552e-09 | 5.547e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.67 | 0xA57E | pH | 2021-03-09 01:18 | 0 | -1 | -1 | -1 | 1205620 | 180 | 341 | 26.66 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.8e-05 | -1.4135 | 0.000131 | 1.356e-08 | 3.559e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.6702 | 0xDA67 | pH | 2021-03-09 01:19 | 0 | -1 | -1 | -1 | 1205621 | 180 | 341 | 26.68 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.4e-05 | -1.4128 | 0.000199 | -6.88e-09 | -2.908e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.6702 | 0xE92D | pH | 2021-03-09 01:20 | 0 | -1 | -1 | -1 | 1205622 | 180 | 341 | 26.68 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 6.7e-05 | -1.4118 | 0.000122 | -1.097e-08 | 1.911e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.239 | 110.4901 | 0x52D6 | pH | 2021-03-09 01:21 | 0 | -1 | -1 | -1 | 1205623 | 180 | 341 | 26.68 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.4e-05 | -1.4214 | 0.000253 | 3.264e-09 | -8.242e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:53 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.237 | 110.2476 | 0x1CD7 | pH | 2021-03-09 01:22 | 0 | -1 | -1 | -1 | 1205624 | 180 | 341 | 26.69 | 25.5 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 6.8e-05 | -1.4289 | 0.000113 | 4.806e-09 | 1.157e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 110.0064 | 0xAD0F | pH | 2021-03-09 01:23 | 0 | -1 | -1 | -1 | 1205625 | 180 | 341 | 26.69 | 25.5 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.5e-05 | -1.434 | 0.000172 | -1.67e-08 | 4.643e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.236 | 109.7661 | 0x9BE8 | pH | 2021-03-09 01:24 | 0 | -1 | -1 | -1 | 1205626 | 180 | 341 | 26.69 | 25.5 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 5.7e-05 | -1.4373 | 0.000107 | -4.104e-09 | -1.448e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.5262 | 0x1B16 | pH | 2021-03-09 01:25 | 0 | -1 | -1 | -1 | 1205627 | 180 | 341 | 26.69 | 25.5 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000109 | -1.4401 | 0.000238 | -1.185e-08 | 1.192e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.235 | 109.2866 | 0x5DD9 | pH | 2021-03-09 01:26 | 0 | -1 | -1 | -1 | 1205628 | 180 | 341 | 26.69 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000118 | -1.4424 | 0.000229 | 1.694e-09 | 9.061e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.0471 | 0xBF5C | pH | 2021-03-09 01:27 | 0 | -1 | -1 | -1 | 1205629 | 180 | 341 | 26.69 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000121 | -1.4436 | 0.000116 | -1.89e-09 | -1.118e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.235 | 108.8077 | 0x7081 | pH | 2021-03-09 01:28 | 0 | -1 | -1 | -1 | 1205630 | 180 | 341 | 26.69 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000115 | -1.4447 | 0.000188 | -2.65e-09 | -1.673e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.007 | 108.6362 | 0xE368 | pH | 2021-03-09 01:29 | 0 | -1 | -1 | -1 | 1205631 | 180 | 341 | 26.69 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000134 | -1.4435 | 0.000142 | -1.274e-09 | 1.9e-08 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.002 | 108.6327 | 0x5993 | pH | 2021-03-09 01:30 | 0 | -1 | -1 | -1 | 1205632 | 180 | 341 | 26.68 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.7e-05 | -1.4405 | 0.000154 | 5.477e-09 | -4.432e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 108.6305 | 0x11B6 | pH | 2021-03-09 01:31 | 0 | -1 | -1 | -1 | 1205633 | 180 | 341 | 26.69 | 25.53 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000158 | -1.4379 | 4.6e-05 | 1.262e-09 | -5.592e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 108.6286 | 0xDC4C | pH | 2021-03-09 01:32 | 0 | -1 | -1 | -1 | 1205634 | 180 | 341 | 26.68 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000129 | -1.4333 | 0.000243 | 5.981e-10 | -2.281e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:54 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 108.6271 | 0x12B6 | pH | 2021-03-09 01:33 | 0 | -1 | -1 | -1 | 1205635 | 180 | 341 | 26.68 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000124 | -1.4339 | 0.000144 | -2.895e-09 | 1.223e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 108.6259 | 0x5EC8 | pH | 2021-03-09 01:34 | 0 | -1 | -1 | -1 | 1205636 | 180 | 341 | 26.68 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000156 | -1.4325 | 0.000252 | -3.973e-09 | 1.582e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 108.6246 | 0x4477 | pH | 2021-03-09 01:35 | 0 | -1 | -1 | -1 | 1205637 | 180 | 341 | 26.66 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.9e-05 | -1.4309 | 0.000176 | -4.644e-09 | -9.196e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 108.6235 | 0xDE9D | pH | 2021-03-09 01:36 | 0 | -1 | -1 | -1 | 1205638 | 180 | 341 | 26.66 | 25.61 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000136 | -1.4295 | 0.000175 | 4.142e-10 | 2.668e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 108.6227 | 0xA7D3 | pH | 2021-03-09 01:37 | 0 | -1 | -1 | -1 | 1205639 | 180 | 341 | 26.66 | 25.61 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000101 | -1.4281 | 0.000168 | 2.011e-09 | 1.615e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.13795 | 0.25 | 0.001 | 108.6217 | 0x2AC0 | pH | 2021-03-09 01:38 | 0 | -1 | -1 | -1 | 1205640 | 180 | 341 | 26.66 | 25.63 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.8e-05 | -1.4268 | 0.000171 | -2.965e-09 | 2.355e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 108.6209 | 0x504E | pH | 2021-03-09 01:39 | 0 | -1 | -1 | -1 | 1205641 | 180 | 341 | 26.66 | 25.61 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.9e-05 | -1.4256 | 0.000152 | 3.563e-09 | -2.868e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 108.6202 | 0x764C | pH | 2021-03-09 01:40 | 0 | -1 | -1 | -1 | 1205642 | 180 | 341 | 26.68 | 25.63 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.9e-05 | -1.4243 | 0.000119 | 2.139e-09 | -1.338e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 108.7956 | 0x3C31 | pH | 2021-03-09 01:41 | 0 | -1 | -1 | -1 | 1205643 | 180 | 341 | 26.68 | 25.63 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000127 | -1.4213 | 0.000124 | -1.392e-09 | -5.656e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.235 | 109.0342 | 0x0FFF | pH | 2021-03-09 01:42 | 0 | -1 | -1 | -1 | 1205644 | 180 | 341 | 26.66 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.9e-05 | -1.4189 | 0.000154 | 4.728e-10 | -5.692e-10 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:55 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.2733 | 0xA050 | pH | 2021-03-09 01:43 | 0 | -1 | -1 | -1 | 1205645 | 180 | 341 | 26.68 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 5.9e-05 | -1.4175 | 0.000161 | -1.097e-08 | -1.011e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.235 | 109.5128 | 0x403B | pH | 2021-03-09 01:44 | 0 | -1 | -1 | -1 | 1205646 | 180 | 341 | 26.66 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000107 | -1.417 | 0.000127 | 2.672e-09 | -2.838e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.236 | 109.7526 | 0x34F4 | pH | 2021-03-09 01:45 | 0 | -1 | -1 | -1 | 1205647 | 180 | 341 | 26.68 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000127 | -1.4156 | 0.000151 | -4.25e-09 | 7.204e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.9932 | 0x2E57 | pH | 2021-03-09 01:46 | 0 | -1 | -1 | -1 | 1205648 | 180 | 341 | 26.69 | 25.63 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000102 | -1.4132 | 0.0002 | 2.297e-09 | 1.411e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.13795 | 0.25 | -0.237 | 110.2347 | 0x5E69 | pH | 2021-03-09 01:47 | 0 | -1 | -1 | -1 | 1205649 | 180 | 341 | 26.7 | 25.61 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000109 | -1.4096 | 0.000106 | -4.068e-09 | -6.525e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.239 | 110.4775 | 0x4FAE | pH | 2021-03-09 01:48 | 0 | -1 | -1 | -1 | 1205650 | 180 | 341 | 26.72 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000122 | -1.4057 | 0.000249 | -1.403e-09 | 8.284e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.008 | 110.6211 | 0x387E | pH | 2021-03-09 01:49 | 0 | -1 | -1 | -1 | 1205651 | 180 | 341 | 26.72 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.00014 | -1.4033 | 0.000264 | -1.067e-08 | -8.621e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.002 | 110.6251 | 0x24CC | pH | 2021-03-09 01:50 | 0 | -1 | -1 | -1 | 1205652 | 180 | 341 | 26.72 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000116 | -1.4026 | 0.000219 | -5.176e-09 | 5.34e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.003 | 110.6276 | 0xEDD3 | pH | 2021-03-09 01:51 | 0 | -1 | -1 | -1 | 1205653 | 180 | 341 | 26.72 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000164 | -1.4018 | 0.000197 | -3.693e-09 | -5.562e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.629 | 0xEA96 | pH | 2021-03-09 01:52 | 0 | -1 | -1 | -1 | 1205654 | 180 | 341 | 26.73 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000168 | -1.4011 | 0.000127 | -7.816e-09 | 4.515e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.6306 | 0xCF3A | pH | 2021-03-09 01:53 | 0 | -1 | -1 | -1 | 1205655 | 180 | 341 | 26.74 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000167 | -1.4004 | 0.000233 | 6.003e-09 | 2.584e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:56 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.6317 | 0xA53F | pH | 2021-03-09 01:54 | 0 | -1 | -1 | -1 | 1205656 | 180 | 341 | 26.74 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000111 | -1.3998 | 0.000272 | 2.422e-09 | -1.203e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.6326 | 0x10D8 | pH | 2021-03-09 01:55 | 0 | -1 | -1 | -1 | 1205657 | 180 | 341 | 26.76 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000144 | -1.3991 | 0.000236 | -5.014e-09 | -4.938e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.6335 | 0xB612 | pH | 2021-03-09 01:56 | 0 | -1 | -1 | -1 | 1205658 | 180 | 341 | 26.77 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000139 | -1.3983 | 0.000251 | 5.835e-10 | -5.068e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.13795 | 0.25 | 0 | 110.6342 | 0x0287 | pH | 2021-03-09 01:57 | 0 | -1 | -1 | -1 | 1205659 | 180 | 341 | 26.77 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.9e-05 | -1.3978 | 0.000128 | -7.468e-09 | 4.219e-09 | 0.0002759 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.8102e-05 | -1.4442 | -1.4249 | -1.4249 | -1.4249 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.635 | 0x147E | pH | 2021-03-09 01:58 | 0 | -1 | -1 | -1 | 1205660 | 180 | 341 | 26.8 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000174 | -1.397 | 0.000225 | -5.359e-09 | 1.797e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.6356 | 0xDCB5 | pH | 2021-03-09 01:59 | 0 | -1 | -1 | -1 | 1205661 | 180 | 341 | 26.8 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000139 | -1.3961 | 0.000195 | -3.986e-09 | 1.178e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.239 | 110.4571 | 0x5A83 | pH | 2021-03-09 02:00 | 0 | -1 | -1 | -1 | 1205662 | 180 | 341 | 26.81 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.4e-05 | -1.4055 | 0.000181 | -4.578e-09 | 2.224e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.237 | 110.2149 | 0x50B1 | pH | 2021-03-09 02:01 | 0 | -1 | -1 | -1 | 1205663 | 180 | 341 | 26.82 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000159 | -1.4139 | 0.000268 | 7.758e-09 | -2.382e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.9739 | 0x44BC | pH | 2021-03-09 02:02 | 0 | -1 | -1 | -1 | 1205664 | 180 | 341 | 26.82 | 25.71 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000143 | -1.4199 | 0.000193 | 4.736e-10 | 1.18e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.7336 | 0xB075 | pH | 2021-03-09 02:03 | 0 | -1 | -1 | -1 | 1205665 | 180 | 341 | 26.85 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000191 | -1.4241 | 0.000358 | -1.463e-09 | 1.404e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:57 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.4938 | 0x9871 | pH | 2021-03-09 02:04 | 0 | -1 | -1 | -1 | 1205666 | 180 | 341 | 26.86 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000113 | -1.4277 | 0.000203 | 1.485e-09 | 9.705e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.2541 | 0x50CE | pH | 2021-03-09 02:05 | 0 | -1 | -1 | -1 | 1205667 | 180 | 341 | 26.85 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000111 | -1.4302 | 0.00014 | -1.863e-09 | 1.2e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.0146 | 0x18BD | pH | 2021-03-09 02:06 | 0 | -1 | -1 | -1 | 1205668 | 180 | 341 | 26.86 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 7.5e-05 | -1.432 | 0.000152 | -5.842e-09 | -1.922e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.7754 | 0xD6A6 | pH | 2021-03-09 02:07 | 0 | -1 | -1 | -1 | 1205669 | 180 | 341 | 26.88 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.2e-05 | -1.4334 | 0.000194 | -2.266e-09 | -1.111e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.5363 | 0x207F | pH | 2021-03-09 02:08 | 0 | -1 | -1 | -1 | 1205670 | 180 | 341 | 26.88 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000171 | -1.4345 | 0.000123 | -1.88e-10 | -3.256e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.2971 | 0x1DBC | pH | 2021-03-09 02:09 | 0 | -1 | -1 | -1 | 1205671 | 180 | 341 | 26.89 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.3e-05 | -1.4354 | 0.000161 | -5.885e-09 | 5.428e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.0582 | 0x8141 | pH | 2021-03-09 02:10 | 0 | -1 | -1 | -1 | 1205672 | 180 | 341 | 26.89 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000111 | -1.4448 | 0.000157 | 2.463e-09 | 1.918e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.234 | 107.8179 | 0x477C | pH | 2021-03-09 02:11 | 0 | -1 | -1 | -1 | 1205673 | 180 | 341 | 26.92 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 7.9e-05 | -1.4444 | 0.000197 | 4.993e-09 | -6.968e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.005 | 107.6833 | 0x58D4 | pH | 2021-03-09 02:12 | 0 | -1 | -1 | -1 | 1205674 | 180 | 341 | 26.92 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000108 | -1.4425 | 0.000113 | 1.464e-09 | 1.47e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.6805 | 0xBBBD | pH | 2021-03-09 02:13 | 0 | -1 | -1 | -1 | 1205675 | 180 | 341 | 26.92 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.2e-05 | -1.4399 | 0.000221 | 3.951e-09 | 9.301e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:58 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6786 | 0x92CD | pH | 2021-03-09 02:14 | 0 | -1 | -1 | -1 | 1205676 | 180 | 341 | 26.92 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000131 | -1.4377 | 0.000135 | 4.434e-09 | -2.293e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.677 | 0xC41B | pH | 2021-03-09 02:16 | 0 | -1 | -1 | -1 | 1205677 | 180 | 341 | 26.93 | 25.71 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 8e-05 | -1.4351 | 0.000134 | -4.917e-09 | -7.104e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.6757 | 0xFF02 | pH | 2021-03-09 02:17 | 0 | -1 | -1 | -1 | 1205678 | 180 | 341 | 26.93 | 25.71 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000128 | -1.4339 | 0.000243 | 9.57e-09 | 1.541e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6745 | 0xA6C0 | pH | 2021-03-09 02:18 | 0 | -1 | -1 | -1 | 1205679 | 180 | 341 | 26.93 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.2e-05 | -1.4324 | 0.000155 | -1.484e-09 | 5.678e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.6733 | 0x1466 | pH | 2021-03-09 02:19 | 0 | -1 | -1 | -1 | 1205680 | 180 | 341 | 26.93 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000142 | -1.4308 | 0.000123 | -1.778e-09 | 2.767e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6724 | 0x3BB7 | pH | 2021-03-09 02:20 | 0 | -1 | -1 | -1 | 1205681 | 180 | 341 | 26.92 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000112 | -1.4295 | 0.000159 | 1.54e-09 | 1.131e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.6713 | 0x19E7 | pH | 2021-03-09 02:21 | 0 | -1 | -1 | -1 | 1205682 | 180 | 341 | 26.91 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000148 | -1.4282 | 0.00014 | -5.672e-10 | -9.663e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6704 | 0xF372 | pH | 2021-03-09 02:22 | 0 | -1 | -1 | -1 | 1205683 | 180 | 341 | 26.91 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 8.5e-05 | -1.4272 | 0.000204 | 3.678e-10 | 8.051e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.6695 | 0x8331 | pH | 2021-03-09 02:23 | 0 | -1 | -1 | -1 | 1205684 | 180 | 341 | 26.91 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.00014 | -1.4259 | 0.000156 | -7.297e-09 | 1.069e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6686 | 0xF8C4 | pH | 2021-03-09 02:24 | 0 | -1 | -1 | -1 | 1205685 | 180 | 341 | 26.91 | 25.73 | 11.72 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000106 | -1.4247 | 0.000271 | -6.147e-09 | 3.195e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.6676 | 0x1808 | pH | 2021-03-09 02:25 | 0 | -1 | -1 | -1 | 1205686 | 180 | 341 | 26.91 | 25.75 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000117 | -1.4237 | 0.000123 | -8.336e-09 | 6.489e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 12:59 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6669 | 0x7F27 | pH | 2021-03-09 02:26 | 0 | -1 | -1 | -1 | 1205687 | 180 | 341 | 26.91 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.00012 | -1.4227 | 0.000125 | -5.837e-09 | -5.518e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6661 | 0xA546 | pH | 2021-03-09 02:27 | 0 | -1 | -1 | -1 | 1205688 | 180 | 341 | 26.89 | 25.75 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 6.1e-05 | -1.4218 | 0.000146 | -1.288e-08 | 2.112e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6654 | 0x7016 | pH | 2021-03-09 02:28 | 0 | -1 | -1 | -1 | 1205689 | 180 | 341 | 26.89 | 25.75 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.0001 | -1.4205 | 0.000188 | -7.091e-09 | -5.953e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6649 | 0xC832 | pH | 2021-03-09 02:29 | 0 | -1 | -1 | -1 | 1205690 | 180 | 341 | 26.89 | 25.78 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 7.6e-05 | -1.4198 | 0.000147 | -3.847e-09 | 4.324e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.6641 | 0x2F63 | pH | 2021-03-09 02:30 | 0 | -1 | -1 | -1 | 1205691 | 180 | 341 | 26.89 | 25.78 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000115 | -1.4189 | 0.00019 | -1.441e-09 | -9.104e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 107.8394 | 0xE9B8 | pH | 2021-03-09 02:31 | 0 | -1 | -1 | -1 | 1205692 | 180 | 341 | 26.88 | 25.78 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000111 | -1.4162 | 9.8e-05 | -5.632e-10 | -4.3e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 108.0776 | 0xA89E | pH | 2021-03-09 02:32 | 0 | -1 | -1 | -1 | 1205693 | 180 | 341 | 26.86 | 25.78 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000131 | -1.4141 | 0.000162 | 7.936e-09 | 5.592e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 108.316 | 0x3B43 | pH | 2021-03-09 02:33 | 0 | -1 | -1 | -1 | 1205694 | 180 | 341 | 26.86 | 25.78 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000125 | -1.4122 | 0.000185 | -2.283e-09 | 4.975e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 108.5548 | 0xD5B1 | pH | 2021-03-09 02:34 | 0 | -1 | -1 | -1 | 1205695 | 180 | 341 | 26.85 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000125 | -1.4107 | 0.000137 | -4.114e-09 | 7.897e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 108.7938 | 0x0CBC | pH | 2021-03-09 02:35 | 0 | -1 | -1 | -1 | 1205696 | 180 | 341 | 26.85 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000103 | -1.4105 | 9.1e-05 | 1.448e-09 | -1.167e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:00 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.0331 | 0x31D8 | pH | 2021-03-09 02:36 | 0 | -1 | -1 | -1 | 1205697 | 180 | 341 | 26.85 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 9e-05 | -1.4117 | 0.000123 | -1.198e-08 | 3.304e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 109.2729 | 0x7D9E | pH | 2021-03-09 02:37 | 0 | -1 | -1 | -1 | 1205698 | 180 | 341 | 26.85 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.00017 | -1.4128 | 0.000133 | 4.711e-09 | 1.165e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.5127 | 0x0658 | pH | 2021-03-09 02:38 | 0 | -1 | -1 | -1 | 1205699 | 180 | 341 | 26.85 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000113 | -1.4124 | 0.000144 | -1.622e-09 | -3.932e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.237 | 109.7528 | 0xA0F1 | pH | 2021-03-09 02:39 | 0 | -1 | -1 | -1 | 1205700 | 180 | 341 | 26.85 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000107 | -1.4098 | 0.00013 | -9.976e-09 | 1.765e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 109.9935 | 0x27E4 | pH | 2021-03-09 02:40 | 0 | -1 | -1 | -1 | 1205701 | 180 | 341 | 26.85 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 8.3e-05 | -1.4061 | 0.000186 | -7.122e-09 | -8.265e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.238 | 110.2351 | 0x6DE2 | pH | 2021-03-09 02:41 | 0 | -1 | -1 | -1 | 1205702 | 180 | 341 | 26.85 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000115 | -1.4022 | 0.000184 | 1.108e-09 | 3.148e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.238 | 110.4781 | 0x1F05 | pH | 2021-03-09 02:42 | 0 | -1 | -1 | -1 | 1205703 | 180 | 341 | 26.85 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000141 | -1.3985 | 0.000148 | -4.814e-09 | 3.191e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.005 | 110.5799 | 0xA24D | pH | 2021-03-09 02:43 | 0 | -1 | -1 | -1 | 1205704 | 180 | 341 | 26.85 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000122 | -1.3958 | 0.000181 | 6.877e-10 | 5.164e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.003 | 110.5842 | 0x0843 | pH | 2021-03-09 02:44 | 0 | -1 | -1 | -1 | 1205705 | 180 | 341 | 26.85 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000167 | -1.3938 | 9.4e-05 | 3.182e-09 | 3.162e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.002 | 110.5869 | 0xF205 | pH | 2021-03-09 02:45 | 0 | -1 | -1 | -1 | 1205706 | 180 | 341 | 26.86 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 6e-05 | -1.3922 | 0.000124 | 1.594e-09 | 1.064e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:01 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.002 | 110.589 | 0xCACB | pH | 2021-03-09 02:46 | 0 | -1 | -1 | -1 | 1205707 | 180 | 341 | 26.88 | 25.81 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.00012 | -1.3908 | 0.000109 | -2.648e-09 | 3.089e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.002 | 110.5908 | 0x45F8 | pH | 2021-03-09 02:47 | 0 | -1 | -1 | -1 | 1205708 | 180 | 341 | 26.88 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000121 | -1.3894 | 9.9e-05 | 7.845e-10 | -3.444e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5924 | 0x970D | pH | 2021-03-09 02:48 | 0 | -1 | -1 | -1 | 1205709 | 180 | 341 | 26.89 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000128 | -1.3882 | 0.000204 | -6.983e-09 | 9.005e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.002 | 110.5936 | 0x7E71 | pH | 2021-03-09 02:49 | 0 | -1 | -1 | -1 | 1205710 | 180 | 341 | 26.91 | 25.83 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.3e-05 | -1.387 | 0.000142 | 6.226e-10 | 7.639e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5948 | 0x1C77 | pH | 2021-03-09 02:50 | 0 | -1 | -1 | -1 | 1205711 | 180 | 341 | 26.89 | 25.86 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000129 | -1.3859 | 0.000161 | -4.403e-09 | -1.133e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5961 | 0xD5F3 | pH | 2021-03-09 02:51 | 0 | -1 | -1 | -1 | 1205712 | 180 | 341 | 26.91 | 25.88 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 9.9e-05 | -1.3847 | 0.000188 | -1.441e-08 | -6.593e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5971 | 0x0CE8 | pH | 2021-03-09 02:52 | 0 | -1 | -1 | -1 | 1205713 | 180 | 341 | 26.91 | 25.91 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000161 | -1.3838 | 0.000141 | -5.641e-09 | 7.811e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5981 | 0x71A1 | pH | 2021-03-09 02:53 | 0 | -1 | -1 | -1 | 1205714 | 180 | 341 | 26.92 | 25.93 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000117 | -1.3829 | 0.000107 | -3.182e-09 | -1.03e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5991 | 0xD0B5 | pH | 2021-03-09 02:54 | 0 | -1 | -1 | -1 | 1205715 | 180 | 341 | 26.93 | 25.96 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000128 | -1.3821 | 0.000205 | -1.611e-10 | -2.142e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5999 | 0x1B5B | pH | 2021-03-09 02:55 | 0 | -1 | -1 | -1 | 1205716 | 180 | 341 | 26.95 | 25.93 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000161 | -1.3812 | 0.000159 | -2.075e-10 | -3.046e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.6009 | 0x3CC5 | pH | 2021-03-09 02:56 | 0 | -1 | -1 | -1 | 1205717 | 180 | 341 | 26.96 | 25.93 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.3e-05 | -1.3806 | 0.000195 | -6.095e-10 | -3.554e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:02 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.6018 | 0x9F04 | pH | 2021-03-09 02:57 | 0 | -1 | -1 | -1 | 1205718 | 180 | 341 | 26.96 | 25.93 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 6.6e-05 | -1.38 | 0.000134 | -3.285e-09 | -1.559e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.6024 | 0x14FE | pH | 2021-03-09 02:58 | 0 | -1 | -1 | -1 | 1205719 | 180 | 341 | 26.97 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 8.9e-05 | -1.3793 | 0.000137 | -6.022e-11 | 8.538e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.6033 | 0xDA36 | pH | 2021-03-09 02:59 | 0 | -1 | -1 | -1 | 1205720 | 180 | 341 | 26.97 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7.5e-05 | -1.3788 | 0.000218 | -1.042e-08 | 9.167e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.239 | 110.424 | 0x599E | pH | 2021-03-09 03:00 | 0 | -1 | -1 | -1 | 1205721 | 180 | 341 | 26.99 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 9e-05 | -1.3875 | 0.00022 | 4.462e-09 | 1.001e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.237 | 110.182 | 0xB24A | pH | 2021-03-09 03:01 | 0 | -1 | -1 | -1 | 1205722 | 180 | 341 | 26.99 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 9.6e-05 | -1.3987 | 6.4e-05 | -8.843e-09 | 1.998e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.9403 | 0xB301 | pH | 2021-03-09 03:02 | 0 | -1 | -1 | -1 | 1205723 | 180 | 341 | 27 | 26.01 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.000116 | -1.406 | 0.000109 | 7.642e-10 | -2.141e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.7006 | 0x5146 | pH | 2021-03-09 03:03 | 0 | -1 | -1 | -1 | 1205724 | 180 | 341 | 27.01 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 9.1e-05 | -1.411 | 7.8e-05 | 7.39e-09 | 9.011e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.4608 | 0xEA74 | pH | 2021-03-09 03:04 | 0 | -1 | -1 | -1 | 1205725 | 180 | 341 | 27.03 | 26.03 | 11.72 | 0.0434 | 7.1234 | 3.619 | -1.2065 | 0.000139 | -1.415 | 0.000148 | -1.679e-09 | 3.082e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.221 | 0xB7CC | pH | 2021-03-09 03:05 | 0 | -1 | -1 | -1 | 1205726 | 180 | 341 | 27.03 | 26.03 | 11.72 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 6.9e-05 | -1.4192 | 0.000105 | 1.981e-09 | 1.129e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.9816 | 0x6B9F | pH | 2021-03-09 03:06 | 0 | -1 | -1 | -1 | 1205727 | 180 | 341 | 27.04 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000117 | -1.4235 | 0.00014 | -2.165e-10 | -2.811e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:03 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.7424 | 0xDB96 | pH | 2021-03-09 03:07 | 0 | -1 | -1 | -1 | 1205728 | 180 | 341 | 27.05 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.619 | -1.2064 | 0.0001 | -1.4265 | 0.00014 | -2.663e-09 | 2.091e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.5031 | 0xC36A | pH | 2021-03-09 03:08 | 0 | -1 | -1 | -1 | 1205729 | 180 | 341 | 27.07 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000106 | -1.4289 | 0.000102 | -3.68e-09 | -1.712e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.264 | 0x2A32 | pH | 2021-03-09 03:09 | 0 | -1 | -1 | -1 | 1205730 | 180 | 341 | 27.05 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000153 | -1.4301 | 0.000208 | -6.003e-09 | 3.464e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.234 | 108.025 | 0xE774 | pH | 2021-03-09 03:10 | 0 | -1 | -1 | -1 | 1205731 | 180 | 341 | 27.05 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000149 | -1.4305 | 0.000177 | -5.507e-09 | -6.469e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 107.7859 | 0x41DD | pH | 2021-03-09 03:11 | 0 | -1 | -1 | -1 | 1205732 | 180 | 341 | 27.07 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000134 | -1.4307 | 0.000197 | -1.225e-08 | 3.857e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.234 | 107.547 | 0x462C | pH | 2021-03-09 03:12 | 0 | -1 | -1 | -1 | 1205733 | 180 | 341 | 27.07 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000118 | -1.431 | 0.000123 | 2.279e-11 | 2.377e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.234 | 107.3081 | 0xDD44 | pH | 2021-03-09 03:13 | 0 | -1 | -1 | -1 | 1205734 | 180 | 341 | 27.07 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000117 | -1.4313 | 0.00018 | -1.2e-08 | -4.654e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.004 | 107.2193 | 0x5519 | pH | 2021-03-09 03:14 | 0 | -1 | -1 | -1 | 1205735 | 180 | 341 | 27.05 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000102 | -1.4297 | 0.000154 | 5.162e-09 | -7.622e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.002 | 107.2165 | 0xB7A9 | pH | 2021-03-09 03:16 | 0 | -1 | -1 | -1 | 1205736 | 180 | 341 | 27.04 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000105 | -1.4279 | 0.000105 | -8.078e-09 | 4.449e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2144 | 0x1203 | pH | 2021-03-09 03:17 | 0 | -1 | -1 | -1 | 1205737 | 180 | 341 | 27.05 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000107 | -1.426 | 0.000157 | -6.978e-09 | -5.715e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2129 | 0x59A7 | pH | 2021-03-09 03:18 | 0 | -1 | -1 | -1 | 1205738 | 180 | 341 | 27.03 | 26.18 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8e-05 | -1.4246 | 0.00019 | -3.277e-09 | 7.897e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:04 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.004 | 107.2114 | 0xB125 | pH | 2021-03-09 03:19 | 0 | -1 | -1 | -1 | 1205739 | 180 | 341 | 27.03 | 26.18 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000107 | -1.4231 | 0.00021 | 2.644e-09 | 2.379e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2104 | 0xFE79 | pH | 2021-03-09 03:20 | 0 | -1 | -1 | -1 | 1205740 | 180 | 341 | 27.03 | 26.16 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7e-05 | -1.4219 | 0.000214 | -1.762e-08 | 5.73e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.209 | 0x43F7 | pH | 2021-03-09 03:21 | 0 | -1 | -1 | -1 | 1205741 | 180 | 341 | 27.01 | 26.18 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000113 | -1.4206 | 0.00015 | -8.464e-11 | -2.408e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.208 | 0xE355 | pH | 2021-03-09 03:22 | 0 | -1 | -1 | -1 | 1205742 | 180 | 341 | 27 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000114 | -1.4194 | 0.00021 | 1.232e-08 | 2.299e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2071 | 0x4364 | pH | 2021-03-09 03:23 | 0 | -1 | -1 | -1 | 1205743 | 180 | 341 | 27 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000121 | -1.4183 | 0.00016 | -7.323e-12 | -4.813e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2062 | 0x95BC | pH | 2021-03-09 03:24 | 0 | -1 | -1 | -1 | 1205744 | 180 | 341 | 26.99 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000159 | -1.4171 | 0.000163 | 7.303e-09 | 1.015e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2053 | 0xF5AF | pH | 2021-03-09 03:25 | 0 | -1 | -1 | -1 | 1205745 | 180 | 341 | 26.97 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000138 | -1.4162 | 0.000119 | -5.302e-09 | -7.521e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2046 | 0x17F5 | pH | 2021-03-09 03:26 | 0 | -1 | -1 | -1 | 1205746 | 180 | 341 | 26.97 | 26.26 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000111 | -1.4152 | 0.000149 | 3.62e-09 | -7.701e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2039 | 0x87B1 | pH | 2021-03-09 03:27 | 0 | -1 | -1 | -1 | 1205747 | 180 | 341 | 26.97 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000145 | -1.4143 | 0.000145 | -7.267e-10 | -2.785e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2033 | 0x2EC9 | pH | 2021-03-09 03:28 | 0 | -1 | -1 | -1 | 1205748 | 180 | 341 | 26.96 | 26.26 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000113 | -1.4135 | 0.000148 | -6.031e-09 | -2.973e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:05 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2029 | 0xE589 | pH | 2021-03-09 03:29 | 0 | -1 | -1 | -1 | 1205749 | 180 | 341 | 26.96 | 26.28 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000131 | -1.4124 | 0.000184 | 2.886e-09 | 1.546e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2025 | 0x89B1 | pH | 2021-03-09 03:30 | 0 | -1 | -1 | -1 | 1205750 | 180 | 341 | 26.95 | 26.26 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00011 | -1.4116 | 0.00016 | 3.581e-10 | -1.356e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2021 | 0x0B20 | pH | 2021-03-09 03:31 | 0 | -1 | -1 | -1 | 1205751 | 180 | 341 | 26.93 | 26.26 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.8e-05 | -1.4108 | 0.000275 | -3.764e-09 | -5.482e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2017 | 0x8E8B | pH | 2021-03-09 03:32 | 0 | -1 | -1 | -1 | 1205752 | 180 | 341 | 26.93 | 26.26 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000133 | -1.4099 | 0.000115 | -8.525e-09 | -9.606e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2015 | 0xF7D1 | pH | 2021-03-09 03:33 | 0 | -1 | -1 | -1 | 1205753 | 180 | 341 | 26.93 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000101 | -1.4091 | 0.000131 | -6.234e-10 | -5.854e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2012 | 0x729C | pH | 2021-03-09 03:34 | 0 | -1 | -1 | -1 | 1205754 | 180 | 341 | 26.93 | 26.24 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000116 | -1.4083 | 0.000155 | -9.788e-09 | -9.741e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2012 | 0x5329 | pH | 2021-03-09 03:35 | 0 | -1 | -1 | -1 | 1205755 | 180 | 341 | 26.92 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000107 | -1.4076 | 0.000114 | 1.027e-09 | 1.782e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2009 | 0x30C3 | pH | 2021-03-09 03:36 | 0 | -1 | -1 | -1 | 1205756 | 180 | 341 | 26.93 | 26.26 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9.8e-05 | -1.4069 | 0.000142 | 4.75e-09 | -1.151e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2007 | 0x2743 | pH | 2021-03-09 03:37 | 0 | -1 | -1 | -1 | 1205757 | 180 | 341 | 26.93 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000118 | -1.4062 | 0.000196 | -1.121e-09 | -7.76e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2006 | 0x308C | pH | 2021-03-09 03:38 | 0 | -1 | -1 | -1 | 1205758 | 180 | 341 | 26.93 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000111 | -1.4054 | 7e-05 | 5.54e-09 | 1.348e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:06 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2006 | 0x2460 | pH | 2021-03-09 03:39 | 0 | -1 | -1 | -1 | 1205759 | 180 | 341 | 26.95 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000164 | -1.4049 | 9.6e-05 | -4.522e-09 | 7.844e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2006 | 0x4750 | pH | 2021-03-09 03:40 | 0 | -1 | -1 | -1 | 1205760 | 180 | 341 | 26.93 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000123 | -1.4043 | 0.000186 | -8.604e-09 | -4.236e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.233 | 107.3755 | 0xA2A4 | pH | 2021-03-09 03:41 | 0 | -1 | -1 | -1 | 1205761 | 180 | 341 | 26.95 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000118 | -1.4021 | 0.000129 | -1.772e-09 | -7.669e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 107.6137 | 0x5DBD | pH | 2021-03-09 03:42 | 0 | -1 | -1 | -1 | 1205762 | 180 | 341 | 26.93 | 26.24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9.6e-05 | -1.4001 | 0.00014 | -7.976e-09 | -1.362e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 107.8518 | 0x06DA | pH | 2021-03-09 03:43 | 0 | -1 | -1 | -1 | 1205763 | 180 | 341 | 26.95 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000153 | -1.3983 | 0.000159 | -5.149e-09 | 6.396e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 108.0903 | 0x3B3A | pH | 2021-03-09 03:44 | 0 | -1 | -1 | -1 | 1205764 | 180 | 341 | 26.96 | 26.24 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000152 | -1.397 | 0.000161 | 4.432e-09 | 6.977e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.3291 | 0x3F64 | pH | 2021-03-09 03:45 | 0 | -1 | -1 | -1 | 1205765 | 180 | 341 | 26.96 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.9e-05 | -1.396 | 0.000226 | 8.946e-09 | 1.613e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.5682 | 0x4BA8 | pH | 2021-03-09 03:46 | 0 | -1 | -1 | -1 | 1205766 | 180 | 341 | 26.97 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000149 | -1.3958 | 0.000125 | -5.952e-09 | 9.827e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.8075 | 0x7D57 | pH | 2021-03-09 03:47 | 0 | -1 | -1 | -1 | 1205767 | 180 | 341 | 26.97 | 26.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000107 | -1.3963 | 0.000243 | -1.015e-08 | 1.359e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.0468 | 0x1ADD | pH | 2021-03-09 03:48 | 0 | -1 | -1 | -1 | 1205768 | 180 | 341 | 26.99 | 26.18 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000116 | -1.3971 | 8e-05 | -1.986e-09 | 1.289e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:07 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 109.2866 | 0xA4BE | pH | 2021-03-09 03:49 | 0 | -1 | -1 | -1 | 1205769 | 180 | 341 | 26.99 | 26.18 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7.7e-05 | -1.3975 | 0.000149 | -3.91e-09 | -6.206e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 109.5265 | 0x6D01 | pH | 2021-03-09 03:50 | 0 | -1 | -1 | -1 | 1205770 | 180 | 341 | 27 | 26.16 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.2e-05 | -1.3966 | 0.000138 | -7.376e-09 | -3.09e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.7667 | 0x5EA5 | pH | 2021-03-09 03:51 | 0 | -1 | -1 | -1 | 1205771 | 180 | 341 | 27 | 26.16 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9.4e-05 | -1.394 | 0.000211 | 4.623e-09 | 1.6e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 110.0075 | 0x9EB2 | pH | 2021-03-09 03:52 | 0 | -1 | -1 | -1 | 1205772 | 180 | 341 | 27.01 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000167 | -1.3907 | 0.000148 | -5.381e-09 | 9.908e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.239 | 110.2495 | 0x4A4A | pH | 2021-03-09 03:53 | 0 | -1 | -1 | -1 | 1205773 | 180 | 341 | 27.03 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000106 | -1.387 | 0.000107 | -5.825e-09 | -3.161e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.239 | 110.4926 | 0x11DE | pH | 2021-03-09 03:54 | 0 | -1 | -1 | -1 | 1205774 | 180 | 341 | 27.04 | 26.13 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 6.6e-05 | -1.3832 | 8.1e-05 | 1.676e-09 | -9.058e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.003 | 110.5666 | 0xA824 | pH | 2021-03-09 03:55 | 0 | -1 | -1 | -1 | 1205775 | 180 | 341 | 27.04 | 26.11 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000121 | -1.3805 | 0.000168 | 9.79e-10 | -3.319e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.003 | 110.57 | 0x0552 | pH | 2021-03-09 03:56 | 0 | -1 | -1 | -1 | 1205776 | 180 | 341 | 27.05 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 5.9e-05 | -1.3789 | 0.000161 | -4.859e-09 | 1.226e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5723 | 0xF8D1 | pH | 2021-03-09 03:57 | 0 | -1 | -1 | -1 | 1205777 | 180 | 341 | 27.07 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000155 | -1.3774 | 0.000138 | -5.39e-09 | -1.066e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4443 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5742 | 0x0374 | pH | 2021-03-09 03:58 | 0 | -1 | -1 | -1 | 1205778 | 180 | 341 | 27.07 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000123 | -1.3762 | 0.0002 | -6.112e-10 | 5.108e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5755 | 0x2463 | pH | 2021-03-09 03:59 | 0 | -1 | -1 | -1 | 1205779 | 180 | 341 | 27.08 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000107 | -1.375 | 0.000156 | -8.868e-09 | 6.465e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:08 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5766 | 0xE9CF | pH | 2021-03-09 04:00 | 0 | -1 | -1 | -1 | 1205780 | 180 | 341 | 27.08 | 26.06 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 9.3e-05 | -1.3742 | 0.000135 | -1.839e-10 | 1.57e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5776 | 0x7763 | pH | 2021-03-09 04:01 | 0 | -1 | -1 | -1 | 1205781 | 180 | 341 | 27.11 | 26.03 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000145 | -1.3732 | 0.00014 | 1.807e-09 | 2.075e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5784 | 0xB653 | pH | 2021-03-09 04:02 | 0 | -1 | -1 | -1 | 1205782 | 180 | 341 | 27.11 | 26.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8e-05 | -1.3725 | 0.000204 | -3.414e-09 | -1.048e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5793 | 0x668A | pH | 2021-03-09 04:03 | 0 | -1 | -1 | -1 | 1205783 | 180 | 341 | 27.12 | 26.01 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000136 | -1.3716 | 9.5e-05 | -7.213e-09 | 8.61e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.58 | 0x6539 | pH | 2021-03-09 04:04 | 0 | -1 | -1 | -1 | 1205784 | 180 | 341 | 27.12 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.1e-05 | -1.3709 | 0.000137 | -6.128e-10 | 9.779e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5807 | 0x6685 | pH | 2021-03-09 04:05 | 0 | -1 | -1 | -1 | 1205785 | 180 | 341 | 27.14 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000131 | -1.3704 | 0.000217 | -1.749e-09 | 9.966e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5813 | 0xE00A | pH | 2021-03-09 04:06 | 0 | -1 | -1 | -1 | 1205786 | 180 | 341 | 27.15 | 25.96 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000113 | -1.3699 | 0.000151 | -6.384e-09 | -1.5e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5817 | 0xDD7C | pH | 2021-03-09 04:07 | 0 | -1 | -1 | -1 | 1205787 | 180 | 341 | 27.15 | 25.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000121 | -1.3692 | 0.000179 | -6.994e-09 | -5.261e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5821 | 0x0E07 | pH | 2021-03-09 04:08 | 0 | -1 | -1 | -1 | 1205788 | 180 | 341 | 27.15 | 25.93 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 6.3e-05 | -1.3689 | 0.000131 | 2.349e-09 | 6.026e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5824 | 0x8792 | pH | 2021-03-09 04:09 | 0 | -1 | -1 | -1 | 1205789 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000126 | -1.3684 | 9.6e-05 | -2.043e-09 | -4.118e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:09 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5829 | 0xEA1A | pH | 2021-03-09 04:10 | 0 | -1 | -1 | -1 | 1205790 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.4e-05 | -1.3679 | 0.000128 | -6.636e-09 | 3.529e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5829 | 0xB492 | pH | 2021-03-09 04:11 | 0 | -1 | -1 | -1 | 1205791 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 9.3e-05 | -1.3676 | 0.000182 | 3.004e-09 | 1.15e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5832 | 0x2C90 | pH | 2021-03-09 04:12 | 0 | -1 | -1 | -1 | 1205792 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 5.3e-05 | -1.3673 | 0.000219 | 2.04e-09 | -2.266e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5835 | 0x2537 | pH | 2021-03-09 04:13 | 0 | -1 | -1 | -1 | 1205793 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000116 | -1.3669 | 0.000157 | -3.673e-09 | 4.056e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5835 | 0xE78C | pH | 2021-03-09 04:14 | 0 | -1 | -1 | -1 | 1205794 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 7e-05 | -1.3665 | 0.000146 | -1.527e-09 | -4.767e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5839 | 0xFD06 | pH | 2021-03-09 04:15 | 0 | -1 | -1 | -1 | 1205795 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000138 | -1.3662 | 0.00015 | -4.484e-10 | 2.592e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.584 | 0xC3CF | pH | 2021-03-09 04:16 | 0 | -1 | -1 | -1 | 1205796 | 180 | 341 | 27.18 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000102 | -1.3658 | 0.000134 | 8.077e-09 | 6.804e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5841 | 0x85A6 | pH | 2021-03-09 04:18 | 0 | -1 | -1 | -1 | 1205797 | 180 | 341 | 27.16 | 25.93 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 6.8e-05 | -1.3656 | 0.00013 | 4.045e-10 | 5.606e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5841 | 0x9346 | pH | 2021-03-09 04:19 | 0 | -1 | -1 | -1 | 1205798 | 180 | 341 | 27.15 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9.4e-05 | -1.3654 | 0.000158 | -6.317e-09 | -1.18e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5842 | 0x4DAD | pH | 2021-03-09 04:20 | 0 | -1 | -1 | -1 | 1205799 | 180 | 341 | 27.15 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000112 | -1.3651 | 0.000186 | 5.112e-09 | 8.008e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.238 | 110.4052 | 0xF74E | pH | 2021-03-09 04:21 | 0 | -1 | -1 | -1 | 1205800 | 180 | 341 | 27.15 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000109 | -1.3739 | 0.00016 | -5.959e-09 | 1.695e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:10 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.237 | 110.1632 | 0x1DBE | pH | 2021-03-09 04:22 | 0 | -1 | -1 | -1 | 1205801 | 180 | 341 | 27.14 | 25.93 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 6.9e-05 | -1.385 | 0.00021 | 6.706e-10 | 6.523e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.9222 | 0x1096 | pH | 2021-03-09 04:23 | 0 | -1 | -1 | -1 | 1205802 | 180 | 341 | 27.14 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 6.7e-05 | -1.3924 | 0.000164 | 2.968e-09 | -1.96e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.682 | 0xEC87 | pH | 2021-03-09 04:24 | 0 | -1 | -1 | -1 | 1205803 | 180 | 341 | 27.14 | 25.96 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000118 | -1.3977 | 0.000177 | 2.729e-09 | 9.847e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 109.4422 | 0xC2A9 | pH | 2021-03-09 04:25 | 0 | -1 | -1 | -1 | 1205804 | 180 | 341 | 27.12 | 25.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000104 | -1.4017 | 0.000124 | -4.666e-09 | 4.667e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.236 | 109.2025 | 0xFC10 | pH | 2021-03-09 04:26 | 0 | -1 | -1 | -1 | 1205805 | 180 | 341 | 27.11 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 6.2e-05 | -1.4063 | 0.00017 | -5.219e-09 | 8.051e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.9629 | 0x11F8 | pH | 2021-03-09 04:27 | 0 | -1 | -1 | -1 | 1205806 | 180 | 341 | 27.11 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000125 | -1.4107 | 0.0001 | 4.256e-10 | 1.305e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.7236 | 0x1CFF | pH | 2021-03-09 04:28 | 0 | -1 | -1 | -1 | 1205807 | 180 | 341 | 27.11 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.3e-05 | -1.4137 | 0.00018 | -1.166e-08 | -1.907e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.4844 | 0x50EE | pH | 2021-03-09 04:29 | 0 | -1 | -1 | -1 | 1205808 | 180 | 341 | 27.09 | 25.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00011 | -1.4163 | 0.000144 | -1.106e-08 | -1.086e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 108.2454 | 0x5A7D | pH | 2021-03-09 04:30 | 0 | -1 | -1 | -1 | 1205809 | 180 | 341 | 27.08 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7.2e-05 | -1.418 | 0.000224 | -4.956e-09 | -4.968e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.234 | 108.0063 | 0x5183 | pH | 2021-03-09 04:31 | 0 | -1 | -1 | -1 | 1205810 | 180 | 341 | 27.08 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000129 | -1.4186 | 0.000134 | -7.708e-09 | 4.721e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:11 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 107.767 | 0x095A | pH | 2021-03-09 04:32 | 0 | -1 | -1 | -1 | 1205811 | 180 | 341 | 27.07 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000126 | -1.4188 | 0.000154 | -4.82e-09 | -1.109e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.235 | 107.5279 | 0x4733 | pH | 2021-03-09 04:33 | 0 | -1 | -1 | -1 | 1205812 | 180 | 341 | 27.07 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9.1e-05 | -1.419 | 0.000181 | -3.484e-09 | 1.055e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.234 | 107.2898 | 0x74C7 | pH | 2021-03-09 04:34 | 0 | -1 | -1 | -1 | 1205813 | 180 | 341 | 27.05 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 7.4e-05 | -1.4189 | 0.000196 | -4.065e-09 | 5.021e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.002 | 107.2117 | 0x3F08 | pH | 2021-03-09 04:35 | 0 | -1 | -1 | -1 | 1205814 | 180 | 341 | 27.05 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000136 | -1.4162 | 0.000144 | 4.539e-09 | -1.685e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.002 | 107.2092 | 0x1F0C | pH | 2021-03-09 04:36 | 0 | -1 | -1 | -1 | 1205815 | 180 | 341 | 27.04 | 26.06 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 6.5e-05 | -1.4154 | 0.000139 | 3.67e-10 | 2.445e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2078 | 0x55BE | pH | 2021-03-09 04:37 | 0 | -1 | -1 | -1 | 1205816 | 180 | 341 | 27.04 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000114 | -1.4142 | 0.000152 | -6.803e-10 | 1.822e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2063 | 0x7867 | pH | 2021-03-09 04:38 | 0 | -1 | -1 | -1 | 1205817 | 180 | 341 | 27.03 | 26.06 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000113 | -1.4128 | 0.000152 | 5.263e-09 | 1.099e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2054 | 0x4FC4 | pH | 2021-03-09 04:39 | 0 | -1 | -1 | -1 | 1205818 | 180 | 341 | 27.04 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000102 | -1.4115 | 0.000156 | -1.541e-08 | 3.031e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2046 | 0xF8E5 | pH | 2021-03-09 04:40 | 0 | -1 | -1 | -1 | 1205819 | 180 | 341 | 27.03 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.1e-05 | -1.4105 | 0.000118 | 1.007e-09 | -6.632e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.204 | 0xB050 | pH | 2021-03-09 04:41 | 0 | -1 | -1 | -1 | 1205820 | 180 | 341 | 27.03 | 26.08 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000105 | -1.4091 | 0.000115 | 1.362e-09 | 1.059e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:12 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2031 | 0x83B7 | pH | 2021-03-09 04:42 | 0 | -1 | -1 | -1 | 1205821 | 180 | 341 | 27.01 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.5e-05 | -1.408 | 0.000159 | 6.008e-09 | 5.179e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.003 | 107.2025 | 0xECE0 | pH | 2021-03-09 04:43 | 0 | -1 | -1 | -1 | 1205822 | 180 | 341 | 27.03 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000144 | -1.407 | 0.000176 | 1.017e-08 | 8.879e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2022 | 0x2563 | pH | 2021-03-09 04:44 | 0 | -1 | -1 | -1 | 1205823 | 180 | 341 | 27.03 | 26.08 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000104 | -1.4061 | 0.000156 | -5.243e-09 | -3.711e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2017 | 0xCB97 | pH | 2021-03-09 04:45 | 0 | -1 | -1 | -1 | 1205824 | 180 | 341 | 27.03 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8.9e-05 | -1.405 | 0.000178 | -3.849e-10 | -1.589e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2014 | 0x1540 | pH | 2021-03-09 04:46 | 0 | -1 | -1 | -1 | 1205825 | 180 | 341 | 27.01 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8.6e-05 | -1.4041 | 0.000154 | -1.986e-10 | -2.146e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2011 | 0xABB3 | pH | 2021-03-09 04:47 | 0 | -1 | -1 | -1 | 1205826 | 180 | 341 | 27.03 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 5.3e-05 | -1.4032 | 0.000135 | 4.085e-10 | 2.438e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2008 | 0x1F9C | pH | 2021-03-09 04:48 | 0 | -1 | -1 | -1 | 1205827 | 180 | 341 | 27.03 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000127 | -1.4026 | 0.00025 | 9.294e-10 | 6.165e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.2003 | 0xDD68 | pH | 2021-03-09 04:49 | 0 | -1 | -1 | -1 | 1205828 | 180 | 341 | 27.03 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8.7e-05 | -1.4016 | 0.000184 | -6.414e-09 | 6.786e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.2002 | 0x8D7F | pH | 2021-03-09 04:50 | 0 | -1 | -1 | -1 | 1205829 | 180 | 341 | 27.03 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000113 | -1.4009 | 0.000185 | -3.505e-09 | 2.211e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1999 | 0xA6F6 | pH | 2021-03-09 04:51 | 0 | -1 | -1 | -1 | 1205830 | 180 | 341 | 27.04 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000124 | -1.4 | 0.000154 | -1.01e-08 | -4.365e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1998 | 0x9D70 | pH | 2021-03-09 04:52 | 0 | -1 | -1 | -1 | 1205831 | 180 | 341 | 27.03 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000123 | -1.3992 | 0.000173 | -9.424e-10 | 3.603e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:13 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1994 | 0x1D80 | pH | 2021-03-09 04:53 | 0 | -1 | -1 | -1 | 1205832 | 180 | 341 | 27.03 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000162 | -1.3986 | 0.000155 | -3.982e-09 | -4.153e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1993 | 0xBF2E | pH | 2021-03-09 04:54 | 0 | -1 | -1 | -1 | 1205833 | 180 | 341 | 27.03 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9e-05 | -1.3978 | 0.000109 | -3.381e-09 | 1.606e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0.001 | 107.1991 | 0xBD82 | pH | 2021-03-09 04:55 | 0 | -1 | -1 | -1 | 1205834 | 180 | 341 | 27.04 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000127 | -1.3971 | 7.2e-05 | -3.677e-09 | 9.351e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1988 | 0x2B36 | pH | 2021-03-09 04:56 | 0 | -1 | -1 | -1 | 1205835 | 180 | 341 | 27.04 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000143 | -1.3965 | 0.000133 | 4.682e-09 | 3.911e-11 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1988 | 0xE62C | pH | 2021-03-09 04:57 | 0 | -1 | -1 | -1 | 1205836 | 180 | 341 | 27.05 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 4.9e-05 | -1.3959 | 9.3e-05 | -2.205e-09 | -3.438e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1985 | 0xE3B3 | pH | 2021-03-09 04:58 | 0 | -1 | -1 | -1 | 1205837 | 180 | 341 | 27.05 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000155 | -1.3952 | 0.000171 | 1.432e-10 | -1.322e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1985 | 0xCCAC | pH | 2021-03-09 04:59 | 0 | -1 | -1 | -1 | 1205838 | 180 | 341 | 27.04 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000141 | -1.3946 | 0.000147 | 1.876e-08 | 8.122e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1982 | 0x2863 | pH | 2021-03-09 05:00 | 0 | -1 | -1 | -1 | 1205839 | 180 | 341 | 27.05 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000124 | -1.3939 | 0.00012 | -3.93e-09 | 4.389e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1982 | 0xCD0B | pH | 2021-03-09 05:01 | 0 | -1 | -1 | -1 | 1205840 | 180 | 341 | 27.07 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.6e-05 | -1.3935 | 0.000139 | -2.873e-09 | 1.237e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1981 | 0xCBCC | pH | 2021-03-09 05:02 | 0 | -1 | -1 | -1 | 1205841 | 180 | 341 | 27.07 | 26.03 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000166 | -1.3927 | 0.000185 | 3.8e-09 | 6.673e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:14 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1979 | 0x2EBE | pH | 2021-03-09 05:03 | 0 | -1 | -1 | -1 | 1205842 | 180 | 341 | 27.08 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.8e-05 | -1.3923 | 9.2e-05 | 1.94e-09 | 4.682e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 107.1979 | 0x5E47 | pH | 2021-03-09 05:04 | 0 | -1 | -1 | -1 | 1205843 | 180 | 341 | 27.07 | 26.01 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 6.3e-05 | -1.3917 | 0.000194 | 1.497e-09 | 3.853e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 107.3731 | 0xCF8B | pH | 2021-03-09 05:05 | 0 | -1 | -1 | -1 | 1205844 | 180 | 341 | 27.08 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.4e-05 | -1.3899 | 0.000184 | 7.254e-09 | 1.758e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 107.6112 | 0x2B00 | pH | 2021-03-09 05:06 | 0 | -1 | -1 | -1 | 1205845 | 180 | 341 | 27.09 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000124 | -1.3884 | 9.9e-05 | -1.523e-09 | -1.17e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.234 | 107.8496 | 0x4C60 | pH | 2021-03-09 05:07 | 0 | -1 | -1 | -1 | 1205846 | 180 | 341 | 27.09 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 7.7e-05 | -1.3869 | 0.000148 | -2.603e-09 | 3.799e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.0882 | 0xD861 | pH | 2021-03-09 05:08 | 0 | -1 | -1 | -1 | 1205847 | 180 | 341 | 27.09 | 25.96 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000125 | -1.3857 | 9.7e-05 | -7.098e-09 | 1.171e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.327 | 0x51CD | pH | 2021-03-09 05:09 | 0 | -1 | -1 | -1 | 1205848 | 180 | 341 | 27.11 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.4e-05 | -1.3849 | 0.000186 | -5.301e-09 | 4.736e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.566 | 0x980D | pH | 2021-03-09 05:10 | 0 | -1 | -1 | -1 | 1205849 | 180 | 341 | 27.11 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000112 | -1.3849 | 0.00019 | 2.01e-09 | -5.795e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 108.8052 | 0x52A5 | pH | 2021-03-09 05:11 | 0 | -1 | -1 | -1 | 1205850 | 180 | 341 | 27.12 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00012 | -1.3851 | 0.000133 | -5.61e-09 | 1.27e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.235 | 109.0445 | 0x3809 | pH | 2021-03-09 05:12 | 0 | -1 | -1 | -1 | 1205851 | 180 | 341 | 27.12 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.00014 | -1.3856 | 0.000202 | -6.971e-09 | -9.631e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:15 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 109.2842 | 0x3F1E | pH | 2021-03-09 05:13 | 0 | -1 | -1 | -1 | 1205852 | 180 | 341 | 27.12 | 25.98 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 7.9e-05 | -1.3856 | 0.000138 | 2.455e-09 | -4.937e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 109.5242 | 0x61E1 | pH | 2021-03-09 05:14 | 0 | -1 | -1 | -1 | 1205853 | 180 | 341 | 27.14 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.4e-05 | -1.3844 | 0.000155 | 9.572e-09 | 9.467e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.237 | 109.7645 | 0xBBDF | pH | 2021-03-09 05:15 | 0 | -1 | -1 | -1 | 1205854 | 180 | 341 | 27.14 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000136 | -1.3817 | 0.000135 | 5.299e-09 | 1.318e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.236 | 110.0053 | 0xE997 | pH | 2021-03-09 05:16 | 0 | -1 | -1 | -1 | 1205855 | 180 | 341 | 27.14 | 25.98 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000125 | -1.3786 | 0.00013 | -2.236e-09 | -6.545e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.238 | 110.247 | 0x51D4 | pH | 2021-03-09 05:18 | 0 | -1 | -1 | -1 | 1205856 | 180 | 341 | 27.15 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000101 | -1.3748 | 0.000186 | 5.941e-11 | 1.614e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.239 | 110.4902 | 0x4CE1 | pH | 2021-03-09 05:19 | 0 | -1 | -1 | -1 | 1205857 | 180 | 341 | 27.15 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000129 | -1.371 | 0.000113 | 4.342e-09 | -1.158e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.005 | 110.5629 | 0xFF39 | pH | 2021-03-09 05:20 | 0 | -1 | -1 | -1 | 1205858 | 180 | 341 | 27.16 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000163 | -1.3684 | 0.000121 | -8.927e-10 | -4.747e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.002 | 110.5661 | 0xAA65 | pH | 2021-03-09 05:21 | 0 | -1 | -1 | -1 | 1205859 | 180 | 341 | 27.16 | 25.98 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000132 | -1.3667 | 0.000173 | -2.479e-09 | -1.285e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.568 | 0x16C0 | pH | 2021-03-09 05:22 | 0 | -1 | -1 | -1 | 1205860 | 180 | 341 | 27.16 | 25.98 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.0001 | -1.3655 | 0.000235 | -2.831e-09 | 5.397e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4443 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5697 | 0xC1C3 | pH | 2021-03-09 05:23 | 0 | -1 | -1 | -1 | 1205861 | 180 | 341 | 27.16 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.1e-05 | -1.3645 | 0.000164 | -4.557e-09 | -8e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.571 | 0xCBAF | pH | 2021-03-09 05:24 | 0 | -1 | -1 | -1 | 1205862 | 180 | 341 | 27.18 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 7.5e-05 | -1.3636 | 0.000209 | 1.273e-09 | -4.968e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:16 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5722 | 0x5ABC | pH | 2021-03-09 05:25 | 0 | -1 | -1 | -1 | 1205863 | 180 | 341 | 27.18 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000123 | -1.3628 | 0.000162 | -7.267e-10 | -6.518e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5731 | 0xDCFF | pH | 2021-03-09 05:26 | 0 | -1 | -1 | -1 | 1205864 | 180 | 341 | 27.19 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000114 | -1.3621 | 0.000212 | -1.139e-09 | 9.066e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5739 | 0x70D3 | pH | 2021-03-09 05:27 | 0 | -1 | -1 | -1 | 1205865 | 180 | 341 | 27.18 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000133 | -1.3616 | 0.000179 | -2.248e-09 | -1.877e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | -0.001 | 110.5747 | 0xAAE1 | pH | 2021-03-09 05:28 | 0 | -1 | -1 | -1 | 1205866 | 180 | 341 | 27.18 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000128 | -1.361 | 0.000169 | -6.383e-09 | 3.479e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5754 | 0x5D74 | pH | 2021-03-09 05:29 | 0 | -1 | -1 | -1 | 1205867 | 180 | 341 | 27.19 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000105 | -1.3605 | 0.000211 | -1.646e-09 | -4.202e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5761 | 0x2ABB | pH | 2021-03-09 05:30 | 0 | -1 | -1 | -1 | 1205868 | 180 | 341 | 27.2 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000113 | -1.36 | 0.00018 | 1.975e-09 | 6.922e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5767 | 0x0641 | pH | 2021-03-09 05:31 | 0 | -1 | -1 | -1 | 1205869 | 180 | 341 | 27.19 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000144 | -1.3596 | 0.000153 | 5.02e-09 | 4.069e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5773 | 0x07A9 | pH | 2021-03-09 05:32 | 0 | -1 | -1 | -1 | 1205870 | 180 | 341 | 27.2 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7.2e-05 | -1.3593 | 0.0001 | 2.1e-10 | 1.259e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5778 | 0x0DC8 | pH | 2021-03-09 05:33 | 0 | -1 | -1 | -1 | 1205871 | 180 | 341 | 27.2 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000117 | -1.3589 | 0.000116 | 3.026e-09 | 1.873e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5781 | 0xB536 | pH | 2021-03-09 05:34 | 0 | -1 | -1 | -1 | 1205872 | 180 | 341 | 27.2 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00012 | -1.3586 | 0.000166 | 3.862e-09 | 6.806e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:17 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5784 | 0x0CA3 | pH | 2021-03-09 05:35 | 0 | -1 | -1 | -1 | 1205873 | 180 | 341 | 27.22 | 26.01 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 7.4e-05 | -1.3582 | 0.000126 | 5.133e-09 | -4.202e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5788 | 0x8C97 | pH | 2021-03-09 05:36 | 0 | -1 | -1 | -1 | 1205874 | 180 | 341 | 27.23 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000102 | -1.358 | 0.000105 | -4.571e-09 | 7.947e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5792 | 0x84F7 | pH | 2021-03-09 05:37 | 0 | -1 | -1 | -1 | 1205875 | 180 | 341 | 27.22 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000146 | -1.3577 | 0.000192 | -1.294e-09 | 2.299e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5795 | 0x9B60 | pH | 2021-03-09 05:38 | 0 | -1 | -1 | -1 | 1205876 | 180 | 341 | 27.23 | 26.01 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000163 | -1.3574 | 0.000113 | 1.135e-09 | 8.945e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5797 | 0x1A7C | pH | 2021-03-09 05:39 | 0 | -1 | -1 | -1 | 1205877 | 180 | 341 | 27.22 | 26.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000137 | -1.3572 | 0.00016 | -4.606e-10 | -6.57e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5799 | 0x47A2 | pH | 2021-03-09 05:40 | 0 | -1 | -1 | -1 | 1205878 | 180 | 341 | 27.23 | 26.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2063 | 6.8e-05 | -1.3569 | 0.000181 | -1.302e-08 | -3.23e-09 | 0.00041385 | -1.2063 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5802 | 0xD426 | pH | 2021-03-09 05:41 | 0 | -1 | -1 | -1 | 1205879 | 180 | 341 | 27.22 | 26.03 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 7.5e-05 | -1.3568 | 8.5e-05 | -4.591e-09 | 6.379e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0x4471 | pH | 2021-03-09 05:42 | 0 | -1 | -1 | -1 | 1205880 | 180 | 341 | 27.22 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000121 | -1.3565 | 0.000137 | -3.509e-09 | 4.661e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0x3D13 | pH | 2021-03-09 05:43 | 0 | -1 | -1 | -1 | 1205881 | 180 | 341 | 27.23 | 26.03 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000134 | -1.3564 | 0.000125 | -6.533e-09 | -6.323e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0x27DA | pH | 2021-03-09 05:44 | 0 | -1 | -1 | -1 | 1205882 | 180 | 341 | 27.22 | 26.03 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000103 | -1.3562 | 0.000191 | -8.989e-09 | -4.317e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:18 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0x92A2 | pH | 2021-03-09 05:45 | 0 | -1 | -1 | -1 | 1205883 | 180 | 341 | 27.23 | 26.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000112 | -1.356 | 0.000138 | -1.041e-08 | -5.665e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0xF32A | pH | 2021-03-09 05:46 | 0 | -1 | -1 | -1 | 1205884 | 180 | 341 | 27.22 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.7e-05 | -1.3558 | 0.000187 | -8.794e-09 | 1.364e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0x41E6 | pH | 2021-03-09 05:47 | 0 | -1 | -1 | -1 | 1205885 | 180 | 341 | 27.2 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8.2e-05 | -1.3557 | 0.000121 | 6.86e-10 | 3.194e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0xDF0A | pH | 2021-03-09 05:48 | 0 | -1 | -1 | -1 | 1205886 | 180 | 341 | 27.22 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.1e-05 | -1.3555 | 0.000119 | -2.991e-09 | 1.04e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0x0AB8 | pH | 2021-03-09 05:49 | 0 | -1 | -1 | -1 | 1205887 | 180 | 341 | 27.2 | 26.06 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 9e-05 | -1.3554 | 0.000232 | -4.753e-09 | 1.768e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 20 | 0.20692 | 0.25 | 0 | 110.5805 | 0xE455 | pH | 2021-03-09 05:50 | 0 | -1 | -1 | -1 | 1205888 | 180 | 341 | 27.2 | 26.08 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 7.2e-05 | -1.3553 | 0.000161 | -8.779e-09 | 2.274e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 110.5805 | 0x6AF2 | pH | 2021-03-09 05:51 | 0 | -1 | -1 | -1 | 1205889 | 180 | 341 | 27.19 | 26.06 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000103 | -1.3551 | 0.000149 | -6.952e-09 | 2.649e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 110.5805 | 0x7E34 | pH | 2021-03-09 05:52 | 0 | -1 | -1 | -1 | 1205890 | 180 | 341 | 27.19 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.8e-05 | -1.3548 | 0.000187 | -1.171e-08 | -3.701e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.001 | 110.5797 | 0x7EB2 | pH | 2021-03-09 05:53 | 0 | -1 | -1 | -1 | 1205891 | 180 | 341 | 27.2 | 26.08 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 9.4e-05 | -1.3548 | 7.7e-05 | 2.442e-12 | -1.021e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 110.579 | 0xE6D2 | pH | 2021-03-09 05:54 | 0 | -1 | -1 | -1 | 1205892 | 180 | 341 | 27.18 | 26.06 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 6.6e-05 | -1.3546 | 0.000133 | -7.155e-09 | -7.731e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 110.579 | 0x89FE | pH | 2021-03-09 05:55 | 0 | -1 | -1 | -1 | 1205893 | 180 | 341 | 27.19 | 26.06 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000136 | -1.3545 | 0.000126 | -2.331e-09 | 2.027e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:19 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.005 | 110.5793 | 0xE955 | pH | 2021-03-09 05:56 | 0 | -1 | -1 | -1 | 1205894 | 180 | 341 | 27.18 | 26.08 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 6e-05 | -1.3546 | 0.000178 | -1.036e-08 | 1.163e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.004 | 110.5839 | 0xE6B5 | pH | 2021-03-09 05:57 | 0 | -1 | -1 | -1 | 1205895 | 180 | 341 | 27.18 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000125 | -1.3548 | 0.000162 | -6.787e-10 | -1.164e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.007 | 110.5896 | 0x5EC5 | pH | 2021-03-09 05:58 | 0 | -1 | -1 | -1 | 1205896 | 180 | 341 | 27.18 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000166 | -1.3555 | 0.000148 | -1.396e-09 | 1.253e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.007 | 110.5969 | 0x9FCC | pH | 2021-03-09 05:59 | 0 | -1 | -1 | -1 | 1205897 | 180 | 341 | 27.15 | 26.08 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000167 | -1.3565 | 0.000143 | -4.167e-09 | 1.56e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.011 | 110.6065 | 0xE28F | pH | 2021-03-09 06:00 | 0 | -1 | -1 | -1 | 1205898 | 180 | 341 | 27.15 | 26.08 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000147 | -1.3574 | 0.000228 | -7.558e-09 | -3.122e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.01 | 110.6178 | 0x3ED3 | pH | 2021-03-09 06:01 | 0 | -1 | -1 | -1 | 1205899 | 180 | 341 | 27.14 | 26.11 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000139 | -1.3583 | 0.000262 | -5.298e-10 | 6.709e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.013 | 110.6312 | 0xF021 | pH | 2021-03-09 06:02 | 0 | -1 | -1 | -1 | 1205900 | 180 | 341 | 27.12 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000197 | -1.3591 | 0.000193 | -6.853e-09 | 1.736e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.013 | 110.6454 | 0xB37E | pH | 2021-03-09 06:03 | 0 | -1 | -1 | -1 | 1205901 | 180 | 341 | 27.09 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000146 | -1.3598 | 0.00029 | 1.904e-10 | 9.881e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.016 | 110.6603 | 0xD3B1 | pH | 2021-03-09 06:04 | 0 | -1 | -1 | -1 | 1205902 | 180 | 341 | 27.09 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000205 | -1.3606 | 0.0002 | 2.507e-09 | 7.352e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.017 | 110.6749 | 0x661D | pH | 2021-03-09 06:05 | 0 | -1 | -1 | -1 | 1205903 | 180 | 341 | 27.08 | 26.13 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000184 | -1.3612 | 8.9e-05 | -8.599e-09 | -1.366e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:20 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.012 | 110.688 | 0x0DDE | pH | 2021-03-09 06:06 | 0 | -1 | -1 | -1 | 1205904 | 180 | 341 | 27.07 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000178 | -1.3614 | 0.000226 | -2.262e-09 | 1.878e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.01 | 110.7006 | 0xB972 | pH | 2021-03-09 06:07 | 0 | -1 | -1 | -1 | 1205905 | 180 | 341 | 27.04 | 26.13 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000174 | -1.3619 | 0.000189 | -8.309e-10 | -1.736e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.013 | 110.7126 | 0x8F47 | pH | 2021-03-09 06:08 | 0 | -1 | -1 | -1 | 1205906 | 180 | 341 | 27.04 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9e-05 | -1.3623 | 0.000183 | -8.064e-09 | 4.598e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.009 | 110.7233 | 0x0C7C | pH | 2021-03-09 06:09 | 0 | -1 | -1 | -1 | 1205907 | 180 | 341 | 27.03 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000142 | -1.3622 | 0.00017 | -5.876e-09 | 1.471e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.01 | 110.7336 | 0xF177 | pH | 2021-03-09 06:10 | 0 | -1 | -1 | -1 | 1205908 | 180 | 341 | 27.01 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000106 | -1.3623 | 0.000227 | 1.697e-09 | 5.41e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.008 | 110.7432 | 0xE642 | pH | 2021-03-09 06:11 | 0 | -1 | -1 | -1 | 1205909 | 180 | 341 | 27.01 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.3e-05 | -1.3621 | 0.00016 | 2.834e-09 | 6.068e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.007 | 110.7522 | 0x57D2 | pH | 2021-03-09 06:12 | 0 | -1 | -1 | -1 | 1205910 | 180 | 341 | 27.01 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000114 | -1.3619 | 0.000184 | 9.021e-09 | -7.239e-10 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.009 | 110.7608 | 0xDD04 | pH | 2021-03-09 06:13 | 0 | -1 | -1 | -1 | 1205911 | 180 | 341 | 27 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000125 | -1.3616 | 0.000148 | 3.861e-09 | -3.746e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.01 | 110.769 | 0xA591 | pH | 2021-03-09 06:14 | 0 | -1 | -1 | -1 | 1205912 | 180 | 341 | 27 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 4.7e-05 | -1.3612 | 0.00013 | -6.42e-09 | -1.114e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.008 | 110.7763 | 0x99D2 | pH | 2021-03-09 06:15 | 0 | -1 | -1 | -1 | 1205913 | 180 | 341 | 27.01 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00015 | -1.3609 | 0.000174 | -4.112e-09 | -9.239e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:21 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.004 | 110.7829 | 0x265E | pH | 2021-03-09 06:16 | 0 | -1 | -1 | -1 | 1205914 | 180 | 341 | 27 | 26.13 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 9.7e-05 | -1.3604 | 0.000101 | -2.121e-09 | -9.569e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.004 | 110.7894 | 0x520A | pH | 2021-03-09 06:17 | 0 | -1 | -1 | -1 | 1205915 | 180 | 341 | 27 | 26.13 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8.7e-05 | -1.3599 | 0.000183 | 9.688e-09 | 4.044e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.005 | 110.7951 | 0x419B | pH | 2021-03-09 06:18 | 0 | -1 | -1 | -1 | 1205916 | 180 | 341 | 27.01 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000108 | -1.3593 | 0.000147 | -7.238e-09 | -7.517e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | -0.005 | 110.8006 | 0x4D16 | pH | 2021-03-09 06:20 | 0 | -1 | -1 | -1 | 1205917 | 180 | 341 | 27.03 | 26.13 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 8.5e-05 | -1.3589 | 0.00014 | -7.509e-09 | 2.259e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.237 | 110.624 | 0xBE06 | pH | 2021-03-09 06:21 | 0 | -1 | -1 | -1 | 1205918 | 180 | 341 | 27.03 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7.4e-05 | -1.3676 | 0.000145 | 1.16e-08 | -8.434e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.235 | 110.3836 | 0xBA2C | pH | 2021-03-09 06:22 | 0 | -1 | -1 | -1 | 1205919 | 180 | 341 | 27.03 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000139 | -1.3755 | 0.000263 | -4.11e-09 | 1.07e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.235 | 110.1438 | 0xFE3B | pH | 2021-03-09 06:23 | 0 | -1 | -1 | -1 | 1205920 | 180 | 341 | 27.04 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000115 | -1.3806 | 0.000143 | 3.825e-11 | 2.014e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 109.9043 | 0x2A9D | pH | 2021-03-09 06:24 | 0 | -1 | -1 | -1 | 1205921 | 180 | 341 | 27.04 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 7.7e-05 | -1.3836 | 6.8e-05 | 1.55e-09 | -1.681e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.235 | 109.6644 | 0x1389 | pH | 2021-03-09 06:25 | 0 | -1 | -1 | -1 | 1205922 | 180 | 341 | 27.04 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000103 | -1.3861 | 0.000168 | -1.16e-08 | 4.361e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.235 | 109.426 | 0xB27B | pH | 2021-03-09 06:26 | 0 | -1 | -1 | -1 | 1205923 | 180 | 341 | 27.05 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000111 | -1.3873 | 0.000158 | -8.306e-09 | 4.518e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 109.1867 | 0x4CF0 | pH | 2021-03-09 06:27 | 0 | -1 | -1 | -1 | 1205924 | 180 | 341 | 27.07 | 26.13 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 6.4e-05 | -1.3884 | 0.000232 | -5.184e-10 | 5.483e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.425 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:22 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 108.9481 | 0xADB8 | pH | 2021-03-09 06:28 | 0 | -1 | -1 | -1 | 1205925 | 180 | 341 | 27.07 | 26.13 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000103 | -1.389 | 0.000179 | -6.473e-09 | -1.861e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 108.7092 | 0xCCB5 | pH | 2021-03-09 06:29 | 0 | -1 | -1 | -1 | 1205926 | 180 | 341 | 27.07 | 26.11 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00013 | -1.3954 | 0.000174 | -2.174e-09 | -3.559e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 108.47 | 0xB66E | pH | 2021-03-09 06:30 | 0 | -1 | -1 | -1 | 1205927 | 180 | 341 | 27.08 | 26.08 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000107 | -1.3958 | 8.2e-05 | 7.183e-09 | 5.959e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 108.2313 | 0x732B | pH | 2021-03-09 06:31 | 0 | -1 | -1 | -1 | 1205928 | 180 | 341 | 27.09 | 26.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 7.5e-05 | -1.3955 | 0.000101 | -6.632e-09 | 7.428e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 107.9929 | 0xE539 | pH | 2021-03-09 06:32 | 0 | -1 | -1 | -1 | 1205929 | 180 | 341 | 27.09 | 26.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000116 | -1.3945 | 0.000229 | -2.279e-10 | 1.37e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.234 | 107.7545 | 0x0840 | pH | 2021-03-09 06:33 | 0 | -1 | -1 | -1 | 1205930 | 180 | 341 | 27.11 | 26.06 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.7e-05 | -1.3931 | 0.000188 | 5.062e-10 | -3.114e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.005 | 107.5782 | 0x43C3 | pH | 2021-03-09 06:34 | 0 | -1 | -1 | -1 | 1205931 | 180 | 341 | 27.12 | 26.03 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000104 | -1.3909 | 7.5e-05 | -4.159e-09 | 8.414e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5764 | 0x3737 | pH | 2021-03-09 06:35 | 0 | -1 | -1 | -1 | 1205932 | 180 | 341 | 27.12 | 26.01 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000109 | -1.3876 | 0.000174 | -6.104e-10 | -4.311e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5759 | 0x3676 | pH | 2021-03-09 06:36 | 0 | -1 | -1 | -1 | 1205933 | 180 | 341 | 27.12 | 26.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.00013 | -1.3846 | 0.000101 | -4.135e-09 | -1.313e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xA2B9 | pH | 2021-03-09 06:37 | 0 | -1 | -1 | -1 | 1205934 | 180 | 341 | 27.14 | 25.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 7.6e-05 | -1.3818 | 0.000141 | -9.521e-11 | 1.546e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:23 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xB504 | pH | 2021-03-09 06:38 | 0 | -1 | -1 | -1 | 1205935 | 180 | 341 | 27.15 | 25.98 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000111 | -1.3797 | 0.000179 | -1.418e-08 | 7.377e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xC8E9 | pH | 2021-03-09 06:39 | 0 | -1 | -1 | -1 | 1205936 | 180 | 341 | 27.15 | 25.98 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 9.6e-05 | -1.3783 | 0.000168 | -2.402e-09 | -4.11e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xB459 | pH | 2021-03-09 06:40 | 0 | -1 | -1 | -1 | 1205937 | 180 | 341 | 27.18 | 25.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000118 | -1.3767 | 0.000167 | 1.249e-08 | -1.261e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x2CB6 | pH | 2021-03-09 06:41 | 0 | -1 | -1 | -1 | 1205938 | 180 | 341 | 27.16 | 25.93 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.3e-05 | -1.3752 | 0.000144 | -1.427e-09 | -7.15e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x6989 | pH | 2021-03-09 06:42 | 0 | -1 | -1 | -1 | 1205939 | 180 | 341 | 27.18 | 25.93 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000121 | -1.3739 | 0.000166 | -4.793e-09 | -7.75e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x4FAC | pH | 2021-03-09 06:43 | 0 | -1 | -1 | -1 | 1205940 | 180 | 341 | 27.19 | 25.91 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000133 | -1.3727 | 0.000177 | -8.391e-09 | 4.903e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xBDC0 | pH | 2021-03-09 06:44 | 0 | -1 | -1 | -1 | 1205941 | 180 | 341 | 27.19 | 25.91 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000103 | -1.3714 | 0.00017 | 3.969e-09 | -7.118e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x12FF | pH | 2021-03-09 06:45 | 0 | -1 | -1 | -1 | 1205942 | 180 | 341 | 27.19 | 25.88 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.00015 | -1.3705 | 0.000231 | -2.47e-09 | 5.193e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4248 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xDD28 | pH | 2021-03-09 06:46 | 0 | -1 | -1 | -1 | 1205943 | 180 | 341 | 27.19 | 25.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.7e-05 | -1.3694 | 0.000209 | -1.882e-09 | 6.786e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x7050 | pH | 2021-03-09 06:47 | 0 | -1 | -1 | -1 | 1205944 | 180 | 341 | 27.19 | 25.86 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000121 | -1.3687 | 0.000109 | -2.793e-09 | 8.348e-10 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xDE4E | pH | 2021-03-09 06:48 | 0 | -1 | -1 | -1 | 1205945 | 180 | 341 | 27.2 | 25.83 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 8.5e-05 | -1.3679 | 0.000146 | -6.109e-09 | 4.989e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:24 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x34DA | pH | 2021-03-09 06:49 | 0 | -1 | -1 | -1 | 1205946 | 180 | 341 | 27.2 | 25.83 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 0.000104 | -1.3673 | 0.000148 | -9.986e-09 | -2.368e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xCFC6 | pH | 2021-03-09 06:50 | 0 | -1 | -1 | -1 | 1205947 | 180 | 341 | 27.2 | 25.83 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2064 | 9.4e-05 | -1.3664 | 0.000165 | 1.896e-09 | -3.968e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x9E78 | pH | 2021-03-09 06:51 | 0 | -1 | -1 | -1 | 1205948 | 180 | 341 | 27.2 | 25.81 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 0.000122 | -1.3657 | 0.000102 | -2.624e-09 | 3.383e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4441 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xF909 | pH | 2021-03-09 06:52 | 0 | -1 | -1 | -1 | 1205949 | 180 | 341 | 27.19 | 25.78 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2064 | 8.6e-05 | -1.3651 | 0.000168 | -3.091e-09 | -2.184e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x2FB3 | pH | 2021-03-09 06:53 | 0 | -1 | -1 | -1 | 1205950 | 180 | 341 | 27.19 | 25.71 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 9.5e-05 | -1.3644 | 0.000188 | -3.456e-09 | -1.406e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x21D1 | pH | 2021-03-09 06:54 | 0 | -1 | -1 | -1 | 1205951 | 180 | 341 | 27.2 | 25.71 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.000126 | -1.3638 | 0.000171 | -5.391e-09 | 4.296e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0xAB11 | pH | 2021-03-09 06:55 | 0 | -1 | -1 | -1 | 1205952 | 180 | 341 | 27.19 | 25.73 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 0.000114 | -1.3633 | 0.000109 | -5.246e-09 | -4.989e-09 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0 | 107.5758 | 0x1E94 | pH | 2021-03-09 06:56 | 0 | -1 | -1 | -1 | 1205953 | 180 | 341 | 27.19 | 25.73 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 9.2e-05 | -1.3626 | 0.000222 | 5.005e-10 | 1.003e-08 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.001 | 107.575 | 0x9308 | pH | 2021-03-09 06:57 | 0 | -1 | -1 | -1 | 1205954 | 180 | 341 | 27.18 | 25.71 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2065 | 8.8e-05 | -1.3623 | 0.000162 | -1.816e-09 | 4.46e-09 | 0.00041385 | -1.2065 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 0.20692 | 0.25 | 0.001 | 107.5741 | 0x0E4A | pH | 2021-03-09 06:58 | 0 | -1 | -1 | -1 | 1205955 | 180 | 341 | 27.18 | 25.73 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2065 | 0.0001 | -1.3617 | 0.000171 | -5.566e-10 | 1.053e-08 | 0.00041385 | -1.2064 | 0.90444 | 17.6664 | 0.23773 | 1.7976e-05 | -1.4442 | -1.4249 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:25 | 20 | 20 | 35 | 5.5869 | 0.25 | 0 | 107.5738 | 0xEB78 | pH | 2021-03-09 06:59 | 0 | -1 | -1 | -1 | 1205956 | 180 | 341 | 27.18 | 25.71 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2071 | 0.000101 | -1.3613 | 0.000197 | -7.192e-09 | 7.944e-09 | 0.011174 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 8.1003e-06 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 38.2809 | 0.25 | 0 | 107.5732 | 0x4171 | pH | 2021-03-09 07:00 | 0 | -1 | -1 | -1 | 1205957 | 180 | 341 | 27.16 | 25.66 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.1e-05 | -1.3606 | 0.00016 | 1.972e-09 | -6.811e-09 | 0.076562 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -5.1911e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 39.1086 | 0.25 | 0 | 107.5729 | 0xBB2B | pH | 2021-03-09 07:01 | 0 | -1 | -1 | -1 | 1205958 | 180 | 341 | 27.16 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6.2e-05 | -1.3602 | 0.000198 | -4.371e-09 | -1.48e-08 | 0.078217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.343e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 38.6258 | 0.25 | 0 | 107.5726 | 0x56E0 | pH | 2021-03-09 07:02 | 0 | -1 | -1 | -1 | 1205959 | 180 | 341 | 27.15 | 25.68 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000111 | -1.3596 | 0.000193 | -5.998e-10 | 1.213e-09 | 0.077252 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -5.2544e-05 | -1.4447 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 38.2809 | 0.25 | 0 | 107.5722 | 0xBFD1 | pH | 2021-03-09 07:03 | 0 | -1 | -1 | -1 | 1205960 | 180 | 341 | 27.15 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.00012 | -1.3592 | 0.000223 | 2.115e-09 | -5.306e-10 | 0.076562 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.1911e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 38.074 | 0.25 | 0 | 107.5694 | 0x96B5 | pH | 2021-03-09 07:04 | 0 | -1 | -1 | -1 | 1205961 | 180 | 341 | 27.14 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000143 | -1.3588 | 0.000167 | 1.115e-10 | 1.936e-08 | 0.076148 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.1531e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 37.936 | 0.25 | -0.233 | 107.745 | 0x0D4D | pH | 2021-03-09 07:05 | 0 | -1 | -1 | -1 | 1205962 | 180 | 341 | 27.14 | 25.63 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000113 | -1.3575 | 0.000129 | -7.55e-09 | 2.077e-10 | 0.075872 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.1278e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 37.7981 | 0.25 | -0.233 | 107.9829 | 0x3649 | pH | 2021-03-09 07:06 | 0 | -1 | -1 | -1 | 1205963 | 180 | 341 | 27.14 | 25.66 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000103 | -1.3563 | 0.000169 | 6.477e-09 | 1.022e-08 | 0.075596 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.1024e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 37.6601 | 0.25 | -0.234 | 108.2209 | 0xE3AD | pH | 2021-03-09 07:07 | 0 | -1 | -1 | -1 | 1205964 | 180 | 341 | 27.14 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6.8e-05 | -1.3554 | 0.000126 | -4.546e-09 | 1.566e-08 | 0.07532 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0771e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 37.5912 | 0.25 | -0.234 | 108.459 | 0x4638 | pH | 2021-03-09 07:08 | 0 | -1 | -1 | -1 | 1205965 | 180 | 341 | 27.11 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.4e-05 | -1.3547 | 9.5e-05 | 1.693e-10 | -5.821e-09 | 0.075182 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -5.0645e-05 | -1.4446 | -1.4253 | -1.4253 | -1.4253 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:26 | 20 | 20 | 35 | 37.5222 | 0.25 | -0.234 | 108.6976 | 0x7C78 | pH | 2021-03-09 07:09 | 0 | -1 | -1 | -1 | 1205966 | 180 | 341 | 27.09 | 25.63 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000106 | -1.3539 | 0.000133 | 7.845e-10 | 1.294e-08 | 0.075044 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -5.0518e-05 | -1.4446 | -1.4253 | -1.4253 | -1.4253 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.4532 | 0.25 | -0.234 | 108.9362 | 0x3E17 | pH | 2021-03-09 07:10 | 0 | -1 | -1 | -1 | 1205967 | 180 | 341 | 27.09 | 25.63 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000111 | -1.3536 | 0.000229 | 5.697e-11 | -6.209e-09 | 0.074906 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0391e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.3842 | 0.25 | -0.235 | 109.1753 | 0x7420 | pH | 2021-03-09 07:11 | 0 | -1 | -1 | -1 | 1205968 | 180 | 341 | 27.08 | 25.63 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.9e-05 | -1.3535 | 0.000153 | -6.599e-09 | 2.789e-09 | 0.074768 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0265e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.3153 | 0.25 | -0.235 | 109.4147 | 0xF2CE | pH | 2021-03-09 07:12 | 0 | -1 | -1 | -1 | 1205969 | 180 | 341 | 27.07 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000109 | -1.3533 | 0.000169 | 2.333e-09 | 9.513e-09 | 0.074631 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.3153 | 0.25 | -0.234 | 109.6541 | 0x747C | pH | 2021-03-09 07:13 | 0 | -1 | -1 | -1 | 1205970 | 180 | 341 | 27.08 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2069 | 0.000133 | -1.3518 | 0.000151 | -5.005e-09 | -9.156e-09 | 0.074631 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.2463 | 0.25 | -0.235 | 109.8937 | 0x7065 | pH | 2021-03-09 07:14 | 0 | -1 | -1 | -1 | 1205971 | 180 | 341 | 27.08 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000159 | -1.3495 | 0.000199 | 2.569e-09 | -6.657e-09 | 0.074493 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.2463 | 0.25 | -0.236 | 110.134 | 0x7463 | pH | 2021-03-09 07:15 | 0 | -1 | -1 | -1 | 1205972 | 180 | 341 | 27.08 | 25.71 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000131 | -1.347 | 0.000149 | 3.126e-09 | 6.06e-10 | 0.074493 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.1773 | 0.25 | -0.236 | 110.3741 | 0x3AEF | pH | 2021-03-09 07:16 | 0 | -1 | -1 | -1 | 1205973 | 180 | 341 | 27.07 | 25.71 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000136 | -1.3436 | 0.000141 | 7.618e-09 | 2.565e-09 | 0.074355 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4446 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.1773 | 0.25 | -0.237 | 110.6154 | 0xAF94 | pH | 2021-03-09 07:17 | 0 | -1 | -1 | -1 | 1205974 | 180 | 341 | 27.08 | 25.71 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2069 | 0.000147 | -1.3397 | 9.4e-05 | 1.187e-08 | -2.509e-09 | 0.074355 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.1083 | 0.25 | -0.012 | 110.818 | 0x23BA | pH | 2021-03-09 07:18 | 0 | -1 | -1 | -1 | 1205975 | 180 | 341 | 27.07 | 25.73 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.00014 | -1.3354 | 0.000153 | -7.118e-09 | -4.744e-10 | 0.074217 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4446 | -1.4253 | -1.4253 | -1.4253 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.1083 | 0.25 | -0.003 | 110.8239 | 0x15FD | pH | 2021-03-09 07:20 | 0 | -1 | -1 | -1 | 1205976 | 180 | 341 | 27.09 | 25.66 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.7e-05 | -1.3335 | 0.000159 | -1.963e-08 | -1.432e-08 | 0.074217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:27 | 20 | 20 | 35 | 37.1083 | 0.25 | 0 | 110.8268 | 0x6F62 | pH | 2021-03-09 07:21 | 0 | -1 | -1 | -1 | 1205977 | 180 | 341 | 27.08 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000121 | -1.3322 | 9.7e-05 | -5.599e-10 | -1.53e-08 | 0.074217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.0394 | 0.25 | -0.001 | 110.8291 | 0xA938 | pH | 2021-03-09 07:22 | 0 | -1 | -1 | -1 | 1205978 | 180 | 341 | 27.07 | 25.66 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9e-05 | -1.3312 | 0.000138 | -2.905e-10 | -4.069e-09 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.0394 | 0.25 | -0.003 | 110.8308 | 0x3D69 | pH | 2021-03-09 07:23 | 0 | -1 | -1 | -1 | 1205979 | 180 | 341 | 27.07 | 25.68 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000126 | -1.3307 | 0.000126 | 4.679e-10 | 1.519e-08 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.4532 | 0.25 | 0 | 110.8321 | 0x76F1 | pH | 2021-03-09 07:24 | 0 | -1 | -1 | -1 | 1205980 | 180 | 341 | 27.08 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.9e-05 | -1.3301 | 0.000172 | -6.885e-10 | -3.174e-09 | 0.074906 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0391e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.3842 | 0.25 | 0 | 110.8333 | 0xCD82 | pH | 2021-03-09 07:25 | 0 | -1 | -1 | -1 | 1205981 | 180 | 341 | 27.08 | 25.66 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.9e-05 | -1.3295 | 0.00018 | -9.033e-11 | -1.03e-10 | 0.074768 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0265e-05 | -1.4446 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.3842 | 0.25 | -0.004 | 110.8345 | 0x8490 | pH | 2021-03-09 07:26 | 0 | -1 | -1 | -1 | 1205982 | 180 | 341 | 27.08 | 25.63 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000146 | -1.3286 | 0.000142 | 1.198e-08 | 7.692e-09 | 0.074768 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0265e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.3153 | 0.25 | 0 | 110.8352 | 0x73F0 | pH | 2021-03-09 07:27 | 0 | -1 | -1 | -1 | 1205983 | 180 | 341 | 27.08 | 25.63 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000113 | -1.3286 | 0.000141 | -8.607e-09 | -3.778e-09 | 0.074631 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.3153 | 0.25 | -0.001 | 110.8363 | 0x9661 | pH | 2021-03-09 07:28 | 0 | -1 | -1 | -1 | 1205984 | 180 | 341 | 27.09 | 25.61 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000137 | -1.3277 | 0.00019 | -4.631e-10 | 1.343e-09 | 0.074631 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.3153 | 0.25 | -0.003 | 110.8371 | 0x9181 | pH | 2021-03-09 07:29 | 0 | -1 | -1 | -1 | 1205985 | 180 | 341 | 27.08 | 25.58 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000127 | -1.3272 | 0.000116 | -3.053e-09 | 5.708e-09 | 0.074631 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.3153 | 0.25 | 0 | 110.8377 | 0x0884 | pH | 2021-03-09 07:30 | 0 | -1 | -1 | -1 | 1205986 | 180 | 341 | 27.09 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5.7e-05 | -1.3263 | 0.000139 | -8.527e-09 | 1.29e-08 | 0.074631 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:28 | 20 | 20 | 35 | 37.2463 | 0.25 | 0 | 110.8383 | 0x5CB2 | pH | 2021-03-09 07:31 | 0 | -1 | -1 | -1 | 1205987 | 180 | 341 | 27.09 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000121 | -1.3256 | 0.00015 | -9.109e-09 | 6.78e-09 | 0.074493 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4446 | -1.4253 | -1.4253 | -1.4253 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.2463 | 0.25 | -0.001 | 110.8392 | 0x6A1D | pH | 2021-03-09 07:32 | 0 | -1 | -1 | -1 | 1205988 | 180 | 341 | 27.09 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.9e-05 | -1.3252 | 0.000142 | 4.493e-09 | -8.838e-09 | 0.074493 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.2463 | 0.25 | 0 | 110.8396 | 0x927A | pH | 2021-03-09 07:33 | 0 | -1 | -1 | -1 | 1205989 | 180 | 341 | 27.11 | 25.58 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.7e-05 | -1.3248 | 0.000128 | -9.079e-09 | 9.065e-09 | 0.074493 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1773 | 0.25 | -0.001 | 110.8404 | 0x2F3B | pH | 2021-03-09 07:34 | 0 | -1 | -1 | -1 | 1205990 | 180 | 341 | 27.11 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.2e-05 | -1.324 | 0.000144 | -1.377e-09 | 6.961e-09 | 0.074355 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4446 | -1.4253 | -1.4253 | -1.4253 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1773 | 0.25 | 0 | 110.8407 | 0x28DD | pH | 2021-03-09 07:35 | 0 | -1 | -1 | -1 | 1205991 | 180 | 341 | 27.11 | 25.55 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2069 | 0.000111 | -1.3237 | 0.000168 | -6.865e-09 | -3.273e-09 | 0.074355 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1773 | 0.25 | 0 | 110.8411 | 0x345E | pH | 2021-03-09 07:36 | 0 | -1 | -1 | -1 | 1205992 | 180 | 341 | 27.11 | 25.5 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000121 | -1.3234 | 0.000133 | -7.144e-09 | -4.211e-09 | 0.074355 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1773 | 0.25 | 0 | 110.8416 | 0x4F42 | pH | 2021-03-09 07:37 | 0 | -1 | -1 | -1 | 1205993 | 180 | 341 | 27.12 | 25.5 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.00013 | -1.3228 | 0.000137 | -1.23e-09 | 1.006e-10 | 0.074355 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1083 | 0.25 | 0 | 110.8419 | 0xECE2 | pH | 2021-03-09 07:38 | 0 | -1 | -1 | -1 | 1205994 | 180 | 341 | 27.12 | 25.5 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 8.7e-05 | -1.3221 | 0.000158 | -7.203e-09 | -8.593e-09 | 0.074217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1083 | 0.25 | 0 | 110.8424 | 0x7BC4 | pH | 2021-03-09 07:39 | 0 | -1 | -1 | -1 | 1205995 | 180 | 341 | 27.11 | 25.5 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000134 | -1.3218 | 0.000134 | -1.397e-08 | 7.952e-09 | 0.074217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1083 | 0.25 | 0 | 110.8428 | 0x3021 | pH | 2021-03-09 07:40 | 0 | -1 | -1 | -1 | 1205996 | 180 | 341 | 27.12 | 25.45 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000107 | -1.3212 | 0.000115 | -6.257e-09 | -1.123e-08 | 0.074217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:29 | 20 | 20 | 35 | 37.1083 | 0.25 | 0 | 110.8433 | 0x7F00 | pH | 2021-03-09 07:41 | 0 | -1 | -1 | -1 | 1205997 | 180 | 341 | 27.12 | 25.48 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.5e-05 | -1.3206 | 0.000151 | -1.366e-09 | 8.558e-09 | 0.074217 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.8435 | 0x888D | pH | 2021-03-09 07:42 | 0 | -1 | -1 | -1 | 1205998 | 180 | 341 | 27.12 | 25.5 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5e-05 | -1.3196 | 0.00018 | -1.104e-09 | -3.507e-09 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.844 | 0x770C | pH | 2021-03-09 07:43 | 0 | -1 | -1 | -1 | 1205999 | 180 | 341 | 27.14 | 25.48 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.6e-05 | -1.3191 | 0.000171 | 3.403e-09 | 1.728e-08 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.8441 | 0xCE6D | pH | 2021-03-09 07:44 | 0 | -1 | -1 | -1 | 1206000 | 180 | 341 | 27.14 | 25.45 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000116 | -1.3178 | 0.000146 | 4.334e-09 | 1.073e-08 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.8445 | 0x52F1 | pH | 2021-03-09 07:45 | 0 | -1 | -1 | -1 | 1206001 | 180 | 341 | 27.12 | 25.45 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.9e-05 | -1.3167 | 0.000197 | -6.72e-09 | -2.626e-09 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.8449 | 0xA32E | pH | 2021-03-09 07:46 | 0 | -1 | -1 | -1 | 1206002 | 180 | 341 | 27.12 | 25.43 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.5e-05 | -1.3156 | 0.000164 | -2.302e-09 | 5.612e-09 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.8452 | 0x8C96 | pH | 2021-03-09 07:47 | 0 | -1 | -1 | -1 | 1206003 | 180 | 341 | 27.14 | 25.45 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000136 | -1.3143 | 0.000271 | -1.107e-10 | -1.179e-08 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 37.0394 | 0.25 | 0 | 110.8454 | 0xD321 | pH | 2021-03-09 07:48 | 0 | -1 | -1 | -1 | 1206004 | 180 | 341 | 27.14 | 25.43 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.7e-05 | -1.3132 | 0.000239 | -5.869e-09 | 1.178e-08 | 0.074079 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 36.9704 | 0.25 | -0.001 | 110.8457 | 0xCCE0 | pH | 2021-03-09 07:49 | 0 | -1 | -1 | -1 | 1206005 | 180 | 341 | 27.14 | 25.43 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000128 | -1.3119 | 0.000245 | 3.433e-09 | 8.949e-10 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:30 | 20 | 20 | 35 | 36.9704 | 0.25 | 0 | 110.8462 | 0x0937 | pH | 2021-03-09 07:50 | 0 | -1 | -1 | -1 | 1206006 | 180 | 341 | 27.14 | 25.41 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.3e-05 | -1.3102 | 0.000115 | -1.349e-08 | 1.223e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 20 | 36.9704 | 0.25 | 0 | 110.8463 | 0x07FE | pH | 2021-03-09 07:51 | 0 | -1 | -1 | -1 | 1206007 | 180 | 341 | 27.16 | 25.43 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000153 | -1.3088 | 0.000101 | -4.881e-09 | -1.105e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:30 | 20 | 20 | 20 | 36.9704 | 0.25 | 0 | 110.8469 | 0x3E8F | pH | 2021-03-09 07:52 | 0 | -1 | -1 | -1 | 1206008 | 180 | 341 | 27.16 | 25.41 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 3.8e-05 | -1.3078 | 0.000149 | -7.62e-09 | -4.89e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9704 | 0.25 | 0 | 110.8469 | 0x751A | pH | 2021-03-09 07:53 | 0 | -1 | -1 | -1 | 1206009 | 180 | 341 | 27.18 | 25.41 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000105 | -1.3054 | 0.000192 | -7.321e-09 | -4.788e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9704 | 0.25 | 0 | 110.8469 | 0x5869 | pH | 2021-03-09 07:54 | 0 | -1 | -1 | -1 | 1206010 | 180 | 341 | 27.16 | 25.38 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.2e-05 | -1.3036 | 0.000179 | 3.209e-09 | 1.248e-08 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9704 | 0.25 | 0.002 | 110.845 | 0xEA70 | pH | 2021-03-09 07:55 | 0 | -1 | -1 | -1 | 1206011 | 180 | 341 | 27.16 | 25.38 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.3e-05 | -1.3022 | 0.000161 | 1.955e-09 | 2.476e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9704 | 0.25 | 0.005 | 110.8418 | 0x2000 | pH | 2021-03-09 07:56 | 0 | -1 | -1 | -1 | 1206012 | 180 | 341 | 27.18 | 25.38 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.6e-05 | -1.3014 | 0.000154 | -2.432e-09 | 1.821e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9704 | 0.25 | 0.002 | 110.838 | 0xCF22 | pH | 2021-03-09 07:57 | 0 | -1 | -1 | -1 | 1206013 | 180 | 341 | 27.19 | 25.38 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000118 | -1.301 | 0.000144 | 6.919e-09 | -3.62e-09 | 0.073941 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.005 | 110.8333 | 0x68C8 | pH | 2021-03-09 07:58 | 0 | -1 | -1 | -1 | 1206014 | 180 | 341 | 27.19 | 25.36 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000122 | -1.3002 | 0.000204 | 7.9e-09 | 4.622e-09 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.006 | 110.828 | 0x40DF | pH | 2021-03-09 07:59 | 0 | -1 | -1 | -1 | 1206015 | 180 | 341 | 27.2 | 25.36 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 9.9e-05 | -1.3001 | 0.000184 | -4.883e-11 | 2.7e-09 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.004 | 110.8221 | 0x5565 | pH | 2021-03-09 08:00 | 0 | -1 | -1 | -1 | 1206016 | 180 | 341 | 27.2 | 25.36 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000112 | -1.2999 | 0.000138 | -5.324e-09 | -4.929e-09 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.006 | 110.8157 | 0x29D5 | pH | 2021-03-09 08:01 | 0 | -1 | -1 | -1 | 1206017 | 180 | 341 | 27.2 | 25.36 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000115 | -1.2997 | 0.000137 | -4.517e-09 | 1.344e-08 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:31 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.007 | 110.8088 | 0xB58D | pH | 2021-03-09 08:02 | 0 | -1 | -1 | -1 | 1206018 | 180 | 341 | 27.22 | 25.33 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000104 | -1.3001 | 0.000168 | 5.043e-09 | 8.446e-10 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.005 | 110.8016 | 0x47D4 | pH | 2021-03-09 08:03 | 0 | -1 | -1 | -1 | 1206019 | 180 | 341 | 27.23 | 25.28 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000128 | -1.3005 | 0.00019 | -2.432e-09 | 1.633e-08 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.9014 | 0.25 | 0.008 | 110.7938 | 0x9D1F | pH | 2021-03-09 08:04 | 0 | -1 | -1 | -1 | 1206020 | 180 | 341 | 27.23 | 25.3 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000104 | -1.3002 | 0.000107 | 2.217e-09 | 5.136e-09 | 0.073803 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.008 | 110.786 | 0x1AD7 | pH | 2021-03-09 08:05 | 0 | -1 | -1 | -1 | 1206021 | 180 | 341 | 27.23 | 25.28 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5.8e-05 | -1.3003 | 0.000138 | -6.299e-10 | -8.364e-09 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.006 | 110.7782 | 0xE6E5 | pH | 2021-03-09 08:06 | 0 | -1 | -1 | -1 | 1206022 | 180 | 341 | 27.23 | 25.28 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000112 | -1.3014 | 0.000149 | -1.124e-08 | 5.777e-09 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.008 | 110.77 | 0x3B25 | pH | 2021-03-09 08:07 | 0 | -1 | -1 | -1 | 1206023 | 180 | 341 | 27.24 | 25.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.00012 | -1.3009 | 0.000164 | -6.021e-09 | 1.318e-08 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.006 | 110.7617 | 0xCD62 | pH | 2021-03-09 08:08 | 0 | -1 | -1 | -1 | 1206024 | 180 | 341 | 27.26 | 25.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000107 | -1.301 | 0.000185 | -2.698e-09 | 1.071e-09 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.008 | 110.7531 | 0xFC2D | pH | 2021-03-09 08:09 | 0 | -1 | -1 | -1 | 1206025 | 180 | 341 | 27.26 | 25.28 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000103 | -1.3015 | 0.000148 | -4.576e-09 | 8.905e-10 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.007 | 110.7446 | 0x5D8E | pH | 2021-03-09 08:10 | 0 | -1 | -1 | -1 | 1206026 | 180 | 341 | 27.26 | 25.23 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5.7e-05 | -1.3012 | 0.000176 | 3.403e-09 | 8.618e-09 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.8324 | 0.25 | 0.007 | 110.7362 | 0xC30E | pH | 2021-03-09 08:11 | 0 | -1 | -1 | -1 | 1206027 | 180 | 341 | 27.27 | 25.28 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.00011 | -1.3013 | 0.000154 | -3.26e-09 | -1.627e-09 | 0.073665 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:32 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.006 | 110.7279 | 0x0C08 | pH | 2021-03-09 08:12 | 0 | -1 | -1 | -1 | 1206028 | 180 | 341 | 27.27 | 25.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000108 | -1.3015 | 0.000146 | -2.968e-09 | 7.08e-09 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.007 | 110.7195 | 0xA4E1 | pH | 2021-03-09 08:13 | 0 | -1 | -1 | -1 | 1206029 | 180 | 341 | 27.28 | 25.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000104 | -1.3015 | 0.000164 | -2.944e-09 | 5.516e-10 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.007 | 110.7111 | 0x154A | pH | 2021-03-09 08:14 | 0 | -1 | -1 | -1 | 1206030 | 180 | 341 | 27.28 | 25.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.4e-05 | -1.302 | 0.000182 | 2.735e-09 | 4.127e-09 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.008 | 110.703 | 0x5471 | pH | 2021-03-09 08:15 | 0 | -1 | -1 | -1 | 1206031 | 180 | 341 | 27.31 | 25.16 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000126 | -1.3017 | 8.5e-05 | -5.93e-09 | 1.497e-08 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.009 | 110.6946 | 0xAA6B | pH | 2021-03-09 08:16 | 0 | -1 | -1 | -1 | 1206032 | 180 | 341 | 27.3 | 25.18 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000126 | -1.302 | 0.000145 | 4.651e-09 | 7.164e-09 | 0.073527 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.008 | 110.6865 | 0x5CA2 | pH | 2021-03-09 08:17 | 0 | -1 | -1 | -1 | 1206033 | 180 | 341 | 27.3 | 25.16 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 9.7e-05 | -1.3016 | 0.000145 | 5.926e-09 | -4.133e-09 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.008 | 110.6782 | 0x9FC1 | pH | 2021-03-09 08:18 | 0 | -1 | -1 | -1 | 1206034 | 180 | 341 | 27.3 | 25.16 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.207 | 0.000139 | -1.3017 | 0.000125 | 1.663e-09 | -1.567e-08 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.007 | 110.6706 | 0x7354 | pH | 2021-03-09 08:19 | 0 | -1 | -1 | -1 | 1206035 | 180 | 341 | 27.31 | 25.16 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000122 | -1.3014 | 0.000199 | 4.045e-10 | 1.12e-08 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.7635 | 0.25 | 0.238 | 110.4898 | 0xC5F4 | pH | 2021-03-09 08:20 | 0 | -1 | -1 | -1 | 1206036 | 180 | 341 | 27.31 | 25.13 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.5e-05 | -1.3094 | 0.000196 | 1.014e-09 | 1.51e-09 | 0.073527 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.237 | 110.2478 | 0x61DF | pH | 2021-03-09 08:22 | 0 | -1 | -1 | -1 | 1206037 | 180 | 341 | 27.32 | 25.13 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000131 | -1.3156 | 0.00019 | -1.09e-10 | -4.374e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.236 | 110.0066 | 0x2356 | pH | 2021-03-09 08:23 | 0 | -1 | -1 | -1 | 1206038 | 180 | 341 | 27.32 | 25.1 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.7e-05 | -1.3211 | 0.000147 | 2.267e-09 | -1.09e-08 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:33 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.236 | 109.7663 | 0x2A3B | pH | 2021-03-09 08:24 | 0 | -1 | -1 | -1 | 1206039 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000129 | -1.3253 | 0.000182 | 5.452e-11 | 1.456e-10 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.236 | 109.5262 | 0x569E | pH | 2021-03-09 08:25 | 0 | -1 | -1 | -1 | 1206040 | 180 | 341 | 27.32 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.3e-05 | -1.3283 | 0.000169 | 2.275e-09 | -2.27e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.236 | 109.2865 | 0x3826 | pH | 2021-03-09 08:26 | 0 | -1 | -1 | -1 | 1206041 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000116 | -1.3315 | 0.000175 | -1.132e-08 | -1.322e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.235 | 109.0466 | 0xB5C9 | pH | 2021-03-09 08:27 | 0 | -1 | -1 | -1 | 1206042 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5.3e-05 | -1.3337 | 0.000158 | -5.47e-09 | -4.512e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.235 | 108.8072 | 0x9646 | pH | 2021-03-09 08:28 | 0 | -1 | -1 | -1 | 1206043 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.7e-05 | -1.3354 | 0.000171 | 1.864e-10 | 1.179e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.235 | 108.5673 | 0x7D85 | pH | 2021-03-09 08:29 | 0 | -1 | -1 | -1 | 1206044 | 180 | 341 | 27.3 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000118 | -1.3366 | 0.000165 | 3.247e-09 | -1.006e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.235 | 108.328 | 0xF0F1 | pH | 2021-03-09 08:30 | 0 | -1 | -1 | -1 | 1206045 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6e-05 | -1.3381 | 0.000146 | -1.44e-09 | -1.153e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.235 | 108.0889 | 0x07F1 | pH | 2021-03-09 08:31 | 0 | -1 | -1 | -1 | 1206046 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000159 | -1.3387 | 0.000195 | -7.359e-09 | -9.861e-10 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6945 | 0.25 | 0.235 | 107.8496 | 0x65E8 | pH | 2021-03-09 08:32 | 0 | -1 | -1 | -1 | 1206047 | 180 | 341 | 27.31 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000142 | -1.3389 | 0.000175 | 6.833e-09 | 8.332e-09 | 0.073389 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.235 | 107.6108 | 0x670E | pH | 2021-03-09 08:33 | 0 | -1 | -1 | -1 | 1206048 | 180 | 341 | 27.3 | 25.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000109 | -1.3393 | 0.000164 | -2.482e-09 | 1.404e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:34 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.235 | 107.3716 | 0x6FA1 | pH | 2021-03-09 08:34 | 0 | -1 | -1 | -1 | 1206049 | 180 | 341 | 27.3 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000125 | -1.3399 | 0.000147 | -4.148e-09 | -6.785e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.005 | 107.2452 | 0x892D | pH | 2021-03-09 08:35 | 0 | -1 | -1 | -1 | 1206050 | 180 | 341 | 27.28 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000121 | -1.3396 | 0.000168 | -9.288e-09 | 5.519e-10 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.004 | 107.2398 | 0xB16E | pH | 2021-03-09 08:36 | 0 | -1 | -1 | -1 | 1206051 | 180 | 341 | 27.3 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.2e-05 | -1.3392 | 0.000147 | -5.518e-09 | -2.096e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.003 | 107.2357 | 0xB584 | pH | 2021-03-09 08:37 | 0 | -1 | -1 | -1 | 1206052 | 180 | 341 | 27.3 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.8e-05 | -1.3389 | 0.000129 | 9.459e-09 | -1.736e-08 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.003 | 107.232 | 0xFC55 | pH | 2021-03-09 08:38 | 0 | -1 | -1 | -1 | 1206053 | 180 | 341 | 27.28 | 25.1 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6.7e-05 | -1.3386 | 9.7e-05 | -9.522e-10 | -4.41e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.003 | 107.2288 | 0x43EA | pH | 2021-03-09 08:39 | 0 | -1 | -1 | -1 | 1206054 | 180 | 341 | 27.28 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000103 | -1.3382 | 9.9e-05 | 1.74e-09 | 6.782e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.001 | 107.2261 | 0x9B3E | pH | 2021-03-09 08:40 | 0 | -1 | -1 | -1 | 1206055 | 180 | 341 | 27.27 | 25.08 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000134 | -1.3382 | 0.000228 | -2.194e-09 | -2.025e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.001 | 107.2235 | 0xB9B3 | pH | 2021-03-09 08:41 | 0 | -1 | -1 | -1 | 1206056 | 180 | 341 | 27.27 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000135 | -1.3377 | 8.8e-05 | 3.823e-09 | 1.098e-08 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.002 | 107.2211 | 0x686C | pH | 2021-03-09 08:42 | 0 | -1 | -1 | -1 | 1206057 | 180 | 341 | 27.27 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.9e-05 | -1.3376 | 0.000179 | 5.945e-09 | -1.107e-08 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.002 | 107.219 | 0x8A85 | pH | 2021-03-09 08:43 | 0 | -1 | -1 | -1 | 1206058 | 180 | 341 | 27.26 | 25.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000123 | -1.3372 | 0.000165 | -1.034e-09 | -1.206e-08 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:35 | 20 | 20 | 20 | 36.6255 | 0.25 | 0.001 | 107.2171 | 0x7373 | pH | 2021-03-09 08:44 | 0 | -1 | -1 | -1 | 1206059 | 180 | 341 | 27.27 | 25.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8e-05 | -1.3368 | 0.000178 | -2.31e-09 | 4.755e-09 | 0.073251 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0.002 | 107.2151 | 0x8D0F | pH | 2021-03-09 08:45 | 0 | -1 | -1 | -1 | 1206060 | 180 | 341 | 27.27 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9e-05 | -1.3365 | 0.000134 | 2.231e-09 | -2.391e-11 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0.001 | 107.2136 | 0x45E8 | pH | 2021-03-09 08:46 | 0 | -1 | -1 | -1 | 1206061 | 180 | 341 | 27.26 | 25.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000137 | -1.3362 | 0.000153 | -1.781e-09 | -8.362e-09 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0 | 107.2121 | 0xEB21 | pH | 2021-03-09 08:47 | 0 | -1 | -1 | -1 | 1206062 | 180 | 341 | 27.24 | 25.05 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000136 | -1.3361 | 9.7e-05 | -5.875e-09 | 6.384e-09 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0.002 | 107.2105 | 0x4880 | pH | 2021-03-09 08:48 | 0 | -1 | -1 | -1 | 1206063 | 180 | 341 | 27.26 | 25.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 6.9e-05 | -1.3356 | 0.000134 | 2.155e-09 | -1.754e-08 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0.001 | 107.2092 | 0xB346 | pH | 2021-03-09 08:49 | 0 | -1 | -1 | -1 | 1206064 | 180 | 341 | 27.24 | 25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000115 | -1.3351 | 0.00016 | -2.477e-09 | -7.205e-09 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0 | 107.208 | 0xC461 | pH | 2021-03-09 08:50 | 0 | -1 | -1 | -1 | 1206065 | 180 | 341 | 27.23 | 25.03 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.1e-05 | -1.3347 | 0.000154 | -1.074e-09 | 2.428e-08 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0 | 107.2068 | 0x9AD7 | pH | 2021-03-09 08:51 | 0 | -1 | -1 | -1 | 1206066 | 180 | 341 | 27.23 | 24.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000136 | -1.3346 | 0.00019 | -1.764e-09 | 4.663e-09 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0 | 107.2056 | 0x31B9 | pH | 2021-03-09 08:52 | 0 | -1 | -1 | -1 | 1206067 | 180 | 341 | 27.23 | 25 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000105 | -1.3342 | 0.000143 | -1.967e-09 | -4.971e-09 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0.002 | 107.2046 | 0x7AA3 | pH | 2021-03-09 08:53 | 0 | -1 | -1 | -1 | 1206068 | 180 | 341 | 27.22 | 24.98 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000178 | -1.3336 | 0.000161 | -3.868e-09 | -1.1e-09 | 0.073113 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0.001 | 107.2036 | 0xF394 | pH | 2021-03-09 08:54 | 0 | -1 | -1 | -1 | 1206069 | 180 | 341 | 27.22 | 25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000132 | -1.3334 | 0.000212 | -1.206e-08 | -6.912e-09 | 0.073113 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:36 | 20 | 20 | 20 | 36.5565 | 0.25 | 0 | 107.2027 | 0xB43E | pH | 2021-03-09 08:55 | 0 | -1 | -1 | -1 | 1206070 | 180 | 341 | 27.22 | 24.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.7e-05 | -1.333 | 0.000129 | -2.258e-09 | -8.648e-09 | 0.073113 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.5565 | 0.25 | 0 | 107.2019 | 0xB997 | pH | 2021-03-09 08:56 | 0 | -1 | -1 | -1 | 1206071 | 180 | 341 | 27.2 | 24.92 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000104 | -1.3327 | 0.000136 | -1.377e-08 | 8.663e-09 | 0.073113 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.2011 | 0xFABB | pH | 2021-03-09 08:57 | 0 | -1 | -1 | -1 | 1206072 | 180 | 341 | 27.19 | 24.92 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.00013 | -1.3322 | 0.00013 | -1.458e-09 | 8.257e-10 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.2004 | 0xCF28 | pH | 2021-03-09 08:58 | 0 | -1 | -1 | -1 | 1206073 | 180 | 341 | 27.19 | 24.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000104 | -1.3324 | 0.000129 | -7.919e-09 | -8.548e-09 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.1996 | 0x7F12 | pH | 2021-03-09 08:59 | 0 | -1 | -1 | -1 | 1206074 | 180 | 341 | 27.18 | 24.92 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.8e-05 | -1.3318 | 0.000157 | -8.564e-09 | -5.775e-09 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.1991 | 0x3D50 | pH | 2021-03-09 09:00 | 0 | -1 | -1 | -1 | 1206075 | 180 | 341 | 27.18 | 24.87 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.0001 | -1.3315 | 9.3e-05 | 1.909e-09 | -1.88e-09 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.1983 | 0x7559 | pH | 2021-03-09 09:01 | 0 | -1 | -1 | -1 | 1206076 | 180 | 341 | 27.18 | 24.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000162 | -1.3308 | 0.000152 | 1.099e-09 | 4.705e-11 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.1977 | 0x6E16 | pH | 2021-03-09 09:02 | 0 | -1 | -1 | -1 | 1206077 | 180 | 341 | 27.18 | 24.8 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 7.9e-05 | -1.3304 | 0.000185 | 5.398e-09 | -5.286e-09 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.1971 | 0x68E1 | pH | 2021-03-09 09:03 | 0 | -1 | -1 | -1 | 1206078 | 180 | 341 | 27.16 | 24.83 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.8e-05 | -1.3301 | 0.000173 | 4.779e-09 | -5.798e-09 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4876 | 0.25 | 0 | 107.1966 | 0xB847 | pH | 2021-03-09 09:04 | 0 | -1 | -1 | -1 | 1206079 | 180 | 341 | 27.15 | 24.78 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000149 | -1.3296 | 0.000132 | -3.762e-09 | 3.986e-09 | 0.072975 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:37 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.233 | 107.3718 | 0xFD7F | pH | 2021-03-09 09:05 | 0 | -1 | -1 | -1 | 1206080 | 180 | 341 | 27.15 | 24.75 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.4e-05 | -1.3285 | 0.00012 | -2.889e-10 | -8.469e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.234 | 107.6098 | 0x1D6F | pH | 2021-03-09 09:06 | 0 | -1 | -1 | -1 | 1206081 | 180 | 341 | 27.15 | 24.7 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6e-05 | -1.3273 | 0.000141 | 3.833e-10 | -1.085e-08 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.234 | 107.8481 | 0x74E4 | pH | 2021-03-09 09:07 | 0 | -1 | -1 | -1 | 1206082 | 180 | 341 | 27.15 | 24.72 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5.5e-05 | -1.3259 | 0.000159 | 2.339e-09 | 7.736e-10 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.235 | 108.0865 | 0x7B5E | pH | 2021-03-09 09:08 | 0 | -1 | -1 | -1 | 1206083 | 180 | 341 | 27.14 | 24.7 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000142 | -1.3244 | 0.00021 | -9.119e-09 | -8.99e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.235 | 108.3253 | 0xCE4C | pH | 2021-03-09 09:09 | 0 | -1 | -1 | -1 | 1206084 | 180 | 341 | 27.12 | 24.67 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000115 | -1.3231 | 0.000217 | 5.848e-09 | -9.877e-09 | 0.072837 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.235 | 108.5645 | 0xA7E0 | pH | 2021-03-09 09:10 | 0 | -1 | -1 | -1 | 1206085 | 180 | 341 | 27.12 | 24.72 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000106 | -1.3216 | 0.000124 | 4.956e-10 | 2.014e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.235 | 108.8035 | 0x190A | pH | 2021-03-09 09:11 | 0 | -1 | -1 | -1 | 1206086 | 180 | 341 | 27.12 | 24.6 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000117 | -1.32 | 0.000104 | 4.577e-09 | 4.113e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.235 | 109.0428 | 0x0F78 | pH | 2021-03-09 09:12 | 0 | -1 | -1 | -1 | 1206087 | 180 | 341 | 27.12 | 24.65 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.5e-05 | -1.3184 | 0.000133 | 8.911e-10 | 7.17e-10 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.236 | 109.2827 | 0x5A2E | pH | 2021-03-09 09:13 | 0 | -1 | -1 | -1 | 1206088 | 180 | 341 | 27.11 | 24.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000118 | -1.3164 | 0.000212 | -4.574e-09 | 2.435e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.235 | 109.5225 | 0x5814 | pH | 2021-03-09 09:14 | 0 | -1 | -1 | -1 | 1206089 | 180 | 341 | 27.12 | 24.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 8.9e-05 | -1.3143 | 0.000121 | -9.237e-10 | 4.622e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.236 | 109.7628 | 0x64ED | pH | 2021-03-09 09:15 | 0 | -1 | -1 | -1 | 1206090 | 180 | 341 | 27.11 | 24.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000133 | -1.3113 | 0.000191 | -3.792e-10 | 1.418e-08 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:38 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.237 | 110.0037 | 0x61F2 | pH | 2021-03-09 09:16 | 0 | -1 | -1 | -1 | 1206091 | 180 | 341 | 27.11 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6.6e-05 | -1.3068 | 0.000213 | -9.098e-10 | -8.953e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.237 | 110.2454 | 0xB7A8 | pH | 2021-03-09 09:17 | 0 | -1 | -1 | -1 | 1206092 | 180 | 341 | 27.11 | 24.65 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.3e-05 | -1.3011 | 0.000174 | 3.007e-09 | 1.731e-09 | 0.072837 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.239 | 110.4886 | 0x44EA | pH | 2021-03-09 09:18 | 0 | -1 | -1 | -1 | 1206093 | 180 | 341 | 27.11 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000129 | -1.2957 | 0.000197 | -2.557e-09 | -2.149e-09 | 0.072837 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.006 | 110.5596 | 0x9BCD | pH | 2021-03-09 09:19 | 0 | -1 | -1 | -1 | 1206094 | 180 | 341 | 27.11 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000123 | -1.2933 | 0.000139 | -1.368e-08 | 8.417e-09 | 0.072837 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.4186 | 0.25 | -0.002 | 110.5626 | 0x27E4 | pH | 2021-03-09 09:20 | 0 | -1 | -1 | -1 | 1206095 | 180 | 341 | 27.11 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.7e-05 | -1.2897 | 0.00011 | -9.758e-10 | -7.894e-10 | 0.072837 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.3496 | 0.25 | -0.002 | 110.5646 | 0xBAA6 | pH | 2021-03-09 09:22 | 0 | -1 | -1 | -1 | 1206096 | 180 | 341 | 27.11 | 24.62 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9e-05 | -1.286 | 0.000138 | -8.284e-09 | -7.492e-09 | 0.072699 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.566 | 0x8AFA | pH | 2021-03-09 09:23 | 0 | -1 | -1 | -1 | 1206097 | 180 | 341 | 27.11 | 24.65 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.2e-05 | -1.2832 | 0.000156 | 1.126e-09 | 2.264e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.3496 | 0.25 | -0.004 | 110.5674 | 0x9481 | pH | 2021-03-09 09:24 | 0 | -1 | -1 | -1 | 1206098 | 180 | 341 | 27.09 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000111 | -1.2813 | 0.000118 | -1.764e-09 | -1.112e-08 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5682 | 0xE1BF | pH | 2021-03-09 09:25 | 0 | -1 | -1 | -1 | 1206099 | 180 | 341 | 27.11 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000152 | -1.279 | 0.000152 | 6.924e-09 | 7.682e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5691 | 0x1CA4 | pH | 2021-03-09 09:26 | 0 | -1 | -1 | -1 | 1206100 | 180 | 341 | 27.12 | 24.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000113 | -1.2759 | 0.000148 | -2.24e-09 | -3.39e-10 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:39 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.57 | 0x744C | pH | 2021-03-09 09:27 | 0 | -1 | -1 | -1 | 1206101 | 180 | 341 | 27.11 | 24.6 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000145 | -1.2739 | 0.000193 | -2.69e-09 | -3.351e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5707 | 0x3F47 | pH | 2021-03-09 09:28 | 0 | -1 | -1 | -1 | 1206102 | 180 | 341 | 27.12 | 24.6 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9e-05 | -1.2718 | 0.000189 | -4.715e-09 | -5.416e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | -0.002 | 110.5715 | 0x8CF0 | pH | 2021-03-09 09:29 | 0 | -1 | -1 | -1 | 1206103 | 180 | 341 | 27.12 | 24.6 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000122 | -1.2702 | 0.000132 | 1.349e-09 | 1.026e-08 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | -0.004 | 110.5722 | 0xAF56 | pH | 2021-03-09 09:30 | 0 | -1 | -1 | -1 | 1206104 | 180 | 341 | 27.14 | 24.58 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.9e-05 | -1.2684 | 0.000202 | 6.136e-10 | -1.259e-08 | 0.072699 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5727 | 0xE737 | pH | 2021-03-09 09:31 | 0 | -1 | -1 | -1 | 1206105 | 180 | 341 | 27.12 | 24.58 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000134 | -1.2671 | 0.000172 | -2.604e-10 | 8.894e-09 | 0.072699 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | -0.001 | 110.5732 | 0x2A99 | pH | 2021-03-09 09:32 | 0 | -1 | -1 | -1 | 1206106 | 180 | 341 | 27.14 | 24.58 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.9e-05 | -1.2651 | 0.000102 | -5.706e-09 | 1.288e-08 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5737 | 0xBC84 | pH | 2021-03-09 09:33 | 0 | -1 | -1 | -1 | 1206107 | 180 | 341 | 27.14 | 24.55 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.5e-05 | -1.264 | 0.000187 | -2.946e-10 | 8.575e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5742 | 0xC2AD | pH | 2021-03-09 09:34 | 0 | -1 | -1 | -1 | 1206108 | 180 | 341 | 27.14 | 24.53 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 6.8e-05 | -1.2628 | 0.000148 | -2.892e-09 | 9.761e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5746 | 0xB89D | pH | 2021-03-09 09:35 | 0 | -1 | -1 | -1 | 1206109 | 180 | 341 | 27.16 | 24.5 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000103 | -1.2619 | 0.000104 | -1.209e-09 | 5.013e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5749 | 0x0ABE | pH | 2021-03-09 09:36 | 0 | -1 | -1 | -1 | 1206110 | 180 | 341 | 27.15 | 24.53 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.9e-05 | -1.2614 | 0.000129 | 6.567e-09 | -3.279e-09 | 0.072699 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:40 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5754 | 0x7FA5 | pH | 2021-03-09 09:37 | 0 | -1 | -1 | -1 | 1206111 | 180 | 341 | 27.16 | 24.5 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000121 | -1.2608 | 0.000143 | -5.192e-09 | 7.169e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5758 | 0x2792 | pH | 2021-03-09 09:38 | 0 | -1 | -1 | -1 | 1206112 | 180 | 341 | 27.16 | 24.5 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.7e-05 | -1.26 | 0.000171 | -7.697e-09 | 2.145e-09 | 0.072699 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5763 | 0x7264 | pH | 2021-03-09 09:39 | 0 | -1 | -1 | -1 | 1206113 | 180 | 341 | 27.18 | 24.5 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000158 | -1.2589 | 0.000175 | -8.507e-09 | 3.427e-10 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5768 | 0x0AF6 | pH | 2021-03-09 09:40 | 0 | -1 | -1 | -1 | 1206114 | 180 | 341 | 27.19 | 24.47 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000107 | -1.2586 | 0.000241 | -7.6e-09 | 8.304e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.577 | 0x5E66 | pH | 2021-03-09 09:41 | 0 | -1 | -1 | -1 | 1206115 | 180 | 341 | 27.19 | 24.42 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000102 | -1.2581 | 0.000107 | 2.757e-09 | -1.566e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.3496 | 0.25 | 0 | 110.5775 | 0x9988 | pH | 2021-03-09 09:42 | 0 | -1 | -1 | -1 | 1206116 | 180 | 341 | 27.19 | 24.45 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2072 | 6.5e-05 | -1.2572 | 0.000139 | 2.879e-09 | -3.546e-09 | 0.072699 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5777 | 0xE0A9 | pH | 2021-03-09 09:43 | 0 | -1 | -1 | -1 | 1206117 | 180 | 341 | 27.2 | 24.45 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000134 | -1.2568 | 0.000175 | -7.482e-09 | 8.183e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.578 | 0xE656 | pH | 2021-03-09 09:44 | 0 | -1 | -1 | -1 | 1206118 | 180 | 341 | 27.2 | 24.42 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8e-05 | -1.2561 | 0.000214 | -1.094e-08 | 8.436e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5785 | 0xBA97 | pH | 2021-03-09 09:45 | 0 | -1 | -1 | -1 | 1206119 | 180 | 341 | 27.22 | 24.4 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.9e-05 | -1.2554 | 0.000128 | -2.771e-09 | 1.335e-08 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5787 | 0xDD6B | pH | 2021-03-09 09:46 | 0 | -1 | -1 | -1 | 1206120 | 180 | 341 | 27.2 | 24.42 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.3e-05 | -1.2553 | 0.000119 | -4.029e-09 | 1.481e-08 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5791 | 0x7838 | pH | 2021-03-09 09:47 | 0 | -1 | -1 | -1 | 1206121 | 180 | 341 | 27.22 | 24.37 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000102 | -1.255 | 0.000132 | 3.776e-10 | 2.885e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:41 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5794 | 0xFA0A | pH | 2021-03-09 09:48 | 0 | -1 | -1 | -1 | 1206122 | 180 | 341 | 27.23 | 24.37 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.8e-05 | -1.2547 | 0.000127 | 1.802e-09 | 4.776e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5798 | 0x41B5 | pH | 2021-03-09 09:49 | 0 | -1 | -1 | -1 | 1206123 | 180 | 341 | 27.23 | 24.37 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 6.8e-05 | -1.2544 | 0.000136 | 9.578e-10 | -1.382e-10 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 20 | 36.2806 | 0.25 | 0 | 110.5799 | 0xD0A0 | pH | 2021-03-09 09:50 | 0 | -1 | -1 | -1 | 1206124 | 180 | 341 | 27.23 | 24.35 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 5.4e-05 | -1.2558 | 0.000108 | 4.613e-09 | 6.933e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | -0.001 | 110.5805 | 0x23D9 | pH | 2021-03-09 09:51 | 0 | -1 | -1 | -1 | 1206125 | 180 | 341 | 27.23 | 24.37 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000109 | -1.2652 | 0.000167 | -1.899e-09 | 8.079e-10 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | -0.007 | 110.5819 | 0x8074 | pH | 2021-03-09 09:52 | 0 | -1 | -1 | -1 | 1206126 | 180 | 341 | 27.24 | 24.35 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.3e-05 | -1.2646 | 0.000101 | -8.009e-09 | 2.448e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | -0.002 | 110.5825 | 0xE161 | pH | 2021-03-09 09:53 | 0 | -1 | -1 | -1 | 1206127 | 180 | 341 | 27.24 | 24.33 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000177 | -1.257 | 0.000135 | -5.731e-09 | -2.942e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | -0.002 | 110.5833 | 0xD1DF | pH | 2021-03-09 09:54 | 0 | -1 | -1 | -1 | 1206128 | 180 | 341 | 27.26 | 24.33 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000128 | -1.2541 | 0.000152 | 2.672e-09 | -4.226e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5836 | 0xA579 | pH | 2021-03-09 09:55 | 0 | -1 | -1 | -1 | 1206129 | 180 | 341 | 27.26 | 24.33 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.4e-05 | -1.2528 | 0.000102 | -8.813e-09 | 1.186e-08 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5838 | 0x54FC | pH | 2021-03-09 09:56 | 0 | -1 | -1 | -1 | 1206130 | 180 | 341 | 27.27 | 24.35 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000115 | -1.2519 | 6.6e-05 | 7.471e-10 | -4.139e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5838 | 0xAEB4 | pH | 2021-03-09 09:57 | 0 | -1 | -1 | -1 | 1206131 | 180 | 341 | 27.27 | 24.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.4e-05 | -1.2515 | 0.000177 | -8.449e-09 | -1.8e-08 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:42 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5838 | 0xFAEA | pH | 2021-03-09 09:58 | 0 | -1 | -1 | -1 | 1206132 | 180 | 341 | 27.27 | 24.33 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000113 | -1.2517 | 0.000124 | -6.614e-09 | -8e-10 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5838 | 0x896A | pH | 2021-03-09 09:59 | 0 | -1 | -1 | -1 | 1206133 | 180 | 341 | 27.28 | 24.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.1e-05 | -1.2512 | 0.000129 | 1.91e-09 | 1.514e-08 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5838 | 0x09EA | pH | 2021-03-09 10:00 | 0 | -1 | -1 | -1 | 1206134 | 180 | 341 | 27.28 | 24.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000156 | -1.2513 | 0.000178 | -1.995e-09 | -1.991e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5819 | 0x4308 | pH | 2021-03-09 10:01 | 0 | -1 | -1 | -1 | 1206135 | 180 | 341 | 27.3 | 24.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000132 | -1.2512 | 0.000102 | -3.496e-09 | 2.271e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.002 | 110.58 | 0xF61D | pH | 2021-03-09 10:02 | 0 | -1 | -1 | -1 | 1206136 | 180 | 341 | 27.3 | 24.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000152 | -1.2519 | 0.000137 | -9.014e-09 | -7.657e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5782 | 0x9B6E | pH | 2021-03-09 10:03 | 0 | -1 | -1 | -1 | 1206137 | 180 | 341 | 27.3 | 24.25 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000137 | -1.2524 | 0.000199 | -1.278e-09 | 7.768e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5764 | 0xB058 | pH | 2021-03-09 10:04 | 0 | -1 | -1 | -1 | 1206138 | 180 | 341 | 27.3 | 24.22 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 4e-05 | -1.2521 | 0.00015 | -3.687e-10 | 8.091e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5745 | 0x253E | pH | 2021-03-09 10:05 | 0 | -1 | -1 | -1 | 1206139 | 180 | 341 | 27.31 | 24.22 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.4e-05 | -1.2531 | 0.000156 | 6.921e-09 | 3.383e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5725 | 0xF39F | pH | 2021-03-09 10:06 | 0 | -1 | -1 | -1 | 1206140 | 180 | 341 | 27.32 | 24.2 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000109 | -1.2552 | 0.000145 | -7.245e-09 | 6.494e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.002 | 110.5704 | 0x129F | pH | 2021-03-09 10:07 | 0 | -1 | -1 | -1 | 1206141 | 180 | 341 | 27.31 | 24.22 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 6.6e-05 | -1.2557 | 0.000163 | -3.223e-10 | 4.226e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:43 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5686 | 0xC6DE | pH | 2021-03-09 10:08 | 0 | -1 | -1 | -1 | 1206142 | 180 | 341 | 27.32 | 24.22 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.4e-05 | -1.2554 | 0.000105 | -2.453e-09 | 8.769e-11 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5668 | 0xF922 | pH | 2021-03-09 10:09 | 0 | -1 | -1 | -1 | 1206143 | 180 | 341 | 27.32 | 24.2 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000109 | -1.2559 | 0.000112 | -9.032e-09 | -5.751e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.002 | 110.5648 | 0x0956 | pH | 2021-03-09 10:10 | 0 | -1 | -1 | -1 | 1206144 | 180 | 341 | 27.34 | 24.17 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000149 | -1.2564 | 0.00012 | -3.126e-09 | -2.409e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.563 | 0x12BC | pH | 2021-03-09 10:11 | 0 | -1 | -1 | -1 | 1206145 | 180 | 341 | 27.34 | 24.15 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000102 | -1.2567 | 0.000104 | -6.453e-09 | 4.531e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5611 | 0x335A | pH | 2021-03-09 10:12 | 0 | -1 | -1 | -1 | 1206146 | 180 | 341 | 27.34 | 24.17 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.2e-05 | -1.2575 | 0.000195 | 4.079e-09 | 9.418e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5594 | 0xA5F6 | pH | 2021-03-09 10:13 | 0 | -1 | -1 | -1 | 1206147 | 180 | 341 | 27.34 | 24.15 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000111 | -1.2586 | 0.000172 | -6.339e-09 | 2.341e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5578 | 0x4B94 | pH | 2021-03-09 10:14 | 0 | -1 | -1 | -1 | 1206148 | 180 | 341 | 27.35 | 24.15 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 5.6e-05 | -1.259 | 0.000216 | 6.925e-09 | 1.652e-08 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5561 | 0x6016 | pH | 2021-03-09 10:15 | 0 | -1 | -1 | -1 | 1206149 | 180 | 341 | 27.34 | 24.12 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.2e-05 | -1.2593 | 9.4e-05 | 3.809e-09 | -1.542e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5547 | 0x249E | pH | 2021-03-09 10:16 | 0 | -1 | -1 | -1 | 1206150 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.00015 | -1.2599 | 0.000227 | -4.184e-09 | 1.095e-08 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.002 | 110.5531 | 0xBF8C | pH | 2021-03-09 10:17 | 0 | -1 | -1 | -1 | 1206151 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.3e-05 | -1.2602 | 0.000126 | 5.042e-09 | 1.152e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.001 | 110.5518 | 0x95B2 | pH | 2021-03-09 10:18 | 0 | -1 | -1 | -1 | 1206152 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000171 | -1.2596 | 0.000129 | -2.747e-09 | -6.149e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:44 | 20 | 20 | 5 | 36.2806 | 0.25 | 0 | 110.5507 | 0xC185 | pH | 2021-03-09 10:19 | 0 | -1 | -1 | -1 | 1206153 | 180 | 341 | 27.34 | 24.09 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000105 | -1.2598 | 0.000187 | -1.953e-11 | 1.552e-08 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.239 | 110.371 | 0x0BC5 | pH | 2021-03-09 10:20 | 0 | -1 | -1 | -1 | 1206154 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000147 | -1.2659 | 0.000139 | -1.473e-10 | 1.486e-08 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.239 | 110.1277 | 0xD126 | pH | 2021-03-09 10:21 | 0 | -1 | -1 | -1 | 1206155 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.8e-05 | -1.2711 | 0.000125 | 3.988e-11 | -4.427e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.237 | 109.8851 | 0x0999 | pH | 2021-03-09 10:22 | 0 | -1 | -1 | -1 | 1206156 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.7e-05 | -1.2757 | 0.000117 | 1.899e-09 | 8.016e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.237 | 109.6435 | 0x9D55 | pH | 2021-03-09 10:24 | 0 | -1 | -1 | -1 | 1206157 | 180 | 341 | 27.35 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000115 | -1.279 | 0.000159 | -5.992e-09 | 3.9e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.236 | 109.4026 | 0x9F57 | pH | 2021-03-09 10:25 | 0 | -1 | -1 | -1 | 1206158 | 180 | 341 | 27.34 | 24.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8e-05 | -1.2814 | 0.000124 | -4.749e-09 | 1.901e-08 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.236 | 109.1624 | 0x8626 | pH | 2021-03-09 10:26 | 0 | -1 | -1 | -1 | 1206159 | 180 | 341 | 27.34 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.00013 | -1.2843 | 0.000179 | -9.359e-09 | 5.081e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.236 | 108.9216 | 0x212E | pH | 2021-03-09 10:27 | 0 | -1 | -1 | -1 | 1206160 | 180 | 341 | 27.34 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 4.4e-05 | -1.2878 | 9.2e-05 | 6.836e-11 | -2.64e-10 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.235 | 108.6822 | 0xA498 | pH | 2021-03-09 10:28 | 0 | -1 | -1 | -1 | 1206161 | 180 | 341 | 27.32 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000145 | -1.292 | 0.000172 | -2.441e-11 | -1.757e-09 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.235 | 108.4424 | 0xC7AD | pH | 2021-03-09 10:29 | 0 | -1 | -1 | -1 | 1206162 | 180 | 341 | 27.32 | 24.07 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.3e-05 | -1.2958 | 0.000195 | 1.549e-09 | 5.537e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:45 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.235 | 108.2027 | 0x3D9E | pH | 2021-03-09 10:30 | 0 | -1 | -1 | -1 | 1206163 | 180 | 341 | 27.31 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.6e-05 | -1.2972 | 0.000112 | -1.787e-09 | -4.944e-10 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.234 | 107.9631 | 0xD2FD | pH | 2021-03-09 10:31 | 0 | -1 | -1 | -1 | 1206164 | 180 | 341 | 27.31 | 24.04 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.1e-05 | -1.2988 | 0.000135 | 6.359e-09 | 6.077e-09 | 0.072423 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.235 | 107.7237 | 0x311A | pH | 2021-03-09 10:32 | 0 | -1 | -1 | -1 | 1206165 | 180 | 341 | 27.3 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000112 | -1.3018 | 0.00028 | -1.216e-08 | -1.577e-09 | 0.072423 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.235 | 107.4847 | 0x1E97 | pH | 2021-03-09 10:33 | 0 | -1 | -1 | -1 | 1206166 | 180 | 341 | 27.3 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.3e-05 | -1.3028 | 0.000146 | 7.205e-09 | -1.38e-08 | 0.072561 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.235 | 107.2454 | 0xBE9E | pH | 2021-03-09 10:34 | 0 | -1 | -1 | -1 | 1206167 | 180 | 341 | 27.3 | 24.02 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000118 | -1.3038 | 0.000197 | -1.535e-09 | -7.285e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2806 | 0.25 | 0.235 | 107.0063 | 0xE1D6 | pH | 2021-03-09 10:35 | 0 | -1 | -1 | -1 | 1206168 | 180 | 341 | 27.28 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000105 | -1.3044 | 0.000156 | 3.264e-09 | 1.054e-09 | 0.072561 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.004 | 106.9842 | 0xFD63 | pH | 2021-03-09 10:36 | 0 | -1 | -1 | -1 | 1206169 | 180 | 341 | 27.3 | 24.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.0001 | -1.3048 | 0.000199 | 1.326e-09 | 4.05e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.003 | 106.9803 | 0x3983 | pH | 2021-03-09 10:37 | 0 | -1 | -1 | -1 | 1206170 | 180 | 341 | 27.28 | 24.02 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000102 | -1.3057 | 0.000102 | -5.909e-09 | 5.318e-09 | 0.072423 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.001 | 106.9776 | 0xE043 | pH | 2021-03-09 10:38 | 0 | -1 | -1 | -1 | 1206171 | 180 | 341 | 27.28 | 24.02 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000146 | -1.3073 | 0.000194 | 3.766e-09 | 3.837e-09 | 0.072423 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.002 | 106.9748 | 0x6CC4 | pH | 2021-03-09 10:39 | 0 | -1 | -1 | -1 | 1206172 | 180 | 341 | 27.27 | 24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.3e-05 | -1.3081 | 0.000206 | -2.651e-09 | 9.202e-09 | 0.072423 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:46 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.002 | 106.9726 | 0xD5D3 | pH | 2021-03-09 10:40 | 0 | -1 | -1 | -1 | 1206173 | 180 | 341 | 27.27 | 24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000108 | -1.3088 | 0.000173 | 2.977e-09 | -1.076e-08 | 0.072423 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.001 | 106.9707 | 0x746D | pH | 2021-03-09 10:41 | 0 | -1 | -1 | -1 | 1206174 | 180 | 341 | 27.27 | 24.02 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.9e-05 | -1.3094 | 0.00018 | 1.108e-09 | -4.922e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.002 | 106.9687 | 0x130A | pH | 2021-03-09 10:42 | 0 | -1 | -1 | -1 | 1206175 | 180 | 341 | 27.26 | 24 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000139 | -1.31 | 0.000213 | -7.165e-09 | 1.121e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.002 | 106.9671 | 0x206A | pH | 2021-03-09 10:43 | 0 | -1 | -1 | -1 | 1206176 | 180 | 341 | 27.26 | 23.97 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000132 | -1.3104 | 0.000247 | 2.909e-09 | 2.047e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0 | 106.9657 | 0xFBB7 | pH | 2021-03-09 10:44 | 0 | -1 | -1 | -1 | 1206177 | 180 | 341 | 27.26 | 23.97 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.1e-05 | -1.3111 | 0.000163 | -7.488e-09 | 4.386e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0 | 106.9643 | 0x52E8 | pH | 2021-03-09 10:45 | 0 | -1 | -1 | -1 | 1206178 | 180 | 341 | 27.24 | 23.97 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.1e-05 | -1.3113 | 0.000179 | 8.451e-09 | -6.593e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0 | 106.963 | 0x743C | pH | 2021-03-09 10:46 | 0 | -1 | -1 | -1 | 1206179 | 180 | 341 | 27.24 | 24 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000114 | -1.3116 | 0.000133 | 5.396e-10 | 1.22e-08 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9618 | 0x5214 | pH | 2021-03-09 10:47 | 0 | -1 | -1 | -1 | 1206180 | 180 | 341 | 27.24 | 24 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000131 | -1.3118 | 0.000172 | -2.615e-09 | -1.416e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0 | 106.9607 | 0x0683 | pH | 2021-03-09 10:48 | 0 | -1 | -1 | -1 | 1206181 | 180 | 341 | 27.23 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000112 | -1.3121 | 0.000167 | -1.095e-08 | -4.952e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.002 | 106.9597 | 0x3137 | pH | 2021-03-09 10:49 | 0 | -1 | -1 | -1 | 1206182 | 180 | 341 | 27.23 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.3e-05 | -1.3122 | 0.000133 | -8.634e-10 | -5.845e-10 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0.001 | 106.9588 | 0x8ECD | pH | 2021-03-09 10:50 | 0 | -1 | -1 | -1 | 1206183 | 180 | 341 | 27.24 | 23.92 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000115 | -1.3122 | 0.000112 | -6.082e-09 | -3.669e-10 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:47 | 20 | 20 | 5 | 36.2117 | 0.25 | 0 | 106.9579 | 0xDF8C | pH | 2021-03-09 10:51 | 0 | -1 | -1 | -1 | 1206184 | 180 | 341 | 27.24 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000162 | -1.3122 | 0.000151 | -2.365e-09 | 2.599e-09 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9571 | 0xD82D | pH | 2021-03-09 10:52 | 0 | -1 | -1 | -1 | 1206185 | 180 | 341 | 27.23 | 23.97 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.00013 | -1.3121 | 9e-05 | -5.588e-09 | -1.184e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.2117 | 0.25 | 0 | 106.9564 | 0x5B9A | pH | 2021-03-09 10:53 | 0 | -1 | -1 | -1 | 1206186 | 180 | 341 | 27.22 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000164 | -1.3115 | 0.000169 | -4.285e-09 | 4.786e-11 | 0.072423 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9558 | 0xCC40 | pH | 2021-03-09 10:54 | 0 | -1 | -1 | -1 | 1206187 | 180 | 341 | 27.23 | 23.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 6e-05 | -1.3109 | 0.000127 | -9.444e-09 | 2.065e-10 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9551 | 0x6EF1 | pH | 2021-03-09 10:55 | 0 | -1 | -1 | -1 | 1206188 | 180 | 341 | 27.22 | 23.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000109 | -1.3108 | 0.000151 | 3.928e-09 | -7.24e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9545 | 0x5F17 | pH | 2021-03-09 10:56 | 0 | -1 | -1 | -1 | 1206189 | 180 | 341 | 27.22 | 23.97 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8e-05 | -1.311 | 0.000176 | 2.941e-09 | -1.063e-08 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9539 | 0x4B3A | pH | 2021-03-09 10:57 | 0 | -1 | -1 | -1 | 1206190 | 180 | 341 | 27.22 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.1e-05 | -1.3111 | 0.00015 | -3.394e-10 | -8.373e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0.001 | 106.9532 | 0x25EE | pH | 2021-03-09 10:58 | 0 | -1 | -1 | -1 | 1206191 | 180 | 341 | 27.22 | 23.92 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.2e-05 | -1.3112 | 9.2e-05 | -6.165e-09 | 4.284e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9527 | 0x2139 | pH | 2021-03-09 10:59 | 0 | -1 | -1 | -1 | 1206192 | 180 | 341 | 27.22 | 23.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000127 | -1.3113 | 0.000136 | -3.12e-09 | -5.629e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9522 | 0x6CB4 | pH | 2021-03-09 11:00 | 0 | -1 | -1 | -1 | 1206193 | 180 | 341 | 27.2 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000128 | -1.3115 | 0.000164 | -9.44e-10 | -2.812e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:48 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9519 | 0x5B22 | pH | 2021-03-09 11:01 | 0 | -1 | -1 | -1 | 1206194 | 180 | 341 | 27.2 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.2e-05 | -1.3115 | 0.000152 | 6.743e-09 | 8.157e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9514 | 0xF9F6 | pH | 2021-03-09 11:02 | 0 | -1 | -1 | -1 | 1206195 | 180 | 341 | 27.2 | 23.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000123 | -1.3113 | 0.000183 | 2.536e-09 | 4.928e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | 0.001 | 106.951 | 0x0D7F | pH | 2021-03-09 11:03 | 0 | -1 | -1 | -1 | 1206196 | 180 | 341 | 27.2 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000157 | -1.311 | 0.00015 | 1.032e-08 | -5.15e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | 0 | 106.9505 | 0x6958 | pH | 2021-03-09 11:04 | 0 | -1 | -1 | -1 | 1206197 | 180 | 341 | 27.2 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000123 | -1.3109 | 0.000178 | 7.992e-10 | 9.026e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.233 | 107.1252 | 0xFB8E | pH | 2021-03-09 11:05 | 0 | -1 | -1 | -1 | 1206198 | 180 | 341 | 27.2 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000138 | -1.3088 | 0.000141 | -1.491e-09 | 5.877e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.234 | 107.3636 | 0xBAA1 | pH | 2021-03-09 11:06 | 0 | -1 | -1 | -1 | 1206199 | 180 | 341 | 27.2 | 23.95 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.1e-05 | -1.3081 | 0.000141 | 1.392e-09 | 5.183e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.234 | 107.6022 | 0xC412 | pH | 2021-03-09 11:07 | 0 | -1 | -1 | -1 | 1206200 | 180 | 341 | 27.2 | 23.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000104 | -1.3075 | 0.000128 | -7.301e-09 | -4.928e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.234 | 107.8412 | 0xDE22 | pH | 2021-03-09 11:08 | 0 | -1 | -1 | -1 | 1206201 | 180 | 341 | 27.2 | 23.95 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000117 | -1.3071 | 0.000115 | -1.155e-09 | 9.482e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.235 | 108.0802 | 0x6037 | pH | 2021-03-09 11:09 | 0 | -1 | -1 | -1 | 1206202 | 180 | 341 | 27.19 | 23.92 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000114 | -1.3068 | 0.000185 | 2.015e-09 | 6.917e-10 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.236 | 108.3193 | 0xC6B1 | pH | 2021-03-09 11:10 | 0 | -1 | -1 | -1 | 1206203 | 180 | 341 | 27.2 | 23.92 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000115 | -1.3066 | 0.000143 | 3.068e-10 | 4.231e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:49 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.236 | 108.5592 | 0xEFB4 | pH | 2021-03-09 11:11 | 0 | -1 | -1 | -1 | 1206204 | 180 | 341 | 27.19 | 23.92 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.9e-05 | -1.3062 | 0.000188 | 1.024e-08 | 2.214e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.236 | 108.7991 | 0xD94C | pH | 2021-03-09 11:12 | 0 | -1 | -1 | -1 | 1206205 | 180 | 341 | 27.19 | 23.92 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000112 | -1.3052 | 0.000225 | -2.293e-09 | 1.263e-08 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.236 | 109.0389 | 0x8DC4 | pH | 2021-03-09 11:13 | 0 | -1 | -1 | -1 | 1206206 | 180 | 341 | 27.19 | 23.92 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000128 | -1.3041 | 0.000211 | 3.899e-09 | 1.199e-08 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.236 | 109.2794 | 0x1BD3 | pH | 2021-03-09 11:14 | 0 | -1 | -1 | -1 | 1206207 | 180 | 341 | 27.19 | 23.92 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000111 | -1.3026 | 0.000169 | 1.265e-09 | 1.512e-08 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.236 | 109.5203 | 0xEFF3 | pH | 2021-03-09 11:15 | 0 | -1 | -1 | -1 | 1206208 | 180 | 341 | 27.19 | 23.89 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000106 | -1.3007 | 0.000206 | -9.41e-09 | 8.521e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.238 | 109.7619 | 0x6BCA | pH | 2021-03-09 11:16 | 0 | -1 | -1 | -1 | 1206209 | 180 | 341 | 27.19 | 23.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000115 | -1.2979 | 0.000204 | 6.966e-10 | -1.287e-09 | 0.072285 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.239 | 110.0055 | 0x8B44 | pH | 2021-03-09 11:17 | 0 | -1 | -1 | -1 | 1206210 | 180 | 341 | 27.18 | 23.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000112 | -1.2948 | 0.000188 | 2.911e-09 | 4.362e-10 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.1427 | 0.25 | -0.239 | 110.2483 | 0xE387 | pH | 2021-03-09 11:18 | 0 | -1 | -1 | -1 | 1206211 | 180 | 341 | 27.19 | 23.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7e-05 | -1.2915 | 0.000174 | -1.621e-09 | 9.18e-09 | 0.072285 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.189 | 110.4893 | 0x0406 | pH | 2021-03-09 11:19 | 0 | -1 | -1 | -1 | 1206212 | 180 | 341 | 27.18 | 23.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000142 | -1.2903 | 0.000176 | -1.212e-08 | -7.956e-09 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.003 | 110.4966 | 0xAF52 | pH | 2021-03-09 11:20 | 0 | -1 | -1 | -1 | 1206213 | 180 | 341 | 27.16 | 23.87 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 6.2e-05 | -1.2918 | 0.000135 | 4.028e-10 | -1.98e-08 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.002 | 110.4992 | 0x62CE | pH | 2021-03-09 11:21 | 0 | -1 | -1 | -1 | 1206214 | 180 | 341 | 27.18 | 23.84 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000194 | -1.2923 | 0.000203 | -8.44e-09 | 1.154e-09 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:50 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.005 | 110.5012 | 0xF11D | pH | 2021-03-09 11:22 | 0 | -1 | -1 | -1 | 1206215 | 180 | 341 | 27.16 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000137 | -1.2925 | 0.000186 | -1.599e-09 | 4.093e-09 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5019 | 0xD619 | pH | 2021-03-09 11:24 | 0 | -1 | -1 | -1 | 1206216 | 180 | 341 | 27.16 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8e-05 | -1.2929 | 0.000175 | -5.298e-10 | -2.164e-09 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5031 | 0x23EC | pH | 2021-03-09 11:25 | 0 | -1 | -1 | -1 | 1206217 | 180 | 341 | 27.16 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000157 | -1.2927 | 0.000162 | -6.406e-09 | -5.349e-09 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.002 | 110.5043 | 0xF661 | pH | 2021-03-09 11:26 | 0 | -1 | -1 | -1 | 1206218 | 180 | 341 | 27.16 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000113 | -1.2916 | 0.000129 | -4.433e-09 | -5.691e-10 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.001 | 110.5046 | 0x4DE7 | pH | 2021-03-09 11:27 | 0 | -1 | -1 | -1 | 1206219 | 180 | 341 | 27.16 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000144 | -1.2907 | 0.000126 | 4.98e-09 | 3.779e-09 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5053 | 0x3C96 | pH | 2021-03-09 11:28 | 0 | -1 | -1 | -1 | 1206220 | 180 | 341 | 27.15 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000143 | -1.2907 | 0.000147 | -9.79e-10 | 1.468e-08 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.001 | 110.5058 | 0x61DB | pH | 2021-03-09 11:29 | 0 | -1 | -1 | -1 | 1206221 | 180 | 341 | 27.16 | 23.82 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.6e-05 | -1.2889 | 0.00018 | -1.009e-10 | -6.482e-10 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5063 | 0xEDB0 | pH | 2021-03-09 11:30 | 0 | -1 | -1 | -1 | 1206222 | 180 | 341 | 27.15 | 23.79 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.7e-05 | -1.2899 | 0.000143 | -1.073e-08 | -1.21e-08 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5067 | 0x9031 | pH | 2021-03-09 11:31 | 0 | -1 | -1 | -1 | 1206223 | 180 | 341 | 27.15 | 23.79 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000127 | -1.2902 | 9e-05 | 1.538e-10 | 1.26e-08 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5071 | 0x6FA4 | pH | 2021-03-09 11:32 | 0 | -1 | -1 | -1 | 1206224 | 180 | 341 | 27.15 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.00012 | -1.2877 | 0.000115 | 9.969e-10 | 1.239e-08 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:51 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5079 | 0x06B0 | pH | 2021-03-09 11:33 | 0 | -1 | -1 | -1 | 1206225 | 180 | 341 | 27.15 | 23.79 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000131 | -1.2877 | 0.000114 | -3.92e-09 | -3.85e-09 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0737 | 0.25 | -0.002 | 110.5082 | 0xA2D9 | pH | 2021-03-09 11:34 | 0 | -1 | -1 | -1 | 1206226 | 180 | 341 | 27.14 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.5e-05 | -1.2884 | 0.000259 | -1.181e-09 | -1.075e-08 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5086 | 0x4293 | pH | 2021-03-09 11:35 | 0 | -1 | -1 | -1 | 1206227 | 180 | 341 | 27.14 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8e-05 | -1.2882 | 9.4e-05 | -3.91e-09 | -2.124e-09 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5091 | 0x44FF | pH | 2021-03-09 11:36 | 0 | -1 | -1 | -1 | 1206228 | 180 | 341 | 27.14 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.4e-05 | -1.2875 | 0.000141 | -7.632e-09 | -9.745e-09 | 0.072147 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0737 | 0.25 | 0 | 110.5094 | 0x458E | pH | 2021-03-09 11:37 | 0 | -1 | -1 | -1 | 1206229 | 180 | 341 | 27.14 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.2e-05 | -1.2857 | 0.000139 | 2.688e-09 | -4.956e-09 | 0.072147 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | -0.001 | 110.5097 | 0x90BA | pH | 2021-03-09 11:38 | 0 | -1 | -1 | -1 | 1206230 | 180 | 341 | 27.15 | 23.75 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000106 | -1.2841 | 0.000148 | -9.549e-09 | -6.608e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4255 | -1.4255 | -1.4255 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5102 | 0x5FAE | pH | 2021-03-09 11:39 | 0 | -1 | -1 | -1 | 1206231 | 180 | 341 | 27.14 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.4e-05 | -1.284 | 0.000191 | 1.059e-09 | -3.896e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5107 | 0x0990 | pH | 2021-03-09 11:40 | 0 | -1 | -1 | -1 | 1206232 | 180 | 341 | 27.14 | 23.77 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000109 | -1.2839 | 0.000194 | 5.507e-09 | 8.003e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5109 | 0x4399 | pH | 2021-03-09 11:41 | 0 | -1 | -1 | -1 | 1206233 | 180 | 341 | 27.14 | 23.75 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000105 | -1.2836 | 0.000176 | 1.417e-09 | -3.649e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.511 | 0x083F | pH | 2021-03-09 11:42 | 0 | -1 | -1 | -1 | 1206234 | 180 | 341 | 27.14 | 23.75 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.8e-05 | -1.2838 | 0.000149 | -3.8e-10 | 1.269e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5117 | 0x7E11 | pH | 2021-03-09 11:43 | 0 | -1 | -1 | -1 | 1206235 | 180 | 341 | 27.14 | 23.75 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.7e-05 | -1.2835 | 0.000142 | -1.48e-09 | 3.838e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:52 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5119 | 0x2218 | pH | 2021-03-09 11:44 | 0 | -1 | -1 | -1 | 1206236 | 180 | 341 | 27.15 | 23.72 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000131 | -1.2833 | 0.000114 | -9.619e-10 | 2.559e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5123 | 0x76B2 | pH | 2021-03-09 11:45 | 0 | -1 | -1 | -1 | 1206237 | 180 | 341 | 27.12 | 23.75 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 7.8e-05 | -1.2827 | 0.000139 | -9.929e-09 | -5.317e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5123 | 0x21A6 | pH | 2021-03-09 11:46 | 0 | -1 | -1 | -1 | 1206238 | 180 | 341 | 27.12 | 23.69 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000129 | -1.2821 | 0.000119 | 2.451e-09 | 5.043e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.513 | 0xBDB0 | pH | 2021-03-09 11:47 | 0 | -1 | -1 | -1 | 1206239 | 180 | 341 | 27.14 | 23.69 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000134 | -1.2815 | 0.000174 | -3.888e-09 | 4.637e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5131 | 0x222F | pH | 2021-03-09 11:48 | 0 | -1 | -1 | -1 | 1206240 | 180 | 341 | 27.12 | 23.72 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.5e-05 | -1.2804 | 0.000164 | -1.294e-10 | -5.502e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5136 | 0xCB1A | pH | 2021-03-09 11:49 | 0 | -1 | -1 | -1 | 1206241 | 180 | 341 | 27.12 | 23.69 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 6.9e-05 | -1.2792 | 0.000192 | -1.262e-09 | 1.005e-08 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 5 | 36.0047 | 0.25 | 0 | 110.5139 | 0xC1FE | pH | 2021-03-09 11:50 | 0 | -1 | -1 | -1 | 1206242 | 180 | 341 | 27.12 | 23.69 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.2e-05 | -1.2775 | 0.000175 | -2.909e-09 | 2.077e-08 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 20 | 36.0047 | 0.25 | 0 | 110.514 | 0xD735 | pH | 2021-03-09 11:51 | 0 | -1 | -1 | -1 | 1206243 | 180 | 341 | 27.12 | 23.72 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 8.8e-05 | -1.2786 | 0.000122 | 1.952e-09 | -6.555e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 20 | 36.0047 | 0.25 | 0 | 110.514 | 0x7CAB | pH | 2021-03-09 11:52 | 0 | -1 | -1 | -1 | 1206244 | 180 | 341 | 27.12 | 23.69 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000116 | -1.2793 | 0.000108 | -9.635e-10 | -5.588e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 20 | 36.0047 | 0.25 | 0.002 | 110.5111 | 0x4968 | pH | 2021-03-09 11:53 | 0 | -1 | -1 | -1 | 1206245 | 180 | 341 | 27.12 | 23.69 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 0.000126 | -1.2799 | 0.00017 | -1.226e-08 | 3.56e-09 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:53 | 20 | 20 | 20 | 36.0047 | 0.25 | 0.004 | 110.5073 | 0xAF6E | pH | 2021-03-09 11:54 | 0 | -1 | -1 | -1 | 1206246 | 180 | 341 | 27.12 | 23.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000127 | -1.2801 | 0.0002 | 8.708e-10 | -5.667e-10 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 36.0047 | 0.25 | 0.003 | 110.5038 | 0x1344 | pH | 2021-03-09 11:55 | 0 | -1 | -1 | -1 | 1206247 | 180 | 341 | 27.14 | 23.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 9.8e-05 | -1.2802 | 0.000125 | -4.641e-09 | -1.783e-08 | 0.072009 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 35.9358 | 0.25 | 0.004 | 110.5004 | 0x5225 | pH | 2021-03-09 11:56 | 0 | -1 | -1 | -1 | 1206248 | 180 | 341 | 27.11 | 23.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.207 | 9.6e-05 | -1.2806 | 0.000253 | 3.517e-09 | 6.429e-09 | 0.071872 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 36.0047 | 0.25 | 0.002 | 110.4975 | 0x74F9 | pH | 2021-03-09 11:57 | 0 | -1 | -1 | -1 | 1206249 | 180 | 341 | 27.11 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2071 | 0.000118 | -1.2805 | 0.000134 | 6.811e-09 | 9.68e-09 | 0.072009 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 42.2125 | 0.25 | 0.001 | 110.495 | 0x2D3D | pH | 2021-03-09 11:58 | 0 | -1 | -1 | -1 | 1206250 | 180 | 341 | 27.11 | 23.67 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000135 | -1.28 | 0.000289 | -1.031e-08 | 6.002e-09 | 0.084425 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.9127e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 76.9067 | 0.25 | 0.001 | 110.4931 | 0xC37C | pH | 2021-03-09 11:59 | 0 | -1 | -1 | -1 | 1206251 | 180 | 341 | 27.11 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000149 | -1.2795 | 0.000194 | 1.952e-09 | -3.818e-09 | 0.15381 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 76.5618 | 0.25 | 0 | 110.4921 | 0x52B1 | pH | 2021-03-09 12:00 | 0 | -1 | -1 | -1 | 1206252 | 180 | 341 | 27.09 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000171 | -1.2793 | 0.000224 | 6.239e-09 | 4.364e-09 | 0.15312 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 76.2169 | 0.25 | 0 | 110.4921 | 0x6A18 | pH | 2021-03-09 12:01 | 0 | -1 | -1 | -1 | 1206253 | 180 | 341 | 27.09 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000138 | -1.2786 | 0.00027 | -2.303e-10 | 6.33e-09 | 0.15243 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012154 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 75.941 | 0.25 | 0 | 110.4923 | 0x764B | pH | 2021-03-09 12:02 | 0 | -1 | -1 | -1 | 1206254 | 180 | 341 | 27.08 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000137 | -1.2774 | 0.000145 | 3.143e-09 | 5.524e-09 | 0.15188 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012104 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 75.8031 | 0.25 | -0.006 | 110.4974 | 0x2C4F | pH | 2021-03-09 12:03 | 0 | -1 | -1 | -1 | 1206255 | 180 | 341 | 27.08 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000136 | -1.2729 | 0.000223 | -3.125e-09 | 7.45e-09 | 0.15161 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:54 | 20 | 20 | 20 | 75.5962 | 0.25 | -0.008 | 110.5028 | 0x5145 | pH | 2021-03-09 12:04 | 0 | -1 | -1 | -1 | 1206256 | 180 | 341 | 27.07 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000139 | -1.269 | 0.000141 | -3.492e-09 | 4.856e-10 | 0.15119 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 75.5272 | 0.25 | -0.006 | 110.5082 | 0x55DB | pH | 2021-03-09 12:05 | 0 | -1 | -1 | -1 | 1206257 | 180 | 341 | 27.08 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000198 | -1.2657 | 0.000225 | -1.14e-08 | -6.778e-09 | 0.15105 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 75.3892 | 0.25 | -0.002 | 110.5133 | 0xA5FB | pH | 2021-03-09 12:06 | 0 | -1 | -1 | -1 | 1206258 | 180 | 341 | 27.08 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000112 | -1.2632 | 0.00024 | -1.082e-09 | 7.453e-09 | 0.15078 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012002 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 75.3203 | 0.25 | -0.004 | 110.5182 | 0x6E1E | pH | 2021-03-09 12:07 | 0 | -1 | -1 | -1 | 1206259 | 180 | 341 | 27.07 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000126 | -1.2642 | 7.1e-05 | -1.144e-08 | 5.297e-09 | 0.15064 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001199 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 75.2513 | 0.25 | -0.005 | 110.5232 | 0x2E14 | pH | 2021-03-09 12:08 | 0 | -1 | -1 | -1 | 1206260 | 180 | 341 | 27.07 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000138 | -1.2624 | 0.000252 | -2.491e-09 | 1.399e-08 | 0.1505 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011977 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 75.1133 | 0.25 | -0.003 | 110.5278 | 0xD812 | pH | 2021-03-09 12:09 | 0 | -1 | -1 | -1 | 1206261 | 180 | 341 | 27.07 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.00013 | -1.2635 | 0.00024 | -7.198e-09 | 4.433e-09 | 0.15023 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 75.0444 | 0.25 | -0.004 | 110.5326 | 0xB7F9 | pH | 2021-03-09 12:10 | 0 | -1 | -1 | -1 | 1206262 | 180 | 341 | 27.08 | 23.64 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8e-05 | -1.2593 | 0.000163 | 6.195e-09 | 1.256e-08 | 0.15009 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011939 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 74.9754 | 0.25 | -0.004 | 110.5374 | 0xEA83 | pH | 2021-03-09 12:11 | 0 | -1 | -1 | -1 | 1206263 | 180 | 341 | 27.08 | 23.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.1e-05 | -1.2565 | 0.000166 | -5.088e-09 | -8.797e-09 | 0.14995 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011926 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 74.9754 | 0.25 | -0.003 | 110.5416 | 0x0D7D | pH | 2021-03-09 12:12 | 0 | -1 | -1 | -1 | 1206264 | 180 | 341 | 27.07 | 23.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 7.1e-05 | -1.2555 | 0.000185 | -5.778e-10 | 5.846e-09 | 0.14995 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011926 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 74.9064 | 0.25 | -0.003 | 110.5458 | 0x67DB | pH | 2021-03-09 12:13 | 0 | -1 | -1 | -1 | 1206265 | 180 | 341 | 27.08 | 23.62 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000119 | -1.2534 | 0.000204 | -4.907e-09 | 1.077e-08 | 0.14981 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011914 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:55 | 20 | 20 | 20 | 74.8374 | 0.25 | -0.003 | 110.5499 | 0xF635 | pH | 2021-03-09 12:14 | 0 | -1 | -1 | -1 | 1206266 | 180 | 341 | 27.08 | 23.59 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000137 | -1.252 | 0.000182 | 1.152e-09 | 4.176e-09 | 0.14967 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.7685 | 0.25 | -0.002 | 110.5537 | 0xDE1F | pH | 2021-03-09 12:15 | 0 | -1 | -1 | -1 | 1206267 | 180 | 341 | 27.09 | 23.57 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.8e-05 | -1.251 | 0.000231 | 4.386e-09 | 1.281e-08 | 0.14954 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.7685 | 0.25 | -0.004 | 110.5575 | 0x0AA3 | pH | 2021-03-09 12:16 | 0 | -1 | -1 | -1 | 1206268 | 180 | 341 | 27.09 | 23.57 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000106 | -1.2502 | 0.000211 | -4.726e-09 | 2.106e-09 | 0.14954 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.6995 | 0.25 | -0.001 | 110.5607 | 0xBD16 | pH | 2021-03-09 12:17 | 0 | -1 | -1 | -1 | 1206269 | 180 | 341 | 27.11 | 23.59 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.8e-05 | -1.2496 | 0.000142 | -1.058e-08 | 5.202e-09 | 0.1494 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011876 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.6305 | 0.25 | -0.001 | 110.5639 | 0x6FD2 | pH | 2021-03-09 12:18 | 0 | -1 | -1 | -1 | 1206270 | 180 | 341 | 27.11 | 23.59 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 5.7e-05 | -1.2491 | 0.000177 | -1.716e-09 | 4.007e-10 | 0.14926 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.00011863 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.6305 | 0.25 | -0.004 | 110.5671 | 0xCB77 | pH | 2021-03-09 12:19 | 0 | -1 | -1 | -1 | 1206271 | 180 | 341 | 27.12 | 23.57 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000101 | -1.2483 | 0.000118 | 3.817e-10 | 3.951e-09 | 0.14926 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.00011863 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.5615 | 0.25 | 0.238 | 110.3895 | 0x663D | pH | 2021-03-09 12:20 | 0 | -1 | -1 | -1 | 1206272 | 180 | 341 | 27.12 | 23.54 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.3e-05 | -1.2527 | 0.000189 | 4.975e-09 | 4.594e-10 | 0.14912 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.5615 | 0.25 | 0.237 | 110.1474 | 0x223F | pH | 2021-03-09 12:21 | 0 | -1 | -1 | -1 | 1206273 | 180 | 341 | 27.14 | 23.54 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9e-05 | -1.2593 | 0.000234 | 3.33e-09 | -1.224e-08 | 0.14912 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.4926 | 0.25 | 0.236 | 109.9064 | 0xFDE4 | pH | 2021-03-09 12:22 | 0 | -1 | -1 | -1 | 1206274 | 180 | 341 | 27.15 | 23.52 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000152 | -1.2649 | 0.000124 | -9.457e-09 | -2.455e-08 | 0.14899 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011838 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.4926 | 0.25 | 0.235 | 109.6665 | 0x8076 | pH | 2021-03-09 12:23 | 0 | -1 | -1 | -1 | 1206275 | 180 | 341 | 27.16 | 23.54 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 7e-05 | -1.2696 | 0.000165 | -5.614e-09 | 1.102e-08 | 0.14899 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011838 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.4236 | 0.25 | 0.235 | 109.4265 | 0x515D | pH | 2021-03-09 12:25 | 0 | -1 | -1 | -1 | 1206276 | 180 | 341 | 27.16 | 23.52 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000137 | -1.2735 | 0.000141 | -7.783e-09 | -5.262e-09 | 0.14885 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:56 | 20 | 20 | 20 | 74.4236 | 0.25 | 0.235 | 109.187 | 0x8165 | pH | 2021-03-09 12:26 | 0 | -1 | -1 | -1 | 1206277 | 180 | 341 | 27.16 | 23.46 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.9e-05 | -1.2774 | 0.000202 | -5.441e-09 | 2.494e-09 | 0.14885 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.3546 | 0.25 | 0.235 | 108.9476 | 0xFC88 | pH | 2021-03-09 12:27 | 0 | -1 | -1 | -1 | 1206278 | 180 | 341 | 27.18 | 23.46 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000111 | -1.2807 | 0.000114 | 3.359e-09 | -2.861e-09 | 0.14871 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.3546 | 0.25 | 0.234 | 108.7084 | 0x7F4F | pH | 2021-03-09 12:28 | 0 | -1 | -1 | -1 | 1206279 | 180 | 341 | 27.19 | 23.44 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.6e-05 | -1.2846 | 0.000165 | 1.411e-09 | -6.509e-10 | 0.14871 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.2856 | 0.25 | 0.235 | 108.4693 | 0x0870 | pH | 2021-03-09 12:29 | 0 | -1 | -1 | -1 | 1206280 | 180 | 341 | 27.2 | 23.42 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000115 | -1.2873 | 0.000166 | -4.93e-09 | 7.376e-09 | 0.14857 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.000118 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.2856 | 0.25 | 0.234 | 108.2304 | 0x3B8A | pH | 2021-03-09 12:30 | 0 | -1 | -1 | -1 | 1206281 | 180 | 341 | 27.2 | 23.44 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000126 | -1.2898 | 0.000138 | -5.089e-09 | -7.428e-09 | 0.14857 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.000118 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.2856 | 0.25 | 0.234 | 107.9915 | 0xF074 | pH | 2021-03-09 12:31 | 0 | -1 | -1 | -1 | 1206282 | 180 | 341 | 27.2 | 23.42 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.9e-05 | -1.2918 | 0.000217 | 5.777e-09 | 1.405e-09 | 0.14857 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.000118 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.2167 | 0.25 | 0.234 | 107.7527 | 0xB270 | pH | 2021-03-09 12:32 | 0 | -1 | -1 | -1 | 1206283 | 180 | 341 | 27.2 | 23.42 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.9e-05 | -1.2937 | 0.000146 | -5.662e-09 | 1.219e-08 | 0.14843 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011787 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.2167 | 0.25 | 0.234 | 107.5138 | 0x7F0E | pH | 2021-03-09 12:33 | 0 | -1 | -1 | -1 | 1206284 | 180 | 341 | 27.22 | 23.42 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000139 | -1.2955 | 0.000131 | -5.545e-09 | 1.806e-08 | 0.14843 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011787 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.1477 | 0.25 | 0.234 | 107.2743 | 0x5B34 | pH | 2021-03-09 12:34 | 0 | -1 | -1 | -1 | 1206285 | 180 | 341 | 27.22 | 23.39 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000114 | -1.2973 | 0.000144 | 1.272e-09 | 8.313e-09 | 0.1483 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.1477 | 0.25 | 0.001 | 107.2128 | 0x1DE3 | pH | 2021-03-09 12:35 | 0 | -1 | -1 | -1 | 1206286 | 180 | 341 | 27.24 | 23.39 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.7e-05 | -1.2981 | 0.000228 | -5.466e-09 | -5.152e-09 | 0.1483 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:57 | 20 | 20 | 20 | 74.1477 | 0.25 | 0 | 107.2124 | 0x0372 | pH | 2021-03-09 12:36 | 0 | -1 | -1 | -1 | 1206287 | 180 | 341 | 27.24 | 23.37 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000152 | -1.2984 | 0.000215 | -5.238e-09 | 1.183e-08 | 0.1483 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0787 | 0.25 | 0 | 107.2124 | 0xAE51 | pH | 2021-03-09 12:37 | 0 | -1 | -1 | -1 | 1206288 | 180 | 341 | 27.26 | 23.37 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000114 | -1.2982 | 0.000205 | -4.826e-10 | 1.062e-08 | 0.14816 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0787 | 0.25 | 0 | 107.2123 | 0x57BC | pH | 2021-03-09 12:38 | 0 | -1 | -1 | -1 | 1206289 | 180 | 341 | 27.24 | 23.37 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000112 | -1.298 | 0.000158 | 1.204e-09 | 5.24e-09 | 0.14816 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0787 | 0.25 | 0 | 107.2123 | 0xFD7E | pH | 2021-03-09 12:39 | 0 | -1 | -1 | -1 | 1206290 | 180 | 341 | 27.27 | 23.37 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000134 | -1.2975 | 0.000104 | -5.623e-10 | 9.868e-09 | 0.14816 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0787 | 0.25 | 0 | 107.2123 | 0x8DFF | pH | 2021-03-09 12:40 | 0 | -1 | -1 | -1 | 1206291 | 180 | 341 | 27.27 | 23.34 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 5.8e-05 | -1.297 | 0.000163 | -4.3e-09 | -1.065e-08 | 0.14816 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0787 | 0.25 | 0 | 107.2123 | 0x26C4 | pH | 2021-03-09 12:41 | 0 | -1 | -1 | -1 | 1206292 | 180 | 341 | 27.27 | 23.34 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000122 | -1.2966 | 0.00018 | -3.079e-09 | -3.977e-11 | 0.14816 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0097 | 0.25 | 0 | 107.2123 | 0xC07A | pH | 2021-03-09 12:42 | 0 | -1 | -1 | -1 | 1206293 | 180 | 341 | 27.27 | 23.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 7.3e-05 | -1.2971 | 0.000197 | 2.142e-09 | 4.682e-09 | 0.14802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011749 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0097 | 0.25 | 0 | 107.2124 | 0xBBD5 | pH | 2021-03-09 12:43 | 0 | -1 | -1 | -1 | 1206294 | 180 | 341 | 27.28 | 23.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.00013 | -1.2974 | 0.000272 | 8.933e-09 | 6.783e-09 | 0.14802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011749 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0097 | 0.25 | 0 | 107.2133 | 0xCF60 | pH | 2021-03-09 12:44 | 0 | -1 | -1 | -1 | 1206295 | 180 | 341 | 27.3 | 23.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000124 | -1.3003 | 0.00022 | -6.453e-10 | -1.099e-08 | 0.14802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011749 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 74.0097 | 0.25 | 0 | 107.2141 | 0x974D | pH | 2021-03-09 12:45 | 0 | -1 | -1 | -1 | 1206296 | 180 | 341 | 27.28 | 23.31 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000102 | -1.3042 | 0.000181 | 3.255e-11 | -4.431e-09 | 0.14802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011749 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 73.9408 | 0.25 | 0 | 107.2144 | 0xEE66 | pH | 2021-03-09 12:46 | 0 | -1 | -1 | -1 | 1206297 | 180 | 341 | 27.3 | 23.31 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 7.1e-05 | -1.3009 | 0.000145 | 1.478e-09 | 3.45e-09 | 0.14788 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:58 | 20 | 20 | 20 | 73.9408 | 0.25 | 0 | 107.2149 | 0xB5DE | pH | 2021-03-09 12:47 | 0 | -1 | -1 | -1 | 1206298 | 180 | 341 | 27.3 | 23.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000161 | -1.2991 | 0.000174 | -4.426e-09 | 1.495e-08 | 0.14788 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.9408 | 0.25 | 0 | 107.2151 | 0x2DCC | pH | 2021-03-09 12:48 | 0 | -1 | -1 | -1 | 1206299 | 180 | 341 | 27.31 | 23.26 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.9e-05 | -1.2965 | 0.000168 | -2.731e-09 | 4.242e-09 | 0.14788 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.9408 | 0.25 | 0 | 107.2153 | 0x986C | pH | 2021-03-09 12:49 | 0 | -1 | -1 | -1 | 1206300 | 180 | 341 | 27.31 | 23.29 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.5e-05 | -1.294 | 0.000171 | -5.793e-09 | 8.393e-10 | 0.14788 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8718 | 0.25 | 0 | 107.2153 | 0xF1A0 | pH | 2021-03-09 12:50 | 0 | -1 | -1 | -1 | 1206301 | 180 | 341 | 27.31 | 23.21 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.5e-05 | -1.2921 | 0.000171 | 3.857e-09 | 9.039e-09 | 0.14774 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8718 | 0.25 | 0 | 107.2157 | 0x08CF | pH | 2021-03-09 12:51 | 0 | -1 | -1 | -1 | 1206302 | 180 | 341 | 27.31 | 23.19 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000105 | -1.2904 | 0.00018 | -6.155e-09 | 5.012e-09 | 0.14774 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8718 | 0.25 | 0 | 107.2157 | 0xA1A5 | pH | 2021-03-09 12:52 | 0 | -1 | -1 | -1 | 1206303 | 180 | 341 | 27.31 | 23.19 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000115 | -1.289 | 0.000213 | -3.625e-09 | -5.922e-09 | 0.14774 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8718 | 0.25 | 0 | 107.2157 | 0x87FE | pH | 2021-03-09 12:53 | 0 | -1 | -1 | -1 | 1206304 | 180 | 341 | 27.31 | 23.19 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8e-05 | -1.2874 | 0.000101 | -3.76e-09 | -1.603e-08 | 0.14774 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8028 | 0.25 | 0 | 107.2157 | 0x3E01 | pH | 2021-03-09 12:54 | 0 | -1 | -1 | -1 | 1206305 | 180 | 341 | 27.31 | 23.21 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 6.6e-05 | -1.2859 | 0.000153 | -2.099e-09 | 2.046e-09 | 0.14761 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8028 | 0.25 | 0 | 107.2157 | 0x29CA | pH | 2021-03-09 12:55 | 0 | -1 | -1 | -1 | 1206306 | 180 | 341 | 27.31 | 23.24 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000127 | -1.2849 | 0.00016 | -6.957e-09 | 5.039e-09 | 0.14761 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8028 | 0.25 | 0 | 107.2158 | 0xFE9E | pH | 2021-03-09 12:56 | 0 | -1 | -1 | -1 | 1206307 | 180 | 341 | 27.3 | 23.19 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.2e-05 | -1.2834 | 0.000154 | -6.917e-10 | 1.052e-09 | 0.14761 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 13:59 | 20 | 20 | 20 | 73.8028 | 0.25 | 0 | 107.2158 | 0xD9F1 | pH | 2021-03-09 12:57 | 0 | -1 | -1 | -1 | 1206308 | 180 | 341 | 27.32 | 23.21 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 7.9e-05 | -1.282 | 0.000132 | 3.854e-09 | 4.89e-09 | 0.14761 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.7338 | 0.25 | 0 | 107.2158 | 0xA5CE | pH | 2021-03-09 12:58 | 0 | -1 | -1 | -1 | 1206309 | 180 | 341 | 27.31 | 23.19 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000119 | -1.2808 | 0.000113 | -4.984e-09 | -7.843e-09 | 0.14747 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.8028 | 0.25 | 0 | 107.2158 | 0x68DD | pH | 2021-03-09 12:59 | 0 | -1 | -1 | -1 | 1206310 | 180 | 341 | 27.31 | 23.21 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000116 | -1.2794 | 0.000112 | -2.664e-09 | -3.952e-09 | 0.14761 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.7338 | 0.25 | 0 | 107.2158 | 0xD65F | pH | 2021-03-09 13:00 | 0 | -1 | -1 | -1 | 1206311 | 180 | 341 | 27.32 | 23.21 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000134 | -1.2781 | 0.00011 | -3.788e-09 | -9.033e-09 | 0.14747 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.7338 | 0.25 | 0 | 107.2158 | 0x4C91 | pH | 2021-03-09 13:01 | 0 | -1 | -1 | -1 | 1206312 | 180 | 341 | 27.31 | 23.24 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 9.2e-05 | -1.277 | 0.000208 | -1.037e-08 | -1.13e-08 | 0.14747 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.7338 | 0.25 | 0 | 107.2158 | 0xF008 | pH | 2021-03-09 13:02 | 0 | -1 | -1 | -1 | 1206313 | 180 | 341 | 27.31 | 23.19 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.2e-05 | -1.2761 | 0.000172 | -3.842e-09 | -3.654e-09 | 0.14747 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.7338 | 0.25 | 0 | 107.2158 | 0xCB9A | pH | 2021-03-09 13:03 | 0 | -1 | -1 | -1 | 1206314 | 180 | 341 | 27.31 | 23.14 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2077 | 0.000166 | -1.2753 | 7.4e-05 | -1.291e-08 | 1.04e-08 | 0.14747 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.6649 | 0.25 | 0 | 107.2158 | 0x9A41 | pH | 2021-03-09 13:04 | 0 | -1 | -1 | -1 | 1206315 | 180 | 341 | 27.31 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000113 | -1.2744 | 0.000281 | -7.324e-11 | 1.688e-09 | 0.14733 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011686 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.6649 | 0.25 | 0 | 107.2158 | 0x3128 | pH | 2021-03-09 13:05 | 0 | -1 | -1 | -1 | 1206316 | 180 | 341 | 27.31 | 23.17 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2077 | 7.6e-05 | -1.2736 | 0.000154 | -7.765e-09 | 2.709e-09 | 0.14733 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011686 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.6649 | 0.25 | -0.234 | 107.3911 | 0x87F1 | pH | 2021-03-09 13:06 | 0 | -1 | -1 | -1 | 1206317 | 180 | 341 | 27.3 | 23.17 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000112 | -1.2719 | 0.000113 | -8.211e-10 | 1.524e-08 | 0.14733 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011686 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:00 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.234 | 107.6291 | 0x4491 | pH | 2021-03-09 13:07 | 0 | -1 | -1 | -1 | 1206318 | 180 | 341 | 27.3 | 23.17 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000104 | -1.2702 | 0.000237 | 2.87e-09 | 1.532e-08 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.234 | 107.8674 | 0x44C7 | pH | 2021-03-09 13:08 | 0 | -1 | -1 | -1 | 1206319 | 180 | 341 | 27.31 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 5.8e-05 | -1.2686 | 0.000162 | -2.174e-09 | -7.842e-09 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 108.1059 | 0xE343 | pH | 2021-03-09 13:09 | 0 | -1 | -1 | -1 | 1206320 | 180 | 341 | 27.3 | 23.14 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.1e-05 | -1.2668 | 0.000155 | -1.508e-09 | -4.89e-10 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 108.3454 | 0x20A2 | pH | 2021-03-09 13:10 | 0 | -1 | -1 | -1 | 1206321 | 180 | 341 | 27.3 | 23.17 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000135 | -1.2654 | 0.000117 | 4.806e-09 | -3.991e-09 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 108.5838 | 0xC084 | pH | 2021-03-09 13:11 | 0 | -1 | -1 | -1 | 1206322 | 180 | 341 | 27.31 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.2e-05 | -1.264 | 0.000103 | 1.351e-09 | 1.774e-08 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 108.8231 | 0xDC5A | pH | 2021-03-09 13:12 | 0 | -1 | -1 | -1 | 1206323 | 180 | 341 | 27.3 | 23.17 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 7.7e-05 | -1.2628 | 0.000222 | -6.457e-09 | 4.802e-09 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 109.0626 | 0xD3E2 | pH | 2021-03-09 13:13 | 0 | -1 | -1 | -1 | 1206324 | 180 | 341 | 27.3 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9e-05 | -1.2621 | 0.00011 | -7.276e-09 | 9.454e-09 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 109.3021 | 0xCA73 | pH | 2021-03-09 13:14 | 0 | -1 | -1 | -1 | 1206325 | 180 | 341 | 27.3 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000141 | -1.2614 | 0.000154 | -8.081e-10 | -3.196e-10 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 109.5422 | 0x3BDF | pH | 2021-03-09 13:15 | 0 | -1 | -1 | -1 | 1206326 | 180 | 341 | 27.3 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000106 | -1.2604 | 0.000118 | 5.927e-09 | 1.036e-08 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5959 | 0.25 | -0.235 | 109.7824 | 0x5F6A | pH | 2021-03-09 13:16 | 0 | -1 | -1 | -1 | 1206327 | 180 | 341 | 27.3 | 23.14 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.9e-05 | -1.2586 | 8.9e-05 | 1.694e-09 | -1.215e-08 | 0.14719 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5269 | 0.25 | -0.237 | 110.0231 | 0xD2EA | pH | 2021-03-09 13:17 | 0 | -1 | -1 | -1 | 1206328 | 180 | 341 | 27.3 | 23.11 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000144 | -1.2567 | 0.000168 | 4.255e-09 | 6.338e-09 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:01 | 20 | 20 | 20 | 73.5269 | 0.25 | -0.237 | 110.2647 | 0x922D | pH | 2021-03-09 13:18 | 0 | -1 | -1 | -1 | 1206329 | 180 | 341 | 27.3 | 23.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 9.8e-05 | -1.2531 | 0.000104 | -5.151e-09 | 1.403e-08 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.5269 | 0.25 | -0.24 | 110.508 | 0xBD1E | pH | 2021-03-09 13:19 | 0 | -1 | -1 | -1 | 1206330 | 180 | 341 | 27.3 | 23.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000126 | -1.2504 | 0.000144 | -7.2e-09 | -1.335e-08 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.5269 | 0.25 | -0.003 | 110.5691 | 0xDB93 | pH | 2021-03-09 13:20 | 0 | -1 | -1 | -1 | 1206331 | 180 | 341 | 27.3 | 23.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.9e-05 | -1.2486 | 0.00026 | 3.565e-09 | 1.315e-08 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.5269 | 0.25 | -0.003 | 110.5721 | 0x2B63 | pH | 2021-03-09 13:21 | 0 | -1 | -1 | -1 | 1206332 | 180 | 341 | 27.28 | 23.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000102 | -1.2454 | 0.000124 | 1.938e-09 | -2.264e-08 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.5269 | 0.25 | -0.001 | 110.5738 | 0x3C88 | pH | 2021-03-09 13:22 | 0 | -1 | -1 | -1 | 1206333 | 180 | 341 | 27.3 | 23.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000111 | -1.2423 | 0.000199 | -1.013e-08 | -9.348e-10 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.5269 | 0.25 | 0 | 110.5751 | 0x8690 | pH | 2021-03-09 13:23 | 0 | -1 | -1 | -1 | 1206334 | 180 | 341 | 27.3 | 23.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000145 | -1.2403 | 0.000148 | -3.959e-09 | 1.519e-09 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.5269 | 0.25 | 0 | 110.5763 | 0xDBFF | pH | 2021-03-09 13:24 | 0 | -1 | -1 | -1 | 1206335 | 180 | 341 | 27.28 | 23.06 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2077 | 9.2e-05 | -1.2394 | 0.000137 | -1.075e-08 | 3.99e-09 | 0.14705 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.4579 | 0.25 | 0 | 110.5772 | 0x7FF8 | pH | 2021-03-09 13:26 | 0 | -1 | -1 | -1 | 1206336 | 180 | 341 | 27.28 | 23.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 7.9e-05 | -1.2388 | 0.000207 | 1.177e-09 | -7.07e-09 | 0.14692 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.4579 | 0.25 | -0.001 | 110.5781 | 0x5C6D | pH | 2021-03-09 13:27 | 0 | -1 | -1 | -1 | 1206337 | 180 | 341 | 27.28 | 23.06 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 9.4e-05 | -1.2371 | 9e-05 | 4.606e-09 | 7.289e-09 | 0.14692 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.4579 | 0.25 | 0 | 110.5787 | 0x39DD | pH | 2021-03-09 13:28 | 0 | -1 | -1 | -1 | 1206338 | 180 | 341 | 27.28 | 23.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000111 | -1.236 | 0.000125 | -1.109e-08 | 1.269e-09 | 0.14692 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:02 | 20 | 20 | 20 | 73.4579 | 0.25 | 0 | 110.5794 | 0x6D2D | pH | 2021-03-09 13:29 | 0 | -1 | -1 | -1 | 1206339 | 180 | 341 | 27.27 | 23.09 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.1e-05 | -1.2355 | 0.000163 | -7.601e-10 | 1.666e-08 | 0.14692 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.4579 | 0.25 | 0 | 110.5799 | 0xCC42 | pH | 2021-03-09 13:30 | 0 | -1 | -1 | -1 | 1206340 | 180 | 341 | 27.27 | 23.04 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000106 | -1.2374 | 0.000104 | 4.91e-09 | 7.249e-09 | 0.14692 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | -0.002 | 110.5804 | 0x199A | pH | 2021-03-09 13:31 | 0 | -1 | -1 | -1 | 1206341 | 180 | 341 | 27.27 | 23.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000128 | -1.2362 | 0.000174 | 1.709e-09 | 1.117e-09 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | -0.002 | 110.581 | 0x2EC8 | pH | 2021-03-09 13:32 | 0 | -1 | -1 | -1 | 1206342 | 180 | 341 | 27.27 | 23.01 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000118 | -1.2352 | 0.000193 | -7.567e-09 | 8.825e-10 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | 0 | 110.5814 | 0x90C7 | pH | 2021-03-09 13:33 | 0 | -1 | -1 | -1 | 1206343 | 180 | 341 | 27.27 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000153 | -1.234 | 0.000164 | -7.192e-09 | 5.6e-09 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | 0 | 110.5816 | 0x2A71 | pH | 2021-03-09 13:34 | 0 | -1 | -1 | -1 | 1206344 | 180 | 341 | 27.26 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000114 | -1.2328 | 0.000189 | 3.7e-09 | 1.618e-08 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | 0 | 110.582 | 0x7593 | pH | 2021-03-09 13:35 | 0 | -1 | -1 | -1 | 1206345 | 180 | 341 | 27.27 | 22.99 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000126 | -1.2322 | 0.000218 | 2.1e-09 | -9.069e-10 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | 0 | 110.5823 | 0xC126 | pH | 2021-03-09 13:36 | 0 | -1 | -1 | -1 | 1206346 | 180 | 341 | 27.27 | 22.99 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 5.4e-05 | -1.2322 | 0.000173 | -9.29e-09 | 4.311e-10 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | 0 | 110.5825 | 0xE78E | pH | 2021-03-09 13:37 | 0 | -1 | -1 | -1 | 1206347 | 180 | 341 | 27.27 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000139 | -1.232 | 0.000161 | 4.003e-09 | 1.152e-08 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.389 | 0.25 | 0 | 110.5829 | 0xF14C | pH | 2021-03-09 13:38 | 0 | -1 | -1 | -1 | 1206348 | 180 | 341 | 27.27 | 22.99 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000138 | -1.2321 | 0.000155 | 3.651e-09 | -9.966e-09 | 0.14678 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:03 | 20 | 20 | 20 | 73.32 | 0.25 | 0 | 110.5829 | 0x4BE6 | pH | 2021-03-09 13:39 | 0 | -1 | -1 | -1 | 1206349 | 180 | 341 | 27.27 | 22.99 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000144 | -1.2314 | 0.000141 | 1.312e-08 | 3.535e-09 | 0.14664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.32 | 0.25 | 0 | 110.5833 | 0xBEBD | pH | 2021-03-09 13:40 | 0 | -1 | -1 | -1 | 1206350 | 180 | 341 | 27.27 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000112 | -1.2308 | 0.000187 | 2.113e-09 | -9.197e-09 | 0.14664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.32 | 0.25 | 0 | 110.5835 | 0xD191 | pH | 2021-03-09 13:41 | 0 | -1 | -1 | -1 | 1206351 | 180 | 341 | 27.26 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000142 | -1.23 | 9.6e-05 | -2.23e-10 | 1.312e-08 | 0.14664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.32 | 0.25 | 0 | 110.5839 | 0x21E2 | pH | 2021-03-09 13:42 | 0 | -1 | -1 | -1 | 1206352 | 180 | 341 | 27.26 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000134 | -1.2291 | 0.000205 | 4.372e-09 | 6.042e-09 | 0.14664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.32 | 0.25 | 0 | 110.5839 | 0x37B3 | pH | 2021-03-09 13:43 | 0 | -1 | -1 | -1 | 1206353 | 180 | 341 | 27.26 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000129 | -1.2286 | 0.000151 | -6.09e-09 | -7.539e-09 | 0.14664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.32 | 0.25 | 0 | 110.5841 | 0x81A3 | pH | 2021-03-09 13:44 | 0 | -1 | -1 | -1 | 1206354 | 180 | 341 | 27.26 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.4e-05 | -1.2285 | 0.000138 | -5.081e-09 | 1.275e-09 | 0.14664 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.251 | 0.25 | 0 | 110.5842 | 0xE87B | pH | 2021-03-09 13:45 | 0 | -1 | -1 | -1 | 1206355 | 180 | 341 | 27.26 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.9e-05 | -1.2279 | 0.000199 | -3.276e-09 | 5.111e-10 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.251 | 0.25 | 0 | 110.5845 | 0x6310 | pH | 2021-03-09 13:46 | 0 | -1 | -1 | -1 | 1206356 | 180 | 341 | 27.27 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 6.1e-05 | -1.2275 | 0.000142 | -4.186e-09 | -1.543e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.251 | 0.25 | 0 | 110.5847 | 0x4512 | pH | 2021-03-09 13:47 | 0 | -1 | -1 | -1 | 1206357 | 180 | 341 | 27.26 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.5e-05 | -1.2273 | 0.000158 | 6.182e-09 | -7.953e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.251 | 0.25 | 0 | 110.5849 | 0x4644 | pH | 2021-03-09 13:48 | 0 | -1 | -1 | -1 | 1206358 | 180 | 341 | 27.26 | 22.94 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000151 | -1.2275 | 0.000161 | -4.6e-09 | 2.273e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.251 | 0.25 | 0 | 110.585 | 0xAF77 | pH | 2021-03-09 13:49 | 0 | -1 | -1 | -1 | 1206359 | 180 | 341 | 27.24 | 22.99 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000135 | -1.2276 | 0.000216 | -1.207e-09 | 8.611e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:04 | 20 | 20 | 20 | 73.251 | 0.25 | 0 | 110.5852 | 0x6F81 | pH | 2021-03-09 13:50 | 0 | -1 | -1 | -1 | 1206360 | 180 | 341 | 27.26 | 22.96 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2078 | 9.8e-05 | -1.2276 | 0.000123 | -1.36e-09 | -2.615e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.251 | 0.25 | 0.238 | 110.4061 | 0x8D7B | pH | 2021-03-09 13:51 | 0 | -1 | -1 | -1 | 1206361 | 180 | 341 | 27.26 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000129 | -1.2299 | 0.000115 | -9.843e-09 | 2.395e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.251 | 0.25 | 0.237 | 110.164 | 0xD21F | pH | 2021-03-09 13:52 | 0 | -1 | -1 | -1 | 1206362 | 180 | 341 | 27.24 | 22.96 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000103 | -1.2347 | 0.000247 | 2.113e-09 | -5.445e-10 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.251 | 0.25 | 0.236 | 109.9233 | 0x9135 | pH | 2021-03-09 13:53 | 0 | -1 | -1 | -1 | 1206363 | 180 | 341 | 27.26 | 22.96 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.7e-05 | -1.2395 | 8.7e-05 | -3.608e-09 | 3.52e-11 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.251 | 0.25 | 0.236 | 109.6831 | 0xE635 | pH | 2021-03-09 13:54 | 0 | -1 | -1 | -1 | 1206364 | 180 | 341 | 27.26 | 22.94 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 3.9e-05 | -1.2439 | 0.00016 | -1.145e-08 | -1.987e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.251 | 0.25 | 0.236 | 109.4433 | 0x476C | pH | 2021-03-09 13:55 | 0 | -1 | -1 | -1 | 1206365 | 180 | 341 | 27.27 | 22.96 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000106 | -1.2462 | 0.000188 | 1.126e-08 | 3.446e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.251 | 0.25 | 0.235 | 109.2037 | 0x5C82 | pH | 2021-03-09 13:56 | 0 | -1 | -1 | -1 | 1206366 | 180 | 341 | 27.26 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 9.3e-05 | -1.2482 | 0.00017 | -9.281e-09 | 6.379e-09 | 0.1465 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.235 | 108.9641 | 0x7B4F | pH | 2021-03-09 13:57 | 0 | -1 | -1 | -1 | 1206367 | 180 | 341 | 27.26 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8e-05 | -1.2503 | 0.000177 | -3.107e-09 | 4.601e-10 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.235 | 108.725 | 0x8535 | pH | 2021-03-09 13:58 | 0 | -1 | -1 | -1 | 1206368 | 180 | 341 | 27.28 | 22.94 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 6.5e-05 | -1.2522 | 0.000149 | 9.924e-09 | -1.192e-09 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.235 | 108.4858 | 0xD944 | pH | 2021-03-09 13:59 | 0 | -1 | -1 | -1 | 1206369 | 180 | 341 | 27.27 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.9e-05 | -1.2538 | 0.000166 | -9.836e-09 | 8.75e-09 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:05 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.235 | 108.2466 | 0x364B | pH | 2021-03-09 14:00 | 0 | -1 | -1 | -1 | 1206370 | 180 | 341 | 27.27 | 22.96 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 4.7e-05 | -1.2553 | 0.000158 | -5.024e-09 | 1.398e-08 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.235 | 108.0071 | 0x58C4 | pH | 2021-03-09 14:01 | 0 | -1 | -1 | -1 | 1206371 | 180 | 341 | 27.27 | 22.96 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 7.6e-05 | -1.2557 | 0.000132 | -1.366e-08 | -5.266e-10 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.235 | 107.7689 | 0x0A82 | pH | 2021-03-09 14:02 | 0 | -1 | -1 | -1 | 1206372 | 180 | 341 | 27.28 | 22.94 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000116 | -1.2566 | 0.000121 | 2.93e-10 | -4.918e-09 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.234 | 107.5301 | 0x62BB | pH | 2021-03-09 14:03 | 0 | -1 | -1 | -1 | 1206373 | 180 | 341 | 27.27 | 22.91 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000134 | -1.2574 | 0.000208 | -6.349e-09 | -7.841e-09 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.234 | 107.291 | 0xEA8F | pH | 2021-03-09 14:04 | 0 | -1 | -1 | -1 | 1206374 | 180 | 341 | 27.27 | 22.91 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000163 | -1.2583 | 0.000154 | -5.622e-09 | 2.63e-09 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1821 | 0.25 | 0.003 | 107.2227 | 0x938A | pH | 2021-03-09 14:05 | 0 | -1 | -1 | -1 | 1206375 | 180 | 341 | 27.27 | 22.91 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.00013 | -1.2585 | 0.000153 | -5.273e-10 | -9.007e-09 | 0.14636 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1131 | 0.25 | 0.001 | 107.2205 | 0x1DBF | pH | 2021-03-09 14:06 | 0 | -1 | -1 | -1 | 1206376 | 180 | 341 | 27.28 | 22.91 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000101 | -1.2586 | 0.000141 | 9.098e-10 | -1.179e-08 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1131 | 0.25 | 0 | 107.2194 | 0x5628 | pH | 2021-03-09 14:07 | 0 | -1 | -1 | -1 | 1206377 | 180 | 341 | 27.28 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.00012 | -1.2583 | 0.000181 | 1.367e-10 | -7.089e-09 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1131 | 0.25 | 0 | 107.2183 | 0xE32B | pH | 2021-03-09 14:08 | 0 | -1 | -1 | -1 | 1206378 | 180 | 341 | 27.28 | 22.91 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 9.8e-05 | -1.2579 | 0.000177 | -1.869e-09 | 4.607e-09 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1131 | 0.25 | 0.001 | 107.2174 | 0xE852 | pH | 2021-03-09 14:09 | 0 | -1 | -1 | -1 | 1206379 | 180 | 341 | 27.28 | 22.91 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000128 | -1.2574 | 0.000173 | -1.89e-09 | -5.298e-10 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1131 | 0.25 | 0.002 | 107.2167 | 0xEA8A | pH | 2021-03-09 14:10 | 0 | -1 | -1 | -1 | 1206380 | 180 | 341 | 27.3 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000126 | -1.2569 | 0.000133 | -1.154e-09 | 7.475e-09 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:06 | 20 | 20 | 20 | 73.1131 | 0.25 | 0 | 107.2162 | 0x5B42 | pH | 2021-03-09 14:11 | 0 | -1 | -1 | -1 | 1206381 | 180 | 341 | 27.3 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000193 | -1.2565 | 0.000154 | -1.25e-09 | 1.047e-08 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.1131 | 0.25 | 0.001 | 107.2157 | 0xEC90 | pH | 2021-03-09 14:12 | 0 | -1 | -1 | -1 | 1206382 | 180 | 341 | 27.3 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000148 | -1.2559 | 0.000127 | 1.159e-09 | -2.513e-09 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.1131 | 0.25 | 0 | 107.2152 | 0x8460 | pH | 2021-03-09 14:13 | 0 | -1 | -1 | -1 | 1206383 | 180 | 341 | 27.3 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000115 | -1.2554 | 0.000223 | 3.874e-10 | -4.6e-09 | 0.14623 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2147 | 0x7CBC | pH | 2021-03-09 14:14 | 0 | -1 | -1 | -1 | 1206384 | 180 | 341 | 27.31 | 22.88 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000128 | -1.2548 | 0.00022 | 4.346e-10 | 5.544e-09 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2143 | 0xAD60 | pH | 2021-03-09 14:15 | 0 | -1 | -1 | -1 | 1206385 | 180 | 341 | 27.31 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000122 | -1.2544 | 0.0002 | 1.541e-09 | -8.598e-10 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0.003 | 107.214 | 0x7744 | pH | 2021-03-09 14:16 | 0 | -1 | -1 | -1 | 1206386 | 180 | 341 | 27.31 | 22.91 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000137 | -1.2538 | 0.000144 | -3.712e-09 | 1.292e-09 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2137 | 0xE364 | pH | 2021-03-09 14:17 | 0 | -1 | -1 | -1 | 1206387 | 180 | 341 | 27.31 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000147 | -1.2531 | 0.000309 | 8.577e-10 | 7.837e-09 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2134 | 0x6229 | pH | 2021-03-09 14:18 | 0 | -1 | -1 | -1 | 1206388 | 180 | 341 | 27.31 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000121 | -1.2528 | 0.000237 | 3.53e-09 | -1.364e-08 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2131 | 0x3E3E | pH | 2021-03-09 14:19 | 0 | -1 | -1 | -1 | 1206389 | 180 | 341 | 27.31 | 22.88 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000125 | -1.2524 | 0.000101 | 2.183e-09 | -1.059e-08 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2131 | 0xC285 | pH | 2021-03-09 14:20 | 0 | -1 | -1 | -1 | 1206390 | 180 | 341 | 27.31 | 22.86 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 9.3e-05 | -1.252 | 0.000219 | 6.022e-11 | -1.838e-09 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:07 | 20 | 20 | 20 | 73.0441 | 0.25 | 0 | 107.2128 | 0xC598 | pH | 2021-03-09 14:21 | 0 | -1 | -1 | -1 | 1206391 | 180 | 341 | 27.32 | 22.86 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000102 | -1.2516 | 0.000183 | -1.88e-09 | 7.463e-09 | 0.14609 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9751 | 0.25 | 0 | 107.2126 | 0x3E02 | pH | 2021-03-09 14:22 | 0 | -1 | -1 | -1 | 1206392 | 180 | 341 | 27.32 | 22.86 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000179 | -1.2512 | 0.00016 | 9.652e-10 | 9.335e-09 | 0.14595 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9751 | 0.25 | 0 | 107.2126 | 0xF141 | pH | 2021-03-09 14:23 | 0 | -1 | -1 | -1 | 1206393 | 180 | 341 | 27.32 | 22.86 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000122 | -1.2509 | 0.000199 | 1.09e-09 | -9.683e-09 | 0.14595 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9751 | 0.25 | 0 | 107.2124 | 0x3FFF | pH | 2021-03-09 14:24 | 0 | -1 | -1 | -1 | 1206394 | 180 | 341 | 27.34 | 22.84 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000141 | -1.2503 | 0.000175 | -1.531e-08 | -6.018e-10 | 0.14595 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9751 | 0.25 | 0 | 107.2123 | 0x6907 | pH | 2021-03-09 14:26 | 0 | -1 | -1 | -1 | 1206395 | 180 | 341 | 27.35 | 22.86 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.00013 | -1.2496 | 0.000196 | -2.865e-09 | -9.882e-09 | 0.14595 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9751 | 0.25 | 0 | 107.2122 | 0xAB32 | pH | 2021-03-09 14:27 | 0 | -1 | -1 | -1 | 1206396 | 180 | 341 | 27.35 | 22.81 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000109 | -1.2495 | 0.000216 | -9.44e-11 | -2.991e-09 | 0.14595 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9751 | 0.25 | 0 | 107.2122 | 0x0AC6 | pH | 2021-03-09 14:28 | 0 | -1 | -1 | -1 | 1206397 | 180 | 341 | 27.37 | 22.81 | 11.716 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000165 | -1.2494 | 0.000221 | 4.202e-09 | -3.319e-09 | 0.14595 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0x3C94 | pH | 2021-03-09 14:29 | 0 | -1 | -1 | -1 | 1206398 | 180 | 341 | 27.37 | 22.81 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000159 | -1.2492 | 0.000193 | 3.556e-09 | -2.503e-09 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0x0C6C | pH | 2021-03-09 14:30 | 0 | -1 | -1 | -1 | 1206399 | 180 | 341 | 27.37 | 22.79 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2078 | 0.000103 | -1.249 | 0.000186 | 2.272e-09 | -1.55e-08 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0x8971 | pH | 2021-03-09 14:31 | 0 | -1 | -1 | -1 | 1206400 | 180 | 341 | 27.38 | 22.79 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000144 | -1.2488 | 0.000211 | -2.205e-10 | 1.091e-08 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:08 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0x7638 | pH | 2021-03-09 14:32 | 0 | -1 | -1 | -1 | 1206401 | 180 | 341 | 27.38 | 22.79 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000108 | -1.2485 | 0.000169 | 9.307e-09 | 1.333e-08 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0x8BC4 | pH | 2021-03-09 14:33 | 0 | -1 | -1 | -1 | 1206402 | 180 | 341 | 27.39 | 22.74 | 11.716 | 0.0432 | 7.1234 | 3.62 | -1.2077 | 0.00012 | -1.2484 | 0.000147 | -2.652e-09 | -7.589e-09 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0xF83A | pH | 2021-03-09 14:34 | 0 | -1 | -1 | -1 | 1206403 | 180 | 341 | 27.39 | 22.76 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000158 | -1.2479 | 0.000233 | -6.853e-09 | -1.316e-10 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.9062 | 0.25 | 0 | 107.2122 | 0xB3D5 | pH | 2021-03-09 14:35 | 0 | -1 | -1 | -1 | 1206404 | 180 | 341 | 27.39 | 22.76 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000127 | -1.2477 | 0.000204 | -2.295e-09 | 2.079e-09 | 0.14581 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.233 | 107.3875 | 0xA7E1 | pH | 2021-03-09 14:36 | 0 | -1 | -1 | -1 | 1206405 | 180 | 341 | 27.39 | 22.74 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000113 | -1.2466 | 0.000182 | -3.153e-09 | 2.101e-09 | 0.14567 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.234 | 107.6255 | 0x8EE2 | pH | 2021-03-09 14:37 | 0 | -1 | -1 | -1 | 1206406 | 180 | 341 | 27.42 | 22.74 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000102 | -1.2453 | 0.000148 | -1.18e-08 | -4.783e-09 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.234 | 107.8639 | 0xF2E7 | pH | 2021-03-09 14:38 | 0 | -1 | -1 | -1 | 1206407 | 180 | 341 | 27.42 | 22.76 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000183 | -1.2442 | 0.000278 | -6.806e-09 | -9.353e-10 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.234 | 108.1024 | 0x20AB | pH | 2021-03-09 14:39 | 0 | -1 | -1 | -1 | 1206408 | 180 | 341 | 27.42 | 22.76 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000123 | -1.2432 | 0.000186 | 7.121e-10 | -1.643e-09 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.235 | 108.3412 | 0x8846 | pH | 2021-03-09 14:40 | 0 | -1 | -1 | -1 | 1206409 | 180 | 341 | 27.43 | 22.76 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000112 | -1.242 | 0.0002 | -4.199e-09 | 3.349e-09 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.235 | 108.5801 | 0x1D85 | pH | 2021-03-09 14:41 | 0 | -1 | -1 | -1 | 1206410 | 180 | 341 | 27.45 | 22.74 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000111 | -1.2408 | 0.000145 | -1.385e-08 | -3.687e-09 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:09 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.235 | 108.8194 | 0x129C | pH | 2021-03-09 14:42 | 0 | -1 | -1 | -1 | 1206411 | 180 | 341 | 27.45 | 22.74 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000119 | -1.2395 | 0.000169 | -1.022e-09 | 3.64e-09 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.235 | 109.0587 | 0x23CB | pH | 2021-03-09 14:43 | 0 | -1 | -1 | -1 | 1206412 | 180 | 341 | 27.45 | 22.74 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000145 | -1.2385 | 0.000207 | -5.378e-09 | 7.428e-09 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.8372 | 0.25 | -0.235 | 109.2985 | 0xA4EB | pH | 2021-03-09 14:44 | 0 | -1 | -1 | -1 | 1206413 | 180 | 341 | 27.46 | 22.74 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000114 | -1.2372 | 0.000152 | -2.718e-10 | 1.438e-08 | 0.14567 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.236 | 109.5385 | 0xADB9 | pH | 2021-03-09 14:45 | 0 | -1 | -1 | -1 | 1206414 | 180 | 341 | 27.46 | 22.71 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 7.9e-05 | -1.2352 | 0.000139 | -4.753e-09 | 4.299e-10 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.236 | 109.7787 | 0x57FA | pH | 2021-03-09 14:46 | 0 | -1 | -1 | -1 | 1206415 | 180 | 341 | 27.46 | 22.71 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000119 | -1.2338 | 0.000216 | -5.407e-09 | 6.551e-09 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.236 | 110.0193 | 0x0694 | pH | 2021-03-09 14:47 | 0 | -1 | -1 | -1 | 1206416 | 180 | 341 | 27.47 | 22.68 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000128 | -1.2318 | 0.000199 | 9.098e-09 | 8.752e-09 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.238 | 110.2611 | 0x9E7B | pH | 2021-03-09 14:48 | 0 | -1 | -1 | -1 | 1206417 | 180 | 341 | 27.47 | 22.68 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.0001 | -1.2293 | 0.000163 | 1.26e-09 | -1.661e-08 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.239 | 110.5043 | 0x3622 | pH | 2021-03-09 14:49 | 0 | -1 | -1 | -1 | 1206418 | 180 | 341 | 27.49 | 22.68 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.2e-05 | -1.2281 | 0.00023 | -1.891e-09 | -1.559e-09 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.002 | 110.5728 | 0xED97 | pH | 2021-03-09 14:50 | 0 | -1 | -1 | -1 | 1206419 | 180 | 341 | 27.49 | 22.63 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000111 | -1.2668 | 0.000157 | 1.612e-09 | 3.485e-09 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | -0.003 | 110.5763 | 0xA3B1 | pH | 2021-03-09 14:51 | 0 | -1 | -1 | -1 | 1206420 | 180 | 341 | 27.49 | 22.61 | 11.72 | 0.0432 | 7.1234 | 3.62 | -1.2077 | 9.3e-05 | -1.2549 | 0.000214 | -1.685e-09 | -1.294e-08 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | 0 | 110.5784 | 0x99C1 | pH | 2021-03-09 14:52 | 0 | -1 | -1 | -1 | 1206421 | 180 | 341 | 27.5 | 22.61 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000115 | -1.2437 | 0.000213 | 6.672e-09 | -6.077e-09 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:10 | 20 | 20 | 20 | 72.7682 | 0.25 | 0 | 110.58 | 0xD457 | pH | 2021-03-09 14:53 | 0 | -1 | -1 | -1 | 1206422 | 180 | 341 | 27.5 | 22.66 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 8.3e-05 | -1.2365 | 0.000162 | 1.624e-09 | -5.08e-09 | 0.14554 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.7682 | 0.25 | 0 | 110.5814 | 0x3FCC | pH | 2021-03-09 14:54 | 0 | -1 | -1 | -1 | 1206423 | 180 | 341 | 27.51 | 22.63 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000104 | -1.2317 | 0.000252 | -8.616e-09 | -1.125e-08 | 0.14554 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | -0.003 | 110.5827 | 0xBB18 | pH | 2021-03-09 14:55 | 0 | -1 | -1 | -1 | 1206424 | 180 | 341 | 27.53 | 22.56 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 0.000164 | -1.2281 | 0.000132 | -1.026e-08 | -4.615e-09 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | 0 | 110.5836 | 0xB54C | pH | 2021-03-09 14:56 | 0 | -1 | -1 | -1 | 1206425 | 180 | 341 | 27.51 | 22.59 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000121 | -1.2254 | 0.000102 | -5.199e-09 | -6.176e-09 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | 0 | 110.5844 | 0xFFA1 | pH | 2021-03-09 14:57 | 0 | -1 | -1 | -1 | 1206426 | 180 | 341 | 27.53 | 22.59 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 7.8e-05 | -1.2232 | 0.000167 | 4.98e-10 | 1.476e-10 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | 0 | 110.5853 | 0xEA6B | pH | 2021-03-09 14:58 | 0 | -1 | -1 | -1 | 1206427 | 180 | 341 | 27.54 | 22.59 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 8.6e-05 | -1.2215 | 0.000139 | -4.918e-09 | -4.434e-09 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | -0.001 | 110.5861 | 0x3E68 | pH | 2021-03-09 14:59 | 0 | -1 | -1 | -1 | 1206428 | 180 | 341 | 27.55 | 22.61 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000129 | -1.22 | 0.000179 | -1.227e-09 | 2.652e-09 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | 0 | 110.5867 | 0x24B2 | pH | 2021-03-09 15:00 | 0 | -1 | -1 | -1 | 1206429 | 180 | 341 | 27.55 | 22.59 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 8.6e-05 | -1.219 | 0.000126 | -6.464e-09 | -1.162e-08 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | 0 | 110.5873 | 0xA9FA | pH | 2021-03-09 15:01 | 0 | -1 | -1 | -1 | 1206430 | 180 | 341 | 27.55 | 22.56 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2078 | 0.000132 | -1.2179 | 0.00014 | -4.465e-09 | 4.656e-09 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6992 | 0.25 | -0.003 | 110.5881 | 0xADD5 | pH | 2021-03-09 15:02 | 0 | -1 | -1 | -1 | 1206431 | 180 | 341 | 27.55 | 22.59 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000147 | -1.2171 | 8.7e-05 | -3.158e-09 | -3.013e-09 | 0.1454 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:11 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5884 | 0xD5FC | pH | 2021-03-09 15:03 | 0 | -1 | -1 | -1 | 1206432 | 180 | 341 | 27.57 | 22.56 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 9.2e-05 | -1.2176 | 0.000127 | -8.921e-09 | 9.97e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5889 | 0x2EA7 | pH | 2021-03-09 15:04 | 0 | -1 | -1 | -1 | 1206433 | 180 | 341 | 27.57 | 22.53 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000132 | -1.2173 | 0.000198 | -4.656e-09 | 8.549e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5895 | 0xA409 | pH | 2021-03-09 15:05 | 0 | -1 | -1 | -1 | 1206434 | 180 | 341 | 27.57 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.62 | -1.2077 | 9.6e-05 | -1.2169 | 0.000146 | 1.471e-09 | 6.069e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5899 | 0x345A | pH | 2021-03-09 15:06 | 0 | -1 | -1 | -1 | 1206435 | 180 | 341 | 27.57 | 22.56 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000134 | -1.2163 | 0.000187 | -3.805e-09 | 1.608e-08 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5901 | 0x6417 | pH | 2021-03-09 15:07 | 0 | -1 | -1 | -1 | 1206436 | 180 | 341 | 27.57 | 22.53 | 11.72 | 0.0432 | 7.1234 | 3.621 | -1.2077 | 0.000121 | -1.2159 | 0.000128 | -3.037e-09 | -2.018e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5905 | 0x0EFF | pH | 2021-03-09 15:08 | 0 | -1 | -1 | -1 | 1206437 | 180 | 341 | 27.57 | 22.53 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000125 | -1.2149 | 0.000208 | -4.019e-09 | -9.399e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | -0.003 | 110.5908 | 0x5FC9 | pH | 2021-03-09 15:09 | 0 | -1 | -1 | -1 | 1206438 | 180 | 341 | 27.58 | 22.53 | 11.72 | 0.0432 | 7.1234 | 3.621 | -1.2077 | 0.000108 | -1.2143 | 0.000145 | -2.047e-09 | 2.727e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.591 | 0xA4B7 | pH | 2021-03-09 15:10 | 0 | -1 | -1 | -1 | 1206439 | 180 | 341 | 27.58 | 22.53 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 6.9e-05 | -1.2137 | 0.000159 | -4.357e-09 | 5.08e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.5912 | 0x9215 | pH | 2021-03-09 15:11 | 0 | -1 | -1 | -1 | 1206440 | 180 | 341 | 27.57 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.00011 | -1.2133 | 0.000151 | 7.426e-09 | 9.571e-09 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5914 | 0xE973 | pH | 2021-03-09 15:12 | 0 | -1 | -1 | -1 | 1206441 | 180 | 341 | 27.58 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 8.2e-05 | -1.2129 | 0.000177 | -7.131e-09 | 5.688e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.5915 | 0xEC3B | pH | 2021-03-09 15:13 | 0 | -1 | -1 | -1 | 1206442 | 180 | 341 | 27.57 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000158 | -1.2126 | 0.000174 | -4.684e-09 | -9.9e-09 | 0.14512 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:12 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.5918 | 0x28C2 | pH | 2021-03-09 15:14 | 0 | -1 | -1 | -1 | 1206443 | 180 | 341 | 27.57 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 8.4e-05 | -1.2207 | 0.000176 | 1.298e-09 | 1.824e-09 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.592 | 0x71E5 | pH | 2021-03-09 15:15 | 0 | -1 | -1 | -1 | 1206444 | 180 | 341 | 27.57 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 9.1e-05 | -1.2168 | 0.00017 | 1.685e-09 | 1.277e-08 | 0.14512 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.5922 | 0xD720 | pH | 2021-03-09 15:16 | 0 | -1 | -1 | -1 | 1206445 | 180 | 341 | 27.58 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 8.7e-05 | -1.2152 | 0.000134 | -3.953e-09 | -1.371e-08 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.5922 | 0x7842 | pH | 2021-03-09 15:17 | 0 | -1 | -1 | -1 | 1206446 | 180 | 341 | 27.57 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000133 | -1.2141 | 0.000231 | -7.239e-09 | -7.342e-09 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.6303 | 0.25 | 0 | 110.5922 | 0xE891 | pH | 2021-03-09 15:18 | 0 | -1 | -1 | -1 | 1206447 | 180 | 341 | 27.57 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000143 | -1.2131 | 0.000261 | 1.44e-09 | 9.196e-09 | 0.14526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0 | 110.5923 | 0xE170 | pH | 2021-03-09 15:19 | 0 | -1 | -1 | -1 | 1206448 | 180 | 341 | 27.57 | 22.48 | 11.72 | 0.0432 | 7.1234 | 3.621 | -1.2077 | 0.000102 | -1.2129 | 0.000191 | 2.764e-09 | 1.172e-09 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0.238 | 110.4133 | 0x072C | pH | 2021-03-09 15:20 | 0 | -1 | -1 | -1 | 1206449 | 180 | 341 | 27.57 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.0001 | -1.2159 | 0.000151 | 7.894e-10 | -5.95e-09 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0.237 | 110.1713 | 0x7EAC | pH | 2021-03-09 15:21 | 0 | -1 | -1 | -1 | 1206450 | 180 | 341 | 27.58 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000103 | -1.2214 | 0.000149 | -7.73e-09 | -1.577e-09 | 0.14512 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.5613 | 0.25 | 0.236 | 109.9305 | 0x2F95 | pH | 2021-03-09 15:22 | 0 | -1 | -1 | -1 | 1206451 | 180 | 341 | 27.57 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 8.8e-05 | -1.2251 | 8.9e-05 | 2.318e-09 | -2.48e-09 | 0.14512 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 109.6902 | 0xE0BC | pH | 2021-03-09 15:23 | 0 | -1 | -1 | -1 | 1206452 | 180 | 341 | 27.57 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 7e-05 | -1.2282 | 0.000129 | -4.142e-10 | -2.045e-09 | 0.14498 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:13 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 109.4504 | 0x5B15 | pH | 2021-03-09 15:24 | 0 | -1 | -1 | -1 | 1206453 | 180 | 341 | 27.57 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000145 | -1.2299 | 0.000162 | -4.618e-09 | 2.434e-09 | 0.14498 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 109.2108 | 0xA4C7 | pH | 2021-03-09 15:25 | 0 | -1 | -1 | -1 | 1206454 | 180 | 341 | 27.57 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000102 | -1.232 | 9.7e-05 | -1.302e-10 | 3.843e-09 | 0.14498 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 108.9712 | 0x4AA5 | pH | 2021-03-09 15:26 | 0 | -1 | -1 | -1 | 1206455 | 180 | 341 | 27.57 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000107 | -1.234 | 0.000141 | -5.387e-09 | -1.214e-08 | 0.14498 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 108.732 | 0x6D0A | pH | 2021-03-09 15:28 | 0 | -1 | -1 | -1 | 1206456 | 180 | 341 | 27.57 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000102 | -1.2361 | 0.000173 | -5.745e-09 | -1.418e-08 | 0.14498 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 108.4927 | 0xA056 | pH | 2021-03-09 15:29 | 0 | -1 | -1 | -1 | 1206457 | 180 | 341 | 27.57 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 8.6e-05 | -1.2377 | 0.000182 | 1.176e-09 | 8.382e-09 | 0.14498 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.234 | 108.2537 | 0x3AE8 | pH | 2021-03-09 15:30 | 0 | -1 | -1 | -1 | 1206458 | 180 | 341 | 27.55 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000125 | -1.2391 | 0.000169 | -4.832e-09 | 4.577e-09 | 0.14498 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.234 | 108.0149 | 0x6659 | pH | 2021-03-09 15:31 | 0 | -1 | -1 | -1 | 1206459 | 180 | 341 | 27.55 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000132 | -1.24 | 0.000179 | 4.412e-09 | -1.519e-08 | 0.14498 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.235 | 107.776 | 0x46D2 | pH | 2021-03-09 15:32 | 0 | -1 | -1 | -1 | 1206460 | 180 | 341 | 27.55 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 9.1e-05 | -1.2411 | 0.000157 | -1.107e-09 | 3.371e-09 | 0.14498 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4923 | 0.25 | 0.004 | 107.6833 | 0x75AA | pH | 2021-03-09 15:33 | 0 | -1 | -1 | -1 | 1206461 | 180 | 341 | 27.55 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000165 | -1.2421 | 0.000184 | -1.064e-08 | 4.926e-09 | 0.14498 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4233 | 0.25 | 0.001 | 107.6804 | 0x4546 | pH | 2021-03-09 15:34 | 0 | -1 | -1 | -1 | 1206462 | 180 | 341 | 27.55 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 7.8e-05 | -1.2422 | 0.000109 | -7.401e-09 | 5.061e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:14 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6749 | 0xD20D | pH | 2021-03-09 15:35 | 0 | -1 | -1 | -1 | 1206463 | 180 | 341 | 27.54 | 22.41 | 11.72 | 0.0432 | 7.1234 | 3.621 | -1.2077 | 9.5e-05 | -1.2425 | 0.000168 | -9.725e-10 | 6.068e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6736 | 0x5625 | pH | 2021-03-09 15:36 | 0 | -1 | -1 | -1 | 1206464 | 180 | 341 | 27.54 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000119 | -1.2423 | 0.000206 | -7.374e-09 | -8.483e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6727 | 0x4246 | pH | 2021-03-09 15:37 | 0 | -1 | -1 | -1 | 1206465 | 180 | 341 | 27.54 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000144 | -1.2422 | 0.000146 | -8.057e-10 | 7.329e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0.001 | 107.6718 | 0x8F1B | pH | 2021-03-09 15:38 | 0 | -1 | -1 | -1 | 1206466 | 180 | 341 | 27.54 | 22.38 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000134 | -1.2423 | 0.000241 | -3.25e-09 | -4.093e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6712 | 0xA285 | pH | 2021-03-09 15:39 | 0 | -1 | -1 | -1 | 1206467 | 180 | 341 | 27.54 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000127 | -1.2422 | 0.000188 | -3.33e-09 | -8.885e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6705 | 0x2194 | pH | 2021-03-09 15:40 | 0 | -1 | -1 | -1 | 1206468 | 180 | 341 | 27.54 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 5.2e-05 | -1.2418 | 0.000132 | -8.623e-09 | -1.555e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6699 | 0x82CE | pH | 2021-03-09 15:41 | 0 | -1 | -1 | -1 | 1206469 | 180 | 341 | 27.55 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000133 | -1.2415 | 0.00012 | 2.174e-09 | -5.577e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6695 | 0xED29 | pH | 2021-03-09 15:42 | 0 | -1 | -1 | -1 | 1206470 | 180 | 341 | 27.54 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 8.7e-05 | -1.2411 | 0.000201 | -1.888e-10 | 7.396e-10 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.669 | 0x8A2D | pH | 2021-03-09 15:43 | 0 | -1 | -1 | -1 | 1206471 | 180 | 341 | 27.53 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000109 | -1.2408 | 0.000224 | 8.912e-09 | -8.241e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6685 | 0x90BA | pH | 2021-03-09 15:44 | 0 | -1 | -1 | -1 | 1206472 | 180 | 341 | 27.53 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000129 | -1.2403 | 0.000221 | 1.211e-09 | 6.733e-10 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0 | 107.6681 | 0x8034 | pH | 2021-03-09 15:45 | 0 | -1 | -1 | -1 | 1206473 | 180 | 341 | 27.53 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 8.4e-05 | -1.2399 | 0.000154 | 3.872e-09 | -3.449e-10 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:15 | 20 | 20 | 20 | 72.4233 | 0.25 | 0.001 | 107.6676 | 0x5539 | pH | 2021-03-09 15:46 | 0 | -1 | -1 | -1 | 1206474 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 9.3e-05 | -1.2395 | 0.000178 | -7.539e-09 | 5.033e-09 | 0.14485 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6673 | 0x7041 | pH | 2021-03-09 15:47 | 0 | -1 | -1 | -1 | 1206475 | 180 | 341 | 27.53 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000106 | -1.2405 | 0.000134 | -5.817e-09 | 1.254e-08 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6672 | 0x7E8B | pH | 2021-03-09 15:48 | 0 | -1 | -1 | -1 | 1206476 | 180 | 341 | 27.53 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000137 | -1.24 | 0.000193 | -5.745e-10 | -3.676e-09 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6667 | 0xD58F | pH | 2021-03-09 15:49 | 0 | -1 | -1 | -1 | 1206477 | 180 | 341 | 27.53 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000132 | -1.2394 | 0.000193 | -2.978e-09 | -6.028e-09 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6665 | 0x9EF1 | pH | 2021-03-09 15:50 | 0 | -1 | -1 | -1 | 1206478 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000123 | -1.2387 | 0.000195 | -1.058e-08 | -4.581e-10 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6661 | 0x9930 | pH | 2021-03-09 15:51 | 0 | -1 | -1 | -1 | 1206479 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0432 | 7.1234 | 3.621 | -1.2078 | 0.000106 | -1.2381 | 0.000158 | -4.278e-09 | 8.13e-09 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.6657 | 0x8F1D | pH | 2021-03-09 15:52 | 0 | -1 | -1 | -1 | 1206480 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000126 | -1.2374 | 0.000197 | -9.721e-09 | 1.078e-08 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6655 | 0xA410 | pH | 2021-03-09 15:53 | 0 | -1 | -1 | -1 | 1206481 | 180 | 341 | 27.5 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 8.4e-05 | -1.2367 | 0.000168 | -3.922e-09 | -9.731e-10 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6652 | 0xFA2A | pH | 2021-03-09 15:54 | 0 | -1 | -1 | -1 | 1206482 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000145 | -1.236 | 0.000123 | -1.326e-08 | -1.064e-08 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.3544 | 0.25 | 0 | 107.665 | 0xD526 | pH | 2021-03-09 15:55 | 0 | -1 | -1 | -1 | 1206483 | 180 | 341 | 27.51 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 9.9e-05 | -1.2357 | 0.000124 | -1.729e-09 | -5.059e-10 | 0.14471 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:16 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6648 | 0x4165 | pH | 2021-03-09 15:56 | 0 | -1 | -1 | -1 | 1206484 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 9.4e-05 | -1.2353 | 0.000177 | 1.873e-09 | -7.715e-09 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6645 | 0x9000 | pH | 2021-03-09 15:57 | 0 | -1 | -1 | -1 | 1206485 | 180 | 341 | 27.5 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000105 | -1.2349 | 0.00015 | 5.395e-09 | -7.016e-09 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6643 | 0x1AB5 | pH | 2021-03-09 15:58 | 0 | -1 | -1 | -1 | 1206486 | 180 | 341 | 27.5 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 9e-05 | -1.2345 | 0.000191 | -4.965e-09 | -2.74e-09 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6641 | 0xF9FD | pH | 2021-03-09 15:59 | 0 | -1 | -1 | -1 | 1206487 | 180 | 341 | 27.5 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 7.8e-05 | -1.2339 | 0.000119 | 5.648e-09 | -1.029e-08 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6639 | 0xA3CD | pH | 2021-03-09 16:00 | 0 | -1 | -1 | -1 | 1206488 | 180 | 341 | 27.5 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000167 | -1.2335 | 9.8e-05 | -2.734e-09 | 1.462e-08 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6636 | 0xF9F6 | pH | 2021-03-09 16:01 | 0 | -1 | -1 | -1 | 1206489 | 180 | 341 | 27.51 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000148 | -1.2332 | 0.000189 | 1.895e-09 | 4.008e-09 | 0.14457 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6635 | 0x389B | pH | 2021-03-09 16:02 | 0 | -1 | -1 | -1 | 1206490 | 180 | 341 | 27.51 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000137 | -1.2331 | 0.000153 | -7.602e-09 | -1.468e-08 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6633 | 0xBC4E | pH | 2021-03-09 16:03 | 0 | -1 | -1 | -1 | 1206491 | 180 | 341 | 27.5 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 9.3e-05 | -1.2327 | 0.000226 | -3.798e-09 | -2.202e-09 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.663 | 0xC9AB | pH | 2021-03-09 16:04 | 0 | -1 | -1 | -1 | 1206492 | 180 | 341 | 27.5 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000155 | -1.2324 | 0.000193 | -9.891e-09 | 1.961e-09 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.663 | 0x7F95 | pH | 2021-03-09 16:05 | 0 | -1 | -1 | -1 | 1206493 | 180 | 341 | 27.49 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 9.1e-05 | -1.2324 | 7e-05 | 1.779e-09 | -3.19e-10 | 0.14457 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:17 | 20 | 20 | 20 | 72.2854 | 0.25 | 0 | 107.6627 | 0x1C51 | pH | 2021-03-09 16:06 | 0 | -1 | -1 | -1 | 1206494 | 180 | 341 | 27.49 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 6.4e-05 | -1.2322 | 0.00024 | -2.878e-09 | 5.251e-09 | 0.14457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6626 | 0xC0F6 | pH | 2021-03-09 16:07 | 0 | -1 | -1 | -1 | 1206495 | 180 | 341 | 27.49 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000162 | -1.2319 | 0.000174 | 1.22e-09 | -6.66e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6624 | 0xBD02 | pH | 2021-03-09 16:08 | 0 | -1 | -1 | -1 | 1206496 | 180 | 341 | 27.49 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000172 | -1.2316 | 0.000142 | 2.646e-09 | -1.167e-08 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6621 | 0x20A2 | pH | 2021-03-09 16:09 | 0 | -1 | -1 | -1 | 1206497 | 180 | 341 | 27.49 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000133 | -1.2313 | 0.000114 | 9.798e-10 | -6.714e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.662 | 0xD110 | pH | 2021-03-09 16:10 | 0 | -1 | -1 | -1 | 1206498 | 180 | 341 | 27.47 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000124 | -1.231 | 0.0002 | 7.633e-09 | -1.607e-08 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6618 | 0x3E19 | pH | 2021-03-09 16:11 | 0 | -1 | -1 | -1 | 1206499 | 180 | 341 | 27.47 | 22.41 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000136 | -1.2309 | 0.000141 | -6.217e-09 | 5.503e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6616 | 0xF215 | pH | 2021-03-09 16:12 | 0 | -1 | -1 | -1 | 1206500 | 180 | 341 | 27.47 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000113 | -1.2308 | 0.000192 | -1.896e-09 | 1.56e-09 | 0.14443 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6614 | 0x5BBB | pH | 2021-03-09 16:13 | 0 | -1 | -1 | -1 | 1206501 | 180 | 341 | 27.49 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000107 | -1.2304 | 9.7e-05 | 1.091e-08 | 5.329e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6613 | 0x2034 | pH | 2021-03-09 16:14 | 0 | -1 | -1 | -1 | 1206502 | 180 | 341 | 27.47 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 5.8e-05 | -1.2302 | 0.000235 | 1.025e-08 | -8.791e-10 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.6612 | 0x7E2D | pH | 2021-03-09 16:15 | 0 | -1 | -1 | -1 | 1206503 | 180 | 341 | 27.47 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000143 | -1.2299 | 0.000159 | -6.881e-09 | -5.624e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.661 | 0x6A6C | pH | 2021-03-09 16:16 | 0 | -1 | -1 | -1 | 1206504 | 180 | 341 | 27.47 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000105 | -1.2296 | 0.00015 | -3.051e-09 | -7.456e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:18 | 20 | 20 | 20 | 72.2164 | 0.25 | 0 | 107.661 | 0x4BE5 | pH | 2021-03-09 16:17 | 0 | -1 | -1 | -1 | 1206505 | 180 | 341 | 27.47 | 22.43 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000134 | -1.2293 | 0.000117 | 2.962e-09 | -5.411e-09 | 0.14443 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.661 | 0xF9B4 | pH | 2021-03-09 16:18 | 0 | -1 | -1 | -1 | 1206506 | 180 | 341 | 27.47 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.00014 | -1.2291 | 0.00019 | -8.551e-09 | 1.026e-08 | 0.14429 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4452 | -1.4259 | -1.4259 | -1.4259 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.661 | 0xBAAC | pH | 2021-03-09 16:19 | 0 | -1 | -1 | -1 | 1206507 | 180 | 341 | 27.47 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 8.4e-05 | -1.2289 | 0.000182 | -1.769e-09 | 1.279e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6609 | 0x52EB | pH | 2021-03-09 16:20 | 0 | -1 | -1 | -1 | 1206508 | 180 | 341 | 27.47 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000108 | -1.2288 | 0.000152 | -5.718e-09 | -5.929e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0xD579 | pH | 2021-03-09 16:21 | 0 | -1 | -1 | -1 | 1206509 | 180 | 341 | 27.46 | 22.46 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2077 | 0.000101 | -1.2286 | 0.00014 | -9.738e-09 | -9.587e-10 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0x9475 | pH | 2021-03-09 16:22 | 0 | -1 | -1 | -1 | 1206510 | 180 | 341 | 27.47 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000151 | -1.2285 | 0.000157 | 2.49e-10 | -3.103e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0xB488 | pH | 2021-03-09 16:23 | 0 | -1 | -1 | -1 | 1206511 | 180 | 341 | 27.47 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 9.6e-05 | -1.2283 | 0.00016 | -2.104e-09 | 7.449e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0x6125 | pH | 2021-03-09 16:24 | 0 | -1 | -1 | -1 | 1206512 | 180 | 341 | 27.46 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000115 | -1.228 | 0.000246 | -2.583e-09 | 6.477e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0x37C0 | pH | 2021-03-09 16:25 | 0 | -1 | -1 | -1 | 1206513 | 180 | 341 | 27.46 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000164 | -1.2278 | 0.000145 | 1.462e-09 | -2.933e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0xEEB7 | pH | 2021-03-09 16:26 | 0 | -1 | -1 | -1 | 1206514 | 180 | 341 | 27.47 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2078 | 0.000106 | -1.2275 | 0.00017 | -2.76e-09 | 1.929e-09 | 0.14429 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:19 | 20 | 20 | 20 | 72.1474 | 0.25 | 0 | 107.6608 | 0x31EF | pH | 2021-03-09 16:28 | 0 | -1 | -1 | -1 | 1206515 | 180 | 341 | 27.47 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2086 | 9.5e-05 | -1.2275 | 9.3e-05 | -8.589e-09 | 3.505e-09 | 0.14429 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 142.1567 | 0.25 | 0 | 107.6608 | 0x1F12 | pH | 2021-03-09 16:29 | 0 | -1 | -1 | -1 | 1206516 | 180 | 341 | 27.47 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2095 | 0.000114 | -1.2274 | 0.000156 | -2.233e-09 | 9.1e-09 | 0.28431 | -1.2095 | 0.90444 | 17.6664 | 0.23773 | -0.00024258 | -1.4469 | -1.4277 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 177.6096 | 0.25 | 0 | 107.6608 | 0x2443 | pH | 2021-03-09 16:30 | 0 | -1 | -1 | -1 | 1206517 | 180 | 341 | 27.47 | 22.48 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2095 | 0.000104 | -1.2271 | 0.000172 | 3.009e-09 | 1.396e-08 | 0.35522 | -1.2095 | 0.90444 | 17.6664 | 0.23773 | -0.00030765 | -1.4469 | -1.4276 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 179.334 | 0.25 | 0 | 107.6608 | 0x1E35 | pH | 2021-03-09 16:31 | 0 | -1 | -1 | -1 | 1206518 | 180 | 341 | 27.47 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2095 | 0.000122 | -1.227 | 0.000203 | -1.398e-08 | -2.101e-09 | 0.35867 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00031082 | -1.4469 | -1.4276 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 178.4373 | 0.25 | 0 | 107.6608 | 0x3AED | pH | 2021-03-09 16:32 | 0 | -1 | -1 | -1 | 1206519 | 180 | 341 | 27.47 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2095 | 0.000128 | -1.2267 | 0.000157 | 1.058e-11 | 9.851e-09 | 0.35687 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00030917 | -1.4469 | -1.4276 | -1.4276 | -1.4276 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 177.8165 | 0.25 | 0 | 107.6607 | 0xBA97 | pH | 2021-03-09 16:33 | 0 | -1 | -1 | -1 | 1206520 | 180 | 341 | 27.49 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2095 | 6.7e-05 | -1.2265 | 0.000262 | 6.917e-09 | -6.875e-09 | 0.35563 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00030803 | -1.4468 | -1.4276 | -1.4276 | -1.4276 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 177.4027 | 0.25 | 0 | 107.6607 | 0x3BAE | pH | 2021-03-09 16:34 | 0 | -1 | -1 | -1 | 1206521 | 180 | 341 | 27.49 | 22.51 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2095 | 0.000133 | -1.2264 | 0.00014 | -9.364e-09 | -9.61e-09 | 0.35481 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00030727 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 | 20 | 177.0578 | 0.25 | 0 | 107.6607 | 0x59F4 | pH | 2021-03-09 16:35 | 0 | -1 | -1 | -1 | 1206522 | 180 | 341 | 27.49 | 22.53 | 11.72 | 0.0434 | 7.1234 | 3.621 | -1.2094 | 0.000203 | -1.2263 | 0.000176 | 3.662e-11 | 7.519e-09 | 0.35412 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030664 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 |  | 176.7819 | [19 empty] | -1.2095 | [5 empty] | 0.35356 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00030613 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 |  | 176.575 | [19 empty] | -1.2094 | [5 empty] | 0.35315 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030576 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 |  | 176.3681 | [19 empty] | -1.2094 | [5 empty] | 0.35274 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00030538 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:20 | 20 | 20 |  | 176.1611 | [19 empty] | -1.2094 | [5 empty] | 0.35232 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.000305 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 176.0232 | [19 empty] | -1.2094 | [5 empty] | 0.35205 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030474 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.8852 | [19 empty] | -1.2094 | [5 empty] | 0.35177 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030449 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.7473 | [19 empty] | -1.2094 | [5 empty] | 0.35149 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030424 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.6093 | [19 empty] | -1.2094 | [5 empty] | 0.35122 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030398 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.5404 | [19 empty] | -1.2094 | [5 empty] | 0.35108 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030386 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.4024 | [19 empty] | -1.2094 | [5 empty] | 0.3508 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0003036 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.2645 | [19 empty] | -1.2094 | [5 empty] | 0.35053 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030335 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.1955 | [19 empty] | -1.2094 | [5 empty] | 0.35039 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030322 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.1265 | [19 empty] | -1.2094 | [5 empty] | 0.35025 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0003031 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:21 | 20 | 20 |  | 175.0575 | [19 empty] | -1.2094 | [5 empty] | 0.35012 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030297 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2094 | [5 empty] | 0.34984 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.8506 | [19 empty] | -1.2093 | [5 empty] | 0.3497 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030259 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2093 | [5 empty] | 0.34956 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2094 | [5 empty] | 0.34943 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2093 | [5 empty] | 0.34929 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2094 | [5 empty] | 0.34915 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2093 | [5 empty] | 0.34901 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2094 | [5 empty] | 0.34901 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2093 | [5 empty] | 0.34887 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2093 | [5 empty] | 0.34874 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2093 | [5 empty] | 0.34874 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2094 | [5 empty] | 0.3486 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2094 | [5 empty] | 0.34846 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2094 | [5 empty] | 0.34846 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2094 | [5 empty] | 0.34832 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.0919 | [19 empty] | -1.2093 | [5 empty] | 0.34818 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0003012 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.0919 | [19 empty] | -1.2094 | [5 empty] | 0.34818 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0003012 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2094 | [5 empty] | 0.34805 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2094 | [5 empty] | 0.34805 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:23 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2094 | [5 empty] | 0.34791 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.885 | [19 empty] | -1.2094 | [5 empty] | 0.34777 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.885 | [19 empty] | -1.2094 | [5 empty] | 0.34777 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.816 | [19 empty] | -1.2094 | [5 empty] | 0.34763 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030069 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.816 | [19 empty] | -1.2093 | [5 empty] | 0.34763 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030069 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.747 | [19 empty] | -1.2094 | [5 empty] | 0.34749 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.747 | [19 empty] | -1.2094 | [5 empty] | 0.34749 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.678 | [19 empty] | -1.2093 | [5 empty] | 0.34736 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.678 | [19 empty] | -1.2094 | [5 empty] | 0.34736 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2094 | [5 empty] | 0.34722 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2094 | [5 empty] | 0.34722 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:24 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2094 | [5 empty] | 0.34722 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2094 | [5 empty] | 0.34708 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2094 | [5 empty] | 0.34708 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2094 | [5 empty] | 0.34694 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2094 | [5 empty] | 0.34694 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2094 | [5 empty] | 0.3468 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2094 | [5 empty] | 0.3468 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2094 | [5 empty] | 0.3468 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2094 | [5 empty] | 0.34667 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2093 | [5 empty] | 0.34667 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:25 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2094 | [5 empty] | 0.34653 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2094 | [5 empty] | 0.34653 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2094 | [5 empty] | 0.34653 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2093 | [5 empty] | 0.34639 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2094 | [5 empty] | 0.34639 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2094 | [5 empty] | 0.34639 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2094 | [5 empty] | 0.34625 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2094 | [5 empty] | 0.34625 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2094 | [5 empty] | 0.34625 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.0573 | [19 empty] | -1.2094 | [5 empty] | 0.34611 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.0573 | [19 empty] | -1.2094 | [5 empty] | 0.34611 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:26 | 20 | 20 |  | 173.0573 | [19 empty] | -1.2094 | [5 empty] | 0.34611 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2094 | [5 empty] | 0.34598 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2094 | [5 empty] | 0.34598 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2093 | [5 empty] | 0.34598 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2093 | [5 empty] | 0.34598 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2093 | [5 empty] | 0.34584 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2093 | [5 empty] | 0.34584 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2094 | [5 empty] | 0.34584 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2094 | [5 empty] | 0.34584 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2094 | [5 empty] | 0.34584 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:27 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2094 | [5 empty] | 0.34584 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2094 | [5 empty] | 0.3457 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2094 | [5 empty] | 0.3457 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2094 | [5 empty] | 0.3457 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2094 | [5 empty] | 0.3457 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.7814 | [19 empty] | -1.2094 | [5 empty] | 0.34556 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029879 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.7814 | [19 empty] | -1.2094 | [5 empty] | 0.34556 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029879 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.7814 | [19 empty] | -1.2094 | [5 empty] | 0.34556 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029879 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.7124 | [19 empty] | -1.2094 | [5 empty] | 0.34542 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.7124 | [19 empty] | -1.2094 | [5 empty] | 0.34542 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:28 | 20 | 20 |  | 172.7124 | [19 empty] | -1.2093 | [5 empty] | 0.34542 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2093 | [5 empty] | 0.34529 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2094 | [5 empty] | 0.34529 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2094 | [5 empty] | 0.34529 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2094 | [5 empty] | 0.34529 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2094 | [5 empty] | 0.34529 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2094 | [5 empty] | 0.34529 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.5744 | [19 empty] | -1.2094 | [5 empty] | 0.34515 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.5744 | [19 empty] | -1.2094 | [5 empty] | 0.34515 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.5744 | [19 empty] | -1.2094 | [5 empty] | 0.34515 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.5744 | [19 empty] | -1.2094 | [5 empty] | 0.34515 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:29 | 20 | 20 |  | 172.5055 | [19 empty] | -1.2094 | [5 empty] | 0.34501 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.5055 | [19 empty] | -1.2094 | [5 empty] | 0.34501 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.5055 | [19 empty] | -1.2094 | [5 empty] | 0.34501 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.4365 | [19 empty] | -1.2094 | [5 empty] | 0.34487 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:30 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2094 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.3675 | [19 empty] | -1.2093 | [5 empty] | 0.34474 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:31 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2094 | [5 empty] | 0.3446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2296 | [19 empty] | -1.2094 | [5 empty] | 0.34446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2296 | [19 empty] | -1.2094 | [5 empty] | 0.34446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2296 | [19 empty] | -1.2094 | [5 empty] | 0.34446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2296 | [19 empty] | -1.2094 | [5 empty] | 0.34446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.2296 | [19 empty] | -1.2094 | [5 empty] | 0.34446 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.1606 | [19 empty] | -1.2094 | [5 empty] | 0.34432 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:32 | 20 | 20 |  | 172.1606 | [19 empty] | -1.2094 | [5 empty] | 0.34432 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.1606 | [19 empty] | -1.2093 | [5 empty] | 0.34432 | -1.2092 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.1606 | [19 empty] | -1.2094 | [5 empty] | 0.34432 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.1606 | [19 empty] | -1.2094 | [5 empty] | 0.34432 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:33 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0916 | [19 empty] | -1.2094 | [5 empty] | 0.34418 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2093 | [5 empty] | 0.34405 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2093 | [5 empty] | 0.34405 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2094 | [5 empty] | 0.34405 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2094 | [5 empty] | 0.34405 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2094 | [5 empty] | 0.34405 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:34 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2094 | [5 empty] | 0.34405 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2093 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.9537 | [19 empty] | -1.2094 | [5 empty] | 0.34391 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:35 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8847 | [19 empty] | -1.2094 | [5 empty] | 0.34377 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2093 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | -1.4274 | -1.4274 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:36 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2093 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:37 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2093 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2093 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2093 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:38 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.8157 | [19 empty] | -1.2094 | [5 empty] | 0.34363 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:39 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:40 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2095 | [5 empty] | 0.34336 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2095 | [5 empty] | 0.34336 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2093 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:41 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2094 | [5 empty] | 0.34349 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2093 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2093 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:42 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2093 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6778 | [19 empty] | -1.2094 | [5 empty] | 0.34336 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | -1.4275 | -1.4275 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:43 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2093 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4274 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2094 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:44 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4468 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2094 | [5 empty] | 0.34322 | -1.2093 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4467 | -1.4275 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 171.6088 | [19 empty] | -1.2117 | [5 empty] | 0.34322 | -1.2116 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4491 | -1.4298 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 330.8712 | [19 empty] | -1.2126 | [5 empty] | 0.66174 | -1.2125 | 0.90444 | 17.6664 | 0.23773 | -0.00058897 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 361.289 | [19 empty] | -1.2125 | [5 empty] | 0.72258 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00064481 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 359.7025 | [19 empty] | -1.2124 | [5 empty] | 0.71941 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00064189 | -1.4495 | -1.4302 | -1.4302 | -1.4302 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 358.9438 | [19 empty] | -1.2124 | [5 empty] | 0.71789 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006405 | -1.4494 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 358.392 | [19 empty] | -1.2124 | [5 empty] | 0.71678 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00063949 | -1.4494 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 357.9092 | [19 empty] | -1.2123 | [5 empty] | 0.71582 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006386 | -1.4494 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:45 | 20 | 20 |  | 357.5643 | [19 empty] | -1.2123 | [5 empty] | 0.71513 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063797 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 357.2884 | [19 empty] | -1.2123 | [5 empty] | 0.71458 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063746 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 357.0125 | [19 empty] | -1.2123 | [5 empty] | 0.71403 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063696 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 356.8056 | [19 empty] | -1.2123 | [5 empty] | 0.71361 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063658 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 356.5987 | [19 empty] | -1.2123 | [5 empty] | 0.7132 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006362 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 356.3918 | [19 empty] | -1.2123 | [5 empty] | 0.71278 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063582 | -1.4493 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 356.1848 | [19 empty] | -1.2123 | [5 empty] | 0.71237 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063544 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 356.0469 | [19 empty] | -1.2123 | [5 empty] | 0.71209 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063518 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 355.9089 | [19 empty] | -1.2123 | [5 empty] | 0.71182 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063493 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 355.771 | [19 empty] | -1.2123 | [5 empty] | 0.71154 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063468 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 355.633 | [19 empty] | -1.2123 | [5 empty] | 0.71127 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063442 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:46 | 20 | 20 |  | 355.4951 | [19 empty] | -1.2123 | [5 empty] | 0.71099 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063417 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 355.4261 | [19 empty] | -1.2123 | [5 empty] | 0.71085 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063404 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 355.2882 | [19 empty] | -1.2123 | [5 empty] | 0.71058 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063379 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 355.2192 | [19 empty] | -1.2122 | [5 empty] | 0.71044 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063366 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 354.8743 | [19 empty] | -1.2123 | [5 empty] | 0.70975 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063303 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 357.8402 | [19 empty] | -1.2123 | [5 empty] | 0.71568 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063847 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 357.7713 | [19 empty] | -1.2123 | [5 empty] | 0.71554 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00063835 | -1.4494 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 357.7023 | [19 empty] | -1.2123 | [5 empty] | 0.7154 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00063822 | -1.4494 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 357.6333 | [19 empty] | -1.2123 | [5 empty] | 0.71527 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063809 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 357.5643 | [19 empty] | -1.2123 | [5 empty] | 0.71513 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063797 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:47 | 20 | 20 |  | 357.4264 | [19 empty] | -1.2123 | [5 empty] | 0.71485 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063772 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 357.3574 | [19 empty] | -1.2123 | [5 empty] | 0.71471 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063759 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 357.2884 | [19 empty] | -1.2123 | [5 empty] | 0.71458 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063746 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 357.2195 | [19 empty] | -1.2123 | [5 empty] | 0.71444 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063734 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 357.1505 | [19 empty] | -1.2123 | [5 empty] | 0.7143 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063721 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 357.0815 | [19 empty] | -1.2123 | [5 empty] | 0.71416 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063708 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 357.0125 | [19 empty] | -1.2123 | [5 empty] | 0.71403 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063696 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 356.9436 | [19 empty] | -1.2123 | [5 empty] | 0.71389 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063683 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 356.8746 | [19 empty] | -1.2122 | [5 empty] | 0.71375 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006367 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 356.8056 | [19 empty] | -1.2123 | [5 empty] | 0.71361 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063658 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:48 | 20 | 20 |  | 356.8056 | [19 empty] | -1.2123 | [5 empty] | 0.71361 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063658 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.6677 | [19 empty] | -1.2123 | [5 empty] | 0.71334 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063632 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.6677 | [19 empty] | -1.2123 | [5 empty] | 0.71334 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063632 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.5987 | [19 empty] | -1.2122 | [5 empty] | 0.7132 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006362 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.5297 | [19 empty] | -1.2123 | [5 empty] | 0.71306 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063607 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.5297 | [19 empty] | -1.2123 | [5 empty] | 0.71306 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063607 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.4607 | [19 empty] | -1.2123 | [5 empty] | 0.71292 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063594 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.3918 | [19 empty] | -1.2123 | [5 empty] | 0.71278 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063582 | -1.4493 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.3918 | [19 empty] | -1.2123 | [5 empty] | 0.71278 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00063582 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.3228 | [19 empty] | -1.2123 | [5 empty] | 0.71265 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063569 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.2538 | [19 empty] | -1.2123 | [5 empty] | 0.71251 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063556 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:49 | 20 | 20 |  | 356.2538 | [19 empty] | -1.2123 | [5 empty] | 0.71251 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063556 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 356.1848 | [19 empty] | -1.2123 | [5 empty] | 0.71237 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063544 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 356.1159 | [19 empty] | -1.2123 | [5 empty] | 0.71223 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063531 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 356.1159 | [19 empty] | -1.2123 | [5 empty] | 0.71223 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063531 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 356.0469 | [19 empty] | -1.2122 | [5 empty] | 0.71209 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063518 | -1.4492 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 356.0469 | [19 empty] | -1.2123 | [5 empty] | 0.71209 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063518 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 355.9779 | [19 empty] | -1.2123 | [5 empty] | 0.71196 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063506 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 355.9779 | [19 empty] | -1.2122 | [5 empty] | 0.71196 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063506 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 355.9089 | [19 empty] | -1.2123 | [5 empty] | 0.71182 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063493 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 355.9089 | [19 empty] | -1.2123 | [5 empty] | 0.71182 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063493 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:50 | 20 | 20 |  | 355.84 | [19 empty] | -1.2123 | [5 empty] | 0.71168 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006348 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.771 | [19 empty] | -1.2123 | [5 empty] | 0.71154 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063468 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.702 | [19 empty] | -1.2123 | [5 empty] | 0.7114 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063455 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.702 | [19 empty] | -1.2123 | [5 empty] | 0.7114 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063455 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.633 | [19 empty] | -1.2122 | [5 empty] | 0.71127 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063442 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.633 | [19 empty] | -1.2123 | [5 empty] | 0.71127 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063442 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.5641 | [19 empty] | -1.2122 | [5 empty] | 0.71113 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006343 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.5641 | [19 empty] | -1.2122 | [5 empty] | 0.71113 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006343 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.4951 | [19 empty] | -1.2122 | [5 empty] | 0.71099 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063417 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.4951 | [19 empty] | -1.2122 | [5 empty] | 0.71099 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063417 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:51 | 20 | 20 |  | 355.4261 | [19 empty] | -1.2122 | [5 empty] | 0.71085 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063404 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.4261 | [19 empty] | -1.2122 | [5 empty] | 0.71085 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063404 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.3571 | [19 empty] | -1.2122 | [5 empty] | 0.71071 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063392 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.3571 | [19 empty] | -1.2122 | [5 empty] | 0.71071 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063392 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.2882 | [19 empty] | -1.2122 | [5 empty] | 0.71058 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063379 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.2882 | [19 empty] | -1.2122 | [5 empty] | 0.71058 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063379 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.2192 | [19 empty] | -1.2122 | [5 empty] | 0.71044 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063366 | -1.4492 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.2192 | [19 empty] | -1.2122 | [5 empty] | 0.71044 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063366 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.2192 | [19 empty] | -1.2122 | [5 empty] | 0.71044 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063366 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.1502 | [19 empty] | -1.2122 | [5 empty] | 0.7103 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063354 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.1502 | [19 empty] | -1.2122 | [5 empty] | 0.7103 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063354 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:52 | 20 | 20 |  | 355.1502 | [19 empty] | -1.2122 | [5 empty] | 0.7103 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063354 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 355.0812 | [19 empty] | -1.2122 | [5 empty] | 0.71016 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063341 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 355.0812 | [19 empty] | -1.2122 | [5 empty] | 0.71016 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063341 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 355.0123 | [19 empty] | -1.2122 | [5 empty] | 0.71002 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063328 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 355.0123 | [19 empty] | -1.2123 | [5 empty] | 0.71002 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063328 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 355.0123 | [19 empty] | -1.2122 | [5 empty] | 0.71002 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063328 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 354.9433 | [19 empty] | -1.2122 | [5 empty] | 0.70989 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063316 | -1.4492 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 354.9433 | [19 empty] | -1.2122 | [5 empty] | 0.70989 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063316 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 354.8743 | [19 empty] | -1.2122 | [5 empty] | 0.70975 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063303 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 354.8743 | [19 empty] | -1.2122 | [5 empty] | 0.70975 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063303 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:53 | 20 | 20 |  | 354.8053 | [19 empty] | -1.2122 | [5 empty] | 0.70961 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006329 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.8053 | [19 empty] | -1.2122 | [5 empty] | 0.70961 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006329 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.8053 | [19 empty] | -1.2122 | [5 empty] | 0.70961 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006329 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.7364 | [19 empty] | -1.2122 | [5 empty] | 0.70947 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063278 | -1.4492 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.7364 | [19 empty] | -1.2122 | [5 empty] | 0.70947 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063278 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.7364 | [19 empty] | -1.2122 | [5 empty] | 0.70947 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063278 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.7364 | [19 empty] | -1.2122 | [5 empty] | 0.70947 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063278 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.6674 | [19 empty] | -1.2122 | [5 empty] | 0.70933 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063265 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.6674 | [19 empty] | -1.2122 | [5 empty] | 0.70933 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063265 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.6674 | [19 empty] | -1.2122 | [5 empty] | 0.70933 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063265 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.6674 | [19 empty] | -1.2122 | [5 empty] | 0.70933 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063265 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:54 | 20 | 20 |  | 354.5984 | [19 empty] | -1.2122 | [5 empty] | 0.7092 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063252 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5984 | [19 empty] | -1.2122 | [5 empty] | 0.7092 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063252 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5984 | [19 empty] | -1.2122 | [5 empty] | 0.7092 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063252 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5984 | [19 empty] | -1.2123 | [5 empty] | 0.7092 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063252 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5294 | [19 empty] | -1.2122 | [5 empty] | 0.70906 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006324 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5294 | [19 empty] | -1.2123 | [5 empty] | 0.70906 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006324 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5294 | [19 empty] | -1.2122 | [5 empty] | 0.70906 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006324 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.5294 | [19 empty] | -1.2122 | [5 empty] | 0.70906 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.0006324 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.4605 | [19 empty] | -1.2122 | [5 empty] | 0.70892 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063227 | -1.4492 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.4605 | [19 empty] | -1.2122 | [5 empty] | 0.70892 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063227 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:55 | 20 | 20 |  | 354.4605 | [19 empty] | -1.2122 | [5 empty] | 0.70892 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063227 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3915 | [19 empty] | -1.2122 | [5 empty] | 0.70878 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063214 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3915 | [19 empty] | -1.2122 | [5 empty] | 0.70878 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063214 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3915 | [19 empty] | -1.2122 | [5 empty] | 0.70878 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063214 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3915 | [19 empty] | -1.2122 | [5 empty] | 0.70878 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063214 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3225 | [19 empty] | -1.2122 | [5 empty] | 0.70865 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063202 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3225 | [19 empty] | -1.2122 | [5 empty] | 0.70865 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063202 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3225 | [19 empty] | -1.2122 | [5 empty] | 0.70865 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063202 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3225 | [19 empty] | -1.2122 | [5 empty] | 0.70865 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063202 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3225 | [19 empty] | -1.2122 | [5 empty] | 0.70865 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063202 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:56 | 20 | 20 |  | 354.3225 | [19 empty] | -1.2123 | [5 empty] | 0.70865 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063202 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2122 | [5 empty] | 0.70851 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2122 | [5 empty] | 0.70851 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2122 | [5 empty] | 0.70851 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2123 | [5 empty] | 0.70851 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2123 | [5 empty] | 0.70851 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2123 | [5 empty] | 0.70851 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.2535 | [19 empty] | -1.2123 | [5 empty] | 0.70851 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063189 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.1846 | [19 empty] | -1.2123 | [5 empty] | 0.70837 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063176 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.1846 | [19 empty] | -1.2122 | [5 empty] | 0.70837 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063176 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:57 | 20 | 20 |  | 354.1846 | [19 empty] | -1.2122 | [5 empty] | 0.70837 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063176 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1846 | [19 empty] | -1.2123 | [5 empty] | 0.70837 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063176 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2122 | [5 empty] | 0.70823 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2122 | [5 empty] | 0.70823 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2122 | [5 empty] | 0.70823 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2123 | [5 empty] | 0.70823 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2122 | [5 empty] | 0.70823 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2122 | [5 empty] | 0.70823 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.1156 | [19 empty] | -1.2122 | [5 empty] | 0.70823 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063164 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:58 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2123 | [5 empty] | 0.70809 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 354.0466 | [19 empty] | -1.2122 | [5 empty] | 0.70809 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063151 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 353.9776 | [19 empty] | -1.2122 | [5 empty] | 0.70796 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063138 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 353.9776 | [19 empty] | -1.2122 | [5 empty] | 0.70796 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063138 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 14:59 | 20 | 20 |  | 353.9776 | [19 empty] | -1.2122 | [5 empty] | 0.70796 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063138 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9776 | [19 empty] | -1.2122 | [5 empty] | 0.70796 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063138 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9776 | [19 empty] | -1.2122 | [5 empty] | 0.70796 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063138 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9776 | [19 empty] | -1.2122 | [5 empty] | 0.70796 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063138 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9087 | [19 empty] | -1.2123 | [5 empty] | 0.70782 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063126 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9087 | [19 empty] | -1.2122 | [5 empty] | 0.70782 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063126 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9087 | [19 empty] | -1.2123 | [5 empty] | 0.70782 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063126 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.9087 | [19 empty] | -1.2122 | [5 empty] | 0.70782 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063126 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:00 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2123 | [5 empty] | 0.70768 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2123 | [5 empty] | 0.70768 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.8397 | [19 empty] | -1.2122 | [5 empty] | 0.70768 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063113 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:01 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2123 | [5 empty] | 0.70754 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2123 | [5 empty] | 0.70754 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7707 | [19 empty] | -1.2122 | [5 empty] | 0.70754 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063101 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:02 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2123 | [5 empty] | 0.7074 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.7017 | [19 empty] | -1.2122 | [5 empty] | 0.7074 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063088 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:03 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:04 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:05 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:06 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | -1.43 | -1.43 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:07 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4492 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2123 | [5 empty] | 0.70713 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2123 | [5 empty] | 0.70727 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4492 | -1.4299 | -1.4299 | -1.4299 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2122 | [5 empty] | 0.70713 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:08 | 20 | 20 |  | 353.6328 | [19 empty] | -1.2122 | [5 empty] | 0.70727 | -1.2121 | 0.90444 | 17.6664 | 0.23773 | -0.00063075 | -1.4492 | -1.4299 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 353.5638 | [19 empty] | -1.2133 | [5 empty] | 0.70713 | -1.2132 | 0.90444 | 17.6664 | 0.23773 | -0.00063063 | -1.4503 | -1.4311 | -1.4311 | -1.4311 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 445.438 | [19 empty] | -1.2149 | [5 empty] | 0.89088 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00079926 | -1.4517 | -1.4324 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 529.1731 | [19 empty] | -1.2149 | [5 empty] | 1.0583 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00095296 | -1.4516 | -1.4323 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 527.7936 | [19 empty] | -1.2149 | [5 empty] | 1.0556 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00095043 | -1.4516 | -1.4323 | -1.4323 | -1.4323 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 527.2418 | [19 empty] | -1.2148 | [5 empty] | 1.0545 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094942 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 526.828 | [19 empty] | -1.2148 | [5 empty] | 1.0537 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00094866 | -1.4515 | -1.4323 | -1.4323 | -1.4323 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 526.5521 | [19 empty] | -1.2148 | [5 empty] | 1.0531 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094815 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 527.4488 | [19 empty] | -1.2148 | [5 empty] | 1.0549 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009498 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 527.3108 | [19 empty] | -1.2149 | [5 empty] | 1.0546 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00094954 | -1.4515 | -1.4323 | -1.4323 | -1.4323 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:09 | 20 | 20 |  | 527.1039 | [19 empty] | -1.2148 | [5 empty] | 1.0542 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094917 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.897 | [19 empty] | -1.2148 | [5 empty] | 1.0538 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094879 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.759 | [19 empty] | -1.2148 | [5 empty] | 1.0535 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094853 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.6211 | [19 empty] | -1.2148 | [5 empty] | 1.0532 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094828 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.4831 | [19 empty] | -1.2148 | [5 empty] | 1.053 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094803 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.4141 | [19 empty] | -1.2147 | [5 empty] | 1.0528 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009479 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.2762 | [19 empty] | -1.2148 | [5 empty] | 1.0526 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094765 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.2072 | [19 empty] | -1.2148 | [5 empty] | 1.0524 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094752 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.1382 | [19 empty] | -1.2147 | [5 empty] | 1.0523 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094739 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 526.0003 | [19 empty] | -1.2148 | [5 empty] | 1.052 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094714 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:10 | 20 | 20 |  | 525.9313 | [19 empty] | -1.2148 | [5 empty] | 1.0519 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094701 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.7934 | [19 empty] | -1.2148 | [5 empty] | 1.0516 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094676 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.7934 | [19 empty] | -1.2148 | [5 empty] | 1.0516 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.6554 | [19 empty] | -1.2148 | [5 empty] | 1.0513 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094651 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.5865 | [19 empty] | -1.2148 | [5 empty] | 1.0512 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094638 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.5175 | [19 empty] | -1.2147 | [5 empty] | 1.051 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094625 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.4485 | [19 empty] | -1.2147 | [5 empty] | 1.0509 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094613 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.3795 | [19 empty] | -1.2147 | [5 empty] | 1.0508 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.000946 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.3106 | [19 empty] | -1.2148 | [5 empty] | 1.0506 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094587 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.3106 | [19 empty] | -1.2148 | [5 empty] | 1.0506 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094587 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.2416 | [19 empty] | -1.2148 | [5 empty] | 1.0505 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094575 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:11 | 20 | 20 |  | 525.1726 | [19 empty] | -1.2147 | [5 empty] | 1.0503 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094562 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 525.1036 | [19 empty] | -1.2147 | [5 empty] | 1.0502 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094549 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 525.0347 | [19 empty] | -1.2147 | [5 empty] | 1.0501 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094537 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 525.0347 | [19 empty] | -1.2147 | [5 empty] | 1.0501 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094537 | -1.4514 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.9657 | [19 empty] | -1.2147 | [5 empty] | 1.0499 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094524 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.8967 | [19 empty] | -1.2147 | [5 empty] | 1.0498 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094511 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.8967 | [19 empty] | -1.2148 | [5 empty] | 1.0498 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094511 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.8277 | [19 empty] | -1.2148 | [5 empty] | 1.0497 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094499 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.7588 | [19 empty] | -1.2148 | [5 empty] | 1.0495 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094486 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.7588 | [19 empty] | -1.2148 | [5 empty] | 1.0495 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094486 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:12 | 20 | 20 |  | 524.6898 | [19 empty] | -1.2148 | [5 empty] | 1.0494 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094473 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.6208 | [19 empty] | -1.2148 | [5 empty] | 1.0492 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094461 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.6208 | [19 empty] | -1.2148 | [5 empty] | 1.0492 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094461 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.5518 | [19 empty] | -1.2148 | [5 empty] | 1.0491 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094448 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.4829 | [19 empty] | -1.2147 | [5 empty] | 1.049 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094435 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.4829 | [19 empty] | -1.2148 | [5 empty] | 1.049 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094435 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.4139 | [19 empty] | -1.2148 | [5 empty] | 1.0488 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094423 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.4139 | [19 empty] | -1.2148 | [5 empty] | 1.0488 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094423 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.2759 | [19 empty] | -1.2147 | [5 empty] | 1.0486 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094397 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.2759 | [19 empty] | -1.2148 | [5 empty] | 1.0486 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094397 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:13 | 20 | 20 |  | 524.2759 | [19 empty] | -1.2148 | [5 empty] | 1.0486 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094397 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524.207 | [19 empty] | -1.2148 | [5 empty] | 1.0484 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094385 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524.138 | [19 empty] | -1.2148 | [5 empty] | 1.0483 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094372 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524.138 | [19 empty] | -1.2148 | [5 empty] | 1.0483 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094372 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524.069 | [19 empty] | -1.2147 | [5 empty] | 1.0481 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094359 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524 | [19 empty] | -1.2147 | [5 empty] | 1.048 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00094347 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524 | [19 empty] | -1.2148 | [5 empty] | 1.048 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094347 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 524 | [19 empty] | -1.2148 | [5 empty] | 1.048 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094347 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 523.9311 | [19 empty] | -1.2148 | [5 empty] | 1.0479 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094334 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 523.8621 | [19 empty] | -1.2148 | [5 empty] | 1.0477 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094321 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 523.8621 | [19 empty] | -1.2148 | [5 empty] | 1.0477 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094321 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:14 | 20 | 20 |  | 523.7931 | [19 empty] | -1.2148 | [5 empty] | 1.0476 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094309 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.7931 | [19 empty] | -1.2148 | [5 empty] | 1.0476 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094309 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.7931 | [19 empty] | -1.2148 | [5 empty] | 1.0476 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094309 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.7241 | [19 empty] | -1.2148 | [5 empty] | 1.0474 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094296 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.7241 | [19 empty] | -1.2148 | [5 empty] | 1.0474 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094296 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.6552 | [19 empty] | -1.2149 | [5 empty] | 1.0473 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00094283 | -1.4515 | -1.4323 | -1.4323 | -1.4323 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.6552 | [19 empty] | -1.2148 | [5 empty] | 1.0473 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094283 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.5862 | [19 empty] | -1.2148 | [5 empty] | 1.0472 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094271 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.5862 | [19 empty] | -1.2148 | [5 empty] | 1.0472 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094271 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.5862 | [19 empty] | -1.2148 | [5 empty] | 1.0472 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094271 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:15 | 20 | 20 |  | 523.5172 | [19 empty] | -1.2148 | [5 empty] | 1.047 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094258 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.5172 | [19 empty] | -1.2148 | [5 empty] | 1.047 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094258 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.4482 | [19 empty] | -1.2148 | [5 empty] | 1.0469 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094245 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.4482 | [19 empty] | -1.2148 | [5 empty] | 1.0469 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094245 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.3793 | [19 empty] | -1.2148 | [5 empty] | 1.0468 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094233 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.3793 | [19 empty] | -1.2148 | [5 empty] | 1.0468 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094233 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.3103 | [19 empty] | -1.2148 | [5 empty] | 1.0466 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009422 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.3103 | [19 empty] | -1.2148 | [5 empty] | 1.0466 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009422 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.3103 | [19 empty] | -1.2148 | [5 empty] | 1.0466 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009422 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.3103 | [19 empty] | -1.2148 | [5 empty] | 1.0466 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009422 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:16 | 20 | 20 |  | 523.2413 | [19 empty] | -1.2148 | [5 empty] | 1.0465 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094208 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.1723 | [19 empty] | -1.2148 | [5 empty] | 1.0463 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094195 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.1723 | [19 empty] | -1.2148 | [5 empty] | 1.0463 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094195 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.1723 | [19 empty] | -1.2148 | [5 empty] | 1.0463 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094195 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.1034 | [19 empty] | -1.2148 | [5 empty] | 1.0462 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094182 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.1034 | [19 empty] | -1.2148 | [5 empty] | 1.0462 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094182 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.1034 | [19 empty] | -1.2148 | [5 empty] | 1.0462 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094182 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 523.0344 | [19 empty] | -1.2148 | [5 empty] | 1.0461 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009417 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 522.9654 | [19 empty] | -1.2148 | [5 empty] | 1.0459 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094157 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 522.9654 | [19 empty] | -1.2148 | [5 empty] | 1.0459 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094157 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 522.9654 | [19 empty] | -1.2148 | [5 empty] | 1.0459 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094157 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:17 | 20 | 20 |  | 522.8964 | [19 empty] | -1.2148 | [5 empty] | 1.0458 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094144 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.8964 | [19 empty] | -1.2148 | [5 empty] | 1.0458 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094144 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.8964 | [19 empty] | -1.2148 | [5 empty] | 1.0458 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094144 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.8964 | [19 empty] | -1.2148 | [5 empty] | 1.0458 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094144 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.8275 | [19 empty] | -1.2148 | [5 empty] | 1.0457 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094132 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.8275 | [19 empty] | -1.2148 | [5 empty] | 1.0457 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094132 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.8275 | [19 empty] | -1.2148 | [5 empty] | 1.0457 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094132 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.7585 | [19 empty] | -1.2148 | [5 empty] | 1.0455 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094119 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.7585 | [19 empty] | -1.2148 | [5 empty] | 1.0455 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094119 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.7585 | [19 empty] | -1.2148 | [5 empty] | 1.0455 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094119 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:18 | 20 | 20 |  | 522.6895 | [19 empty] | -1.2148 | [5 empty] | 1.0454 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094106 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.6895 | [19 empty] | -1.2148 | [5 empty] | 1.0454 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094106 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.6895 | [19 empty] | -1.2148 | [5 empty] | 1.0454 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094106 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.6895 | [19 empty] | -1.2148 | [5 empty] | 1.0454 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094106 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.6205 | [19 empty] | -1.2148 | [5 empty] | 1.0452 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094094 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.6205 | [19 empty] | -1.2148 | [5 empty] | 1.0452 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094094 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.6205 | [19 empty] | -1.2148 | [5 empty] | 1.0452 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094094 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.5516 | [19 empty] | -1.2148 | [5 empty] | 1.0451 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094081 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.5516 | [19 empty] | -1.2148 | [5 empty] | 1.0451 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094081 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.5516 | [19 empty] | -1.2148 | [5 empty] | 1.0451 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094081 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:19 | 20 | 20 |  | 522.5516 | [19 empty] | -1.2148 | [5 empty] | 1.0451 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094081 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.4826 | [19 empty] | -1.2148 | [5 empty] | 1.045 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094068 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.4826 | [19 empty] | -1.2148 | [5 empty] | 1.045 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094068 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.4826 | [19 empty] | -1.2148 | [5 empty] | 1.045 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094068 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.4136 | [19 empty] | -1.2148 | [5 empty] | 1.0448 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094056 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.4136 | [19 empty] | -1.2148 | [5 empty] | 1.0448 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094056 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.4136 | [19 empty] | -1.2148 | [5 empty] | 1.0448 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094056 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.3446 | [19 empty] | -1.2148 | [5 empty] | 1.0447 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094043 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.3446 | [19 empty] | -1.2148 | [5 empty] | 1.0447 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094043 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.2757 | [19 empty] | -1.2148 | [5 empty] | 1.0446 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009403 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.2757 | [19 empty] | -1.2148 | [5 empty] | 1.0446 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009403 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:20 | 20 | 20 |  | 522.2757 | [19 empty] | -1.2148 | [5 empty] | 1.0446 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009403 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.2757 | [19 empty] | -1.2148 | [5 empty] | 1.0446 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009403 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.2757 | [19 empty] | -1.2148 | [5 empty] | 1.0446 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009403 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.2757 | [19 empty] | -1.2148 | [5 empty] | 1.0446 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009403 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.2067 | [19 empty] | -1.2148 | [5 empty] | 1.0444 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094018 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.2067 | [19 empty] | -1.2148 | [5 empty] | 1.0444 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094018 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.2067 | [19 empty] | -1.2148 | [5 empty] | 1.0444 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094018 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.1377 | [19 empty] | -1.2148 | [5 empty] | 1.0443 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094005 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.1377 | [19 empty] | -1.2148 | [5 empty] | 1.0443 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094005 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.1377 | [19 empty] | -1.2148 | [5 empty] | 1.0443 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094005 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:21 | 20 | 20 |  | 522.1377 | [19 empty] | -1.2148 | [5 empty] | 1.0443 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094005 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 522.1377 | [19 empty] | -1.2148 | [5 empty] | 1.0443 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00094005 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 522.0687 | [19 empty] | -1.2148 | [5 empty] | 1.0441 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093992 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 522.0687 | [19 empty] | -1.2148 | [5 empty] | 1.0441 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093992 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 522.0687 | [19 empty] | -1.2148 | [5 empty] | 1.0441 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093992 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 522.0687 | [19 empty] | -1.2148 | [5 empty] | 1.0441 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093992 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 522.0687 | [19 empty] | -1.2148 | [5 empty] | 1.0441 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093992 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 521.9998 | [19 empty] | -1.2148 | [5 empty] | 1.044 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009398 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 521.9998 | [19 empty] | -1.2147 | [5 empty] | 1.044 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009398 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 521.9998 | [19 empty] | -1.2147 | [5 empty] | 1.044 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009398 | -1.4514 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 521.9998 | [19 empty] | -1.2148 | [5 empty] | 1.044 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009398 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:22 | 20 | 20 |  | 521.9998 | [19 empty] | -1.2147 | [5 empty] | 1.044 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009398 | -1.4514 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2147 | [5 empty] | 1.0439 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2147 | [5 empty] | 1.0439 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4514 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2148 | [5 empty] | 1.0439 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2148 | [5 empty] | 1.0439 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2148 | [5 empty] | 1.0439 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2148 | [5 empty] | 1.0439 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.9308 | [19 empty] | -1.2148 | [5 empty] | 1.0439 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093967 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:23 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.8618 | [19 empty] | -1.2148 | [5 empty] | 1.0437 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093954 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.7928 | [19 empty] | -1.2148 | [5 empty] | 1.0436 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093942 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.7928 | [19 empty] | -1.2147 | [5 empty] | 1.0436 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093942 | -1.4514 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.7928 | [19 empty] | -1.2148 | [5 empty] | 1.0436 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093942 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.7928 | [19 empty] | -1.2148 | [5 empty] | 1.0436 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093942 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2147 | [5 empty] | 1.0434 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4514 | -1.4321 | -1.4321 | -1.4321 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:24 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2148 | [5 empty] | 1.0434 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2148 | [5 empty] | 1.0434 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2148 | [5 empty] | 1.0434 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2148 | [5 empty] | 1.0434 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2149 | [5 empty] | 1.0434 | -1.2148 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4516 | -1.4323 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.7239 | [19 empty] | -1.2148 | [5 empty] | 1.0434 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093929 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2148 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2148 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2148 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2148 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4323 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:25 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2148 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2148 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.6549 | [19 empty] | -1.2149 | [5 empty] | 1.0433 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093916 | -1.4515 | -1.4323 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.5859 | [19 empty] | -1.2148 | [5 empty] | 1.0432 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093904 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.5859 | [19 empty] | -1.2148 | [5 empty] | 1.0432 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093904 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.5859 | [19 empty] | -1.2148 | [5 empty] | 1.0432 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093904 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.5859 | [19 empty] | -1.2148 | [5 empty] | 1.0432 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093904 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.5859 | [19 empty] | -1.2148 | [5 empty] | 1.0432 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093904 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.5859 | [19 empty] | -1.2148 | [5 empty] | 1.0432 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093904 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.517 | [19 empty] | -1.2148 | [5 empty] | 1.043 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093891 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.517 | [19 empty] | -1.2148 | [5 empty] | 1.043 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093891 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:26 | 20 | 20 |  | 521.517 | [19 empty] | -1.2148 | [5 empty] | 1.043 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093891 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.517 | [19 empty] | -1.2148 | [5 empty] | 1.043 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093891 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.517 | [19 empty] | -1.2148 | [5 empty] | 1.043 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093891 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.517 | [19 empty] | -1.2148 | [5 empty] | 1.043 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093891 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:27 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.448 | [19 empty] | -1.2148 | [5 empty] | 1.0429 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093878 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.379 | [19 empty] | -1.2148 | [5 empty] | 1.0428 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093866 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.379 | [19 empty] | -1.2148 | [5 empty] | 1.0428 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093866 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.379 | [19 empty] | -1.2148 | [5 empty] | 1.0428 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093866 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.379 | [19 empty] | -1.2148 | [5 empty] | 1.0428 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093866 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.379 | [19 empty] | -1.2148 | [5 empty] | 1.0428 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093866 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.31 | [19 empty] | -1.2148 | [5 empty] | 1.0426 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093853 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.379 | [19 empty] | -1.2148 | [5 empty] | 1.0428 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093866 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:28 | 20 | 20 |  | 521.31 | [19 empty] | -1.2148 | [5 empty] | 1.0426 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093853 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.31 | [19 empty] | -1.2148 | [5 empty] | 1.0426 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093853 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.31 | [19 empty] | -1.2148 | [5 empty] | 1.0426 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093853 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.31 | [19 empty] | -1.2148 | [5 empty] | 1.0426 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093853 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.31 | [19 empty] | -1.2148 | [5 empty] | 1.0426 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093853 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.2411 | [19 empty] | -1.2148 | [5 empty] | 1.0425 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009384 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.2411 | [19 empty] | -1.2148 | [5 empty] | 1.0425 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009384 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.2411 | [19 empty] | -1.2148 | [5 empty] | 1.0425 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009384 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.2411 | [19 empty] | -1.2148 | [5 empty] | 1.0425 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009384 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.2411 | [19 empty] | -1.2148 | [5 empty] | 1.0425 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009384 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:29 | 20 | 20 |  | 521.2411 | [19 empty] | -1.2148 | [5 empty] | 1.0425 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009384 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1721 | [19 empty] | -1.2148 | [5 empty] | 1.0423 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093828 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1721 | [19 empty] | -1.2148 | [5 empty] | 1.0423 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093828 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1721 | [19 empty] | -1.2148 | [5 empty] | 1.0423 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093828 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1721 | [19 empty] | -1.2148 | [5 empty] | 1.0423 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093828 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1721 | [19 empty] | -1.2148 | [5 empty] | 1.0423 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093828 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1721 | [19 empty] | -1.2148 | [5 empty] | 1.0423 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093828 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1031 | [19 empty] | -1.2148 | [5 empty] | 1.0422 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093815 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1031 | [19 empty] | -1.2148 | [5 empty] | 1.0422 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093815 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1031 | [19 empty] | -1.2148 | [5 empty] | 1.0422 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093815 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:30 | 20 | 20 |  | 521.1031 | [19 empty] | -1.2148 | [5 empty] | 1.0422 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093815 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.1031 | [19 empty] | -1.2148 | [5 empty] | 1.0422 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093815 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 521.0341 | [19 empty] | -1.2148 | [5 empty] | 1.0421 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093802 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 520.9652 | [19 empty] | -1.2148 | [5 empty] | 1.0419 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009379 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:31 | 20 | 20 |  | 520.9652 | [19 empty] | -1.2148 | [5 empty] | 1.0419 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.0009379 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.9652 | [19 empty] | -1.2148 | [5 empty] | 1.0419 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009379 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.9652 | [19 empty] | -1.2148 | [5 empty] | 1.0419 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009379 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.9652 | [19 empty] | -1.2148 | [5 empty] | 1.0419 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.0009379 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:32 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8962 | [19 empty] | -1.2148 | [5 empty] | 1.0418 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093777 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:33 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.8272 | [19 empty] | -1.2148 | [5 empty] | 1.0417 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093764 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2148 | [5 empty] | 1.0415 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2148 | [5 empty] | 1.0414 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2148 | [5 empty] | 1.0414 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:34 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2148 | [5 empty] | 1.0414 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2148 | [5 empty] | 1.0414 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2148 | [5 empty] | 1.0414 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:35 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2148 | [5 empty] | 1.0412 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2148 | [5 empty] | 1.0411 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:36 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2148 | [5 empty] | 1.041 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:37 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2148 | [5 empty] | 1.0408 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:38 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2147 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4515 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2148 | [5 empty] | 1.0407 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4514 | -1.4322 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2148 | [5 empty] | 1.0406 | -1.2146 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4514 | -1.4322 | -1.4322 | -1.4322 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 541.0368 | [19 empty] | -1.2172 | [5 empty] | 1.0821 | -1.2171 | 0.90444 | 17.6664 | 0.23773 | -0.00097474 | -1.4538 | -1.4346 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 703.6099 | [19 empty] | -1.2172 | [5 empty] | 1.4072 | -1.2171 | 0.90444 | 17.6664 | 0.23773 | -0.0012731 | -1.4536 | -1.4343 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 702.9891 | [19 empty] | -1.2172 | [5 empty] | 1.406 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001272 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 702.3684 | [19 empty] | -1.2171 | [5 empty] | 1.4047 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012709 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:39 | 20 | 20 |  | 702.0235 | [19 empty] | -1.2172 | [5 empty] | 1.404 | -1.2171 | 0.90444 | 17.6664 | 0.23773 | -0.0012702 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 701.6786 | [19 empty] | -1.2172 | [5 empty] | 1.4034 | -1.2171 | 0.90444 | 17.6664 | 0.23773 | -0.0012696 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 701.4717 | [19 empty] | -1.2171 | [5 empty] | 1.4029 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012692 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 701.2648 | [19 empty] | -1.2171 | [5 empty] | 1.4025 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012688 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 701.1268 | [19 empty] | -1.2171 | [5 empty] | 1.4023 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012686 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.9199 | [19 empty] | -1.2171 | [5 empty] | 1.4018 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012682 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.8509 | [19 empty] | -1.2171 | [5 empty] | 1.4017 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012681 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.713 | [19 empty] | -1.2171 | [5 empty] | 1.4014 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012678 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.575 | [19 empty] | -1.2171 | [5 empty] | 1.4012 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012676 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.4371 | [19 empty] | -1.2171 | [5 empty] | 1.4009 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012673 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.3681 | [19 empty] | -1.2171 | [5 empty] | 1.4007 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012672 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:40 | 20 | 20 |  | 700.2301 | [19 empty] | -1.2171 | [5 empty] | 1.4005 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012669 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 700.1612 | [19 empty] | -1.2171 | [5 empty] | 1.4003 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012668 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 700.0922 | [19 empty] | -1.2171 | [5 empty] | 1.4002 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012667 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 700.0232 | [19 empty] | -1.2171 | [5 empty] | 1.4 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012666 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.8853 | [19 empty] | -1.2171 | [5 empty] | 1.3998 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012663 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.8163 | [19 empty] | -1.2171 | [5 empty] | 1.3996 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012662 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.7473 | [19 empty] | -1.2171 | [5 empty] | 1.3995 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012661 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.6783 | [19 empty] | -1.217 | [5 empty] | 1.3994 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012659 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.6094 | [19 empty] | -1.2171 | [5 empty] | 1.3992 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012658 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.5404 | [19 empty] | -1.2171 | [5 empty] | 1.3991 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012657 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:41 | 20 | 20 |  | 699.4714 | [19 empty] | -1.2171 | [5 empty] | 1.3989 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012656 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.4024 | [19 empty] | -1.2171 | [5 empty] | 1.3988 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012654 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.3335 | [19 empty] | -1.2171 | [5 empty] | 1.3987 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012653 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.2645 | [19 empty] | -1.2171 | [5 empty] | 1.3985 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012652 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.1955 | [19 empty] | -1.217 | [5 empty] | 1.3984 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.001265 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.1265 | [19 empty] | -1.2171 | [5 empty] | 1.3983 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012649 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.1265 | [19 empty] | -1.2171 | [5 empty] | 1.3983 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012649 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 699.0576 | [19 empty] | -1.2171 | [5 empty] | 1.3981 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012648 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 698.9886 | [19 empty] | -1.2171 | [5 empty] | 1.398 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012647 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 698.9196 | [19 empty] | -1.2171 | [5 empty] | 1.3978 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012645 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 698.8506 | [19 empty] | -1.2171 | [5 empty] | 1.3977 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012644 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:42 | 20 | 20 |  | 698.8506 | [19 empty] | -1.2171 | [5 empty] | 1.3977 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012644 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.7817 | [19 empty] | -1.2171 | [5 empty] | 1.3976 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012643 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.7127 | [19 empty] | -1.2171 | [5 empty] | 1.3974 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012642 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.6437 | [19 empty] | -1.2171 | [5 empty] | 1.3973 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.001264 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.5747 | [19 empty] | -1.2171 | [5 empty] | 1.3971 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012639 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.5747 | [19 empty] | -1.217 | [5 empty] | 1.3971 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012639 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.5058 | [19 empty] | -1.2171 | [5 empty] | 1.397 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012638 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.5058 | [19 empty] | -1.2171 | [5 empty] | 1.397 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012638 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.4368 | [19 empty] | -1.2171 | [5 empty] | 1.3969 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012637 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.3678 | [19 empty] | -1.217 | [5 empty] | 1.3967 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012635 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:43 | 20 | 20 |  | 698.2989 | [19 empty] | -1.217 | [5 empty] | 1.3966 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012634 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.2989 | [19 empty] | -1.2171 | [5 empty] | 1.3966 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012634 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.2299 | [19 empty] | -1.2171 | [5 empty] | 1.3965 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012633 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.2299 | [19 empty] | -1.2171 | [5 empty] | 1.3965 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012633 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.1609 | [19 empty] | -1.2171 | [5 empty] | 1.3963 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012631 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.1609 | [19 empty] | -1.217 | [5 empty] | 1.3963 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012631 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.0919 | [19 empty] | -1.217 | [5 empty] | 1.3962 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.001263 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.0919 | [19 empty] | -1.217 | [5 empty] | 1.3962 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.001263 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 698.023 | [19 empty] | -1.217 | [5 empty] | 1.396 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012629 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 697.954 | [19 empty] | -1.2171 | [5 empty] | 1.3959 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012628 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:44 | 20 | 20 |  | 697.954 | [19 empty] | -1.2171 | [5 empty] | 1.3959 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012628 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.954 | [19 empty] | -1.2171 | [5 empty] | 1.3959 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012628 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.885 | [19 empty] | -1.2171 | [5 empty] | 1.3958 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012626 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.816 | [19 empty] | -1.2171 | [5 empty] | 1.3956 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012625 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.816 | [19 empty] | -1.2171 | [5 empty] | 1.3956 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012625 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.816 | [19 empty] | -1.2171 | [5 empty] | 1.3956 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012625 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.7471 | [19 empty] | -1.2171 | [5 empty] | 1.3955 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012624 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.6781 | [19 empty] | -1.217 | [5 empty] | 1.3954 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012623 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.6781 | [19 empty] | -1.2171 | [5 empty] | 1.3954 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012623 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.6091 | [19 empty] | -1.217 | [5 empty] | 1.3952 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012621 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.6091 | [19 empty] | -1.2171 | [5 empty] | 1.3952 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012621 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:45 | 20 | 20 |  | 697.5401 | [19 empty] | -1.2171 | [5 empty] | 1.3951 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001262 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.5401 | [19 empty] | -1.2171 | [5 empty] | 1.3951 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001262 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.5401 | [19 empty] | -1.2171 | [5 empty] | 1.3951 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001262 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.4712 | [19 empty] | -1.2171 | [5 empty] | 1.3949 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012619 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.4712 | [19 empty] | -1.2171 | [5 empty] | 1.3949 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012619 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.4022 | [19 empty] | -1.2171 | [5 empty] | 1.3948 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012618 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.4022 | [19 empty] | -1.2171 | [5 empty] | 1.3948 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012618 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.3332 | [19 empty] | -1.2171 | [5 empty] | 1.3947 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012616 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.3332 | [19 empty] | -1.2171 | [5 empty] | 1.3947 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012616 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.2642 | [19 empty] | -1.2171 | [5 empty] | 1.3945 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012615 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:46 | 20 | 20 |  | 697.2642 | [19 empty] | -1.2171 | [5 empty] | 1.3945 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012615 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.1953 | [19 empty] | -1.217 | [5 empty] | 1.3944 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2171 | [5 empty] | 1.3944 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2171 | [5 empty] | 1.3944 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2171 | [5 empty] | 1.3943 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2171 | [5 empty] | 1.3943 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2171 | [5 empty] | 1.3941 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 697.0573 | [19 empty] | -1.217 | [5 empty] | 1.3941 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2171 | [5 empty] | 1.394 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 696.9883 | [19 empty] | -1.217 | [5 empty] | 1.394 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:47 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2171 | [5 empty] | 1.3938 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.9194 | [19 empty] | -1.217 | [5 empty] | 1.3938 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.9194 | [19 empty] | -1.217 | [5 empty] | 1.3938 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2171 | [5 empty] | 1.3937 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2171 | [5 empty] | 1.3937 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.8504 | [19 empty] | -1.217 | [5 empty] | 1.3937 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2171 | [5 empty] | 1.3936 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2171 | [5 empty] | 1.3936 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2171 | [5 empty] | 1.3936 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2171 | [5 empty] | 1.3934 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2171 | [5 empty] | 1.3934 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:48 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2171 | [5 empty] | 1.3933 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2171 | [5 empty] | 1.3933 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2171 | [5 empty] | 1.3931 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2171 | [5 empty] | 1.3931 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2171 | [5 empty] | 1.3931 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2171 | [5 empty] | 1.3931 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2171 | [5 empty] | 1.3931 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2171 | [5 empty] | 1.393 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2171 | [5 empty] | 1.393 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2171 | [5 empty] | 1.3929 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:49 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2171 | [5 empty] | 1.3929 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2171 | [5 empty] | 1.3929 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.3676 | [19 empty] | -1.217 | [5 empty] | 1.3927 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.3676 | [19 empty] | -1.217 | [5 empty] | 1.3927 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2171 | [5 empty] | 1.3927 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2171 | [5 empty] | 1.3926 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2171 | [5 empty] | 1.3926 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2171 | [5 empty] | 1.3926 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2171 | [5 empty] | 1.3925 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2171 | [5 empty] | 1.3925 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:50 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2171 | [5 empty] | 1.3925 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2171 | [5 empty] | 1.3923 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2171 | [5 empty] | 1.3923 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2171 | [5 empty] | 1.3923 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2171 | [5 empty] | 1.3923 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2171 | [5 empty] | 1.3922 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2171 | [5 empty] | 1.3922 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2171 | [5 empty] | 1.3922 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2171 | [5 empty] | 1.392 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2171 | [5 empty] | 1.392 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2171 | [5 empty] | 1.392 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:51 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2171 | [5 empty] | 1.3919 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2171 | [5 empty] | 1.3919 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2171 | [5 empty] | 1.3919 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2171 | [5 empty] | 1.3919 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2171 | [5 empty] | 1.3918 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2171 | [5 empty] | 1.3918 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.8158 | [19 empty] | -1.2171 | [5 empty] | 1.3916 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012588 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.8158 | [19 empty] | -1.217 | [5 empty] | 1.3916 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012588 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.7468 | [19 empty] | -1.217 | [5 empty] | 1.3915 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2171 | [5 empty] | 1.3915 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:52 | 20 | 20 |  | 695.7468 | [19 empty] | -1.217 | [5 empty] | 1.3915 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.7468 | [19 empty] | -1.217 | [5 empty] | 1.3915 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2171 | [5 empty] | 1.3915 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2171 | [5 empty] | 1.3914 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2171 | [5 empty] | 1.3914 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.6778 | [19 empty] | -1.217 | [5 empty] | 1.3914 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.6088 | [19 empty] | -1.217 | [5 empty] | 1.3912 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2171 | [5 empty] | 1.3912 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2171 | [5 empty] | 1.3912 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2171 | [5 empty] | 1.3911 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:53 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2171 | [5 empty] | 1.3911 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2171 | [5 empty] | 1.3911 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2171 | [5 empty] | 1.3911 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2171 | [5 empty] | 1.3911 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2171 | [5 empty] | 1.3909 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2171 | [5 empty] | 1.3909 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2171 | [5 empty] | 1.3909 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2171 | [5 empty] | 1.3909 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4019 | [19 empty] | -1.217 | [5 empty] | 1.3908 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2171 | [5 empty] | 1.3909 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4019 | [19 empty] | -1.2171 | [5 empty] | 1.3908 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:54 | 20 | 20 |  | 695.4019 | [19 empty] | -1.2171 | [5 empty] | 1.3908 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.4019 | [19 empty] | -1.217 | [5 empty] | 1.3908 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2171 | [5 empty] | 1.3907 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2171 | [5 empty] | 1.3907 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2171 | [5 empty] | 1.3907 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2171 | [5 empty] | 1.3907 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2171 | [5 empty] | 1.3907 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.264 | [19 empty] | -1.2171 | [5 empty] | 1.3905 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.264 | [19 empty] | -1.2171 | [5 empty] | 1.3905 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.264 | [19 empty] | -1.2171 | [5 empty] | 1.3905 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:55 | 20 | 20 |  | 695.264 | [19 empty] | -1.2171 | [5 empty] | 1.3905 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.2171 | [5 empty] | 1.3904 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.2171 | [5 empty] | 1.3904 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.2171 | [5 empty] | 1.3904 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.2171 | [5 empty] | 1.3904 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.217 | [5 empty] | 1.3904 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.2171 | [5 empty] | 1.3904 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.195 | [19 empty] | -1.2171 | [5 empty] | 1.3904 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.126 | [19 empty] | -1.2171 | [5 empty] | 1.3903 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.126 | [19 empty] | -1.2171 | [5 empty] | 1.3903 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.126 | [19 empty] | -1.2171 | [5 empty] | 1.3903 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:56 | 20 | 20 |  | 695.126 | [19 empty] | -1.2171 | [5 empty] | 1.3903 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 695.057 | [19 empty] | -1.2171 | [5 empty] | 1.3901 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 694.9881 | [19 empty] | -1.217 | [5 empty] | 1.39 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 694.9881 | [19 empty] | -1.217 | [5 empty] | 1.39 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:57 | 20 | 20 |  | 694.9881 | [19 empty] | -1.2171 | [5 empty] | 1.39 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9881 | [19 empty] | -1.2171 | [5 empty] | 1.39 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9881 | [19 empty] | -1.2171 | [5 empty] | 1.39 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2171 | [5 empty] | 1.3898 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2171 | [5 empty] | 1.3898 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2171 | [5 empty] | 1.3898 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9191 | [19 empty] | -1.217 | [5 empty] | 1.3898 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2171 | [5 empty] | 1.3898 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:58 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2171 | [5 empty] | 1.3897 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2171 | [5 empty] | 1.3896 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2171 | [5 empty] | 1.3896 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2171 | [5 empty] | 1.3896 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2171 | [5 empty] | 1.3896 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2171 | [5 empty] | 1.3896 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 15:59 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2172 | [5 empty] | 1.3894 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2171 | [5 empty] | 1.3894 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:00 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2171 | [5 empty] | 1.3893 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2171 | [5 empty] | 1.3891 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4534 | -1.4341 | -1.4341 | -1.4341 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2171 | [5 empty] | 1.3891 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2171 | [5 empty] | 1.3891 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:01 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2171 | [5 empty] | 1.3891 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2171 | [5 empty] | 1.3891 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2171 | [5 empty] | 1.3891 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2171 | [5 empty] | 1.389 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2171 | [5 empty] | 1.389 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2171 | [5 empty] | 1.389 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2171 | [5 empty] | 1.389 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2171 | [5 empty] | 1.389 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2171 | [5 empty] | 1.389 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4534 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:02 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4534 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.2169 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4534 | -1.4341 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2172 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2172 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2171 | [5 empty] | 1.3889 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2171 | [5 empty] | 1.3887 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2171 | [5 empty] | 1.3887 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2172 | [5 empty] | 1.3887 | -1.2171 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4535 | -1.4343 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:03 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2172 | [5 empty] | 1.3887 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2172 | [5 empty] | 1.3887 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2171 | [5 empty] | 1.3886 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2172 | [5 empty] | 1.3886 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2172 | [5 empty] | 1.3886 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2171 | [5 empty] | 1.3886 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2172 | [5 empty] | 1.3886 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2172 | [5 empty] | 1.3886 | -1.2171 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4536 | -1.4343 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2172 | [5 empty] | 1.3885 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4535 | -1.4342 | -1.4342 | -1.4342 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2171 | [5 empty] | 1.3885 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:04 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2171 | [5 empty] | 1.3885 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2171 | [5 empty] | 1.3885 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2171 | [5 empty] | 1.3885 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2171 | [5 empty] | 1.3885 | -1.217 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4535 | -1.4342 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 690.5047 | [19 empty] | -1.2183 | [5 empty] | 1.381 | -1.2182 | 0.90444 | 17.6664 | 0.23773 | -0.0012491 | -1.4547 | -1.4354 | -1.4354 | -1.4354 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 809.5549 | [19 empty] | -1.2205 | [5 empty] | 1.6191 | -1.2204 | 0.90444 | 17.6664 | 0.23773 | -0.0014676 | -1.4567 | -1.4374 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 1043.3788 | [19 empty] | -1.2208 | [5 empty] | 2.0868 | -1.2207 | 0.90444 | 17.6664 | 0.23773 | -0.0018968 | -1.4565 | -1.4373 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 1054.2768 | [19 empty] | -1.2208 | [5 empty] | 2.1086 | -1.2207 | 0.90444 | 17.6664 | 0.23773 | -0.0019168 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 1053.3111 | [19 empty] | -1.2208 | [5 empty] | 2.1066 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001915 | -1.4565 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 1052.6903 | [19 empty] | -1.2208 | [5 empty] | 2.1054 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019139 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 1052.2075 | [19 empty] | -1.2207 | [5 empty] | 2.1044 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001913 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:05 | 20 | 20 |  | 1051.7937 | [19 empty] | -1.2207 | [5 empty] | 2.1036 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019123 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1051.4488 | [19 empty] | -1.2207 | [5 empty] | 2.1029 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019116 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1051.1729 | [19 empty] | -1.2207 | [5 empty] | 2.1023 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019111 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1050.966 | [19 empty] | -1.2207 | [5 empty] | 2.1019 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019107 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1050.6901 | [19 empty] | -1.2207 | [5 empty] | 2.1014 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019102 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1050.4832 | [19 empty] | -1.2207 | [5 empty] | 2.101 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019098 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1050.2762 | [19 empty] | -1.2207 | [5 empty] | 2.1006 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019095 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1050.0693 | [19 empty] | -1.2206 | [5 empty] | 2.1001 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019091 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1049.8624 | [19 empty] | -1.2207 | [5 empty] | 2.0997 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019087 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1049.7244 | [19 empty] | -1.2206 | [5 empty] | 2.0994 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019085 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:06 | 20 | 20 |  | 1049.5865 | [19 empty] | -1.2207 | [5 empty] | 2.0992 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019082 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1049.4485 | [19 empty] | -1.2207 | [5 empty] | 2.0989 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019079 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1049.3106 | [19 empty] | -1.2206 | [5 empty] | 2.0986 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019077 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1049.1037 | [19 empty] | -1.2206 | [5 empty] | 2.0982 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019073 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.9657 | [19 empty] | -1.2206 | [5 empty] | 2.0979 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019071 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.8967 | [19 empty] | -1.2206 | [5 empty] | 2.0978 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019069 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.7588 | [19 empty] | -1.2206 | [5 empty] | 2.0975 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019067 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.6208 | [19 empty] | -1.2207 | [5 empty] | 2.0972 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019064 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.5519 | [19 empty] | -1.2207 | [5 empty] | 2.0971 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019063 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.4139 | [19 empty] | -1.2207 | [5 empty] | 2.0968 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019061 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:07 | 20 | 20 |  | 1048.276 | [19 empty] | -1.2206 | [5 empty] | 2.0966 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019058 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1048.207 | [19 empty] | -1.2207 | [5 empty] | 2.0964 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019057 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1048.069 | [19 empty] | -1.2206 | [5 empty] | 2.0961 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019054 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1048.0001 | [19 empty] | -1.2206 | [5 empty] | 2.096 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019053 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1047.8621 | [19 empty] | -1.2207 | [5 empty] | 2.0957 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001905 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1047.7931 | [19 empty] | -1.2207 | [5 empty] | 2.0956 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019049 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1047.7242 | [19 empty] | -1.2207 | [5 empty] | 2.0954 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019048 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1047.2414 | [19 empty] | -1.2207 | [5 empty] | 2.0945 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019039 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1049.3106 | [19 empty] | -1.2207 | [5 empty] | 2.0986 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019077 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1049.1726 | [19 empty] | -1.2207 | [5 empty] | 2.0983 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019074 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1049.1037 | [19 empty] | -1.2207 | [5 empty] | 2.0982 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019073 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:08 | 20 | 20 |  | 1049.0347 | [19 empty] | -1.2207 | [5 empty] | 2.0981 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019072 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.9657 | [19 empty] | -1.2207 | [5 empty] | 2.0979 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019071 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.8278 | [19 empty] | -1.2207 | [5 empty] | 2.0977 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019068 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.7588 | [19 empty] | -1.2206 | [5 empty] | 2.0975 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019067 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.6898 | [19 empty] | -1.2207 | [5 empty] | 2.0974 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019066 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.5519 | [19 empty] | -1.2207 | [5 empty] | 2.0971 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019063 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.4829 | [19 empty] | -1.2206 | [5 empty] | 2.097 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019062 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.4139 | [19 empty] | -1.2207 | [5 empty] | 2.0968 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019061 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.3449 | [19 empty] | -1.2207 | [5 empty] | 2.0967 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019059 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.276 | [19 empty] | -1.2206 | [5 empty] | 2.0966 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019058 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:09 | 20 | 20 |  | 1048.207 | [19 empty] | -1.2207 | [5 empty] | 2.0964 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019057 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1048.069 | [19 empty] | -1.2207 | [5 empty] | 2.0961 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019054 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1048.0001 | [19 empty] | -1.2207 | [5 empty] | 2.096 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019053 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.9311 | [19 empty] | -1.2207 | [5 empty] | 2.0959 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019052 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.8621 | [19 empty] | -1.2207 | [5 empty] | 2.0957 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001905 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.7931 | [19 empty] | -1.2207 | [5 empty] | 2.0956 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019049 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.7242 | [19 empty] | -1.2207 | [5 empty] | 2.0954 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019048 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.6552 | [19 empty] | -1.2207 | [5 empty] | 2.0953 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019047 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.5862 | [19 empty] | -1.2207 | [5 empty] | 2.0952 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019045 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.5173 | [19 empty] | -1.2207 | [5 empty] | 2.095 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019044 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.4483 | [19 empty] | -1.2207 | [5 empty] | 2.0949 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019043 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:10 | 20 | 20 |  | 1047.3793 | [19 empty] | -1.2207 | [5 empty] | 2.0948 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019042 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.3793 | [19 empty] | -1.2207 | [5 empty] | 2.0948 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019042 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.3103 | [19 empty] | -1.2207 | [5 empty] | 2.0946 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001904 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.2414 | [19 empty] | -1.2208 | [5 empty] | 2.0945 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019039 | -1.4565 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.1724 | [19 empty] | -1.2207 | [5 empty] | 2.0943 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019038 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.1034 | [19 empty] | -1.2207 | [5 empty] | 2.0942 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019036 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.1034 | [19 empty] | -1.2207 | [5 empty] | 2.0942 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019036 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1047.0344 | [19 empty] | -1.2207 | [5 empty] | 2.0941 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019035 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1046.9655 | [19 empty] | -1.2207 | [5 empty] | 2.0939 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019034 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1046.8965 | [19 empty] | -1.2207 | [5 empty] | 2.0938 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019033 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:11 | 20 | 20 |  | 1046.8275 | [19 empty] | -1.2207 | [5 empty] | 2.0937 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019031 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.8275 | [19 empty] | -1.2207 | [5 empty] | 2.0937 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019031 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.7585 | [19 empty] | -1.2207 | [5 empty] | 2.0935 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001903 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.6896 | [19 empty] | -1.2207 | [5 empty] | 2.0934 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019029 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.6206 | [19 empty] | -1.2207 | [5 empty] | 2.0932 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019028 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.6206 | [19 empty] | -1.2207 | [5 empty] | 2.0932 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019028 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.5516 | [19 empty] | -1.2207 | [5 empty] | 2.0931 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019026 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.4826 | [19 empty] | -1.2207 | [5 empty] | 2.093 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019025 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.4137 | [19 empty] | -1.2207 | [5 empty] | 2.0928 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019024 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.3447 | [19 empty] | -1.2207 | [5 empty] | 2.0927 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019023 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:12 | 20 | 20 |  | 1046.3447 | [19 empty] | -1.2206 | [5 empty] | 2.0927 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019023 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1046.2757 | [19 empty] | -1.2207 | [5 empty] | 2.0926 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019021 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1046.2067 | [19 empty] | -1.2207 | [5 empty] | 2.0924 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001902 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1046.1378 | [19 empty] | -1.2206 | [5 empty] | 2.0923 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019019 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1046.1378 | [19 empty] | -1.2207 | [5 empty] | 2.0923 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019019 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1046.0688 | [19 empty] | -1.2207 | [5 empty] | 2.0921 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019017 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1045.9998 | [19 empty] | -1.2207 | [5 empty] | 2.092 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019016 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1045.9308 | [19 empty] | -1.2207 | [5 empty] | 2.0919 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019015 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1045.9308 | [19 empty] | -1.2207 | [5 empty] | 2.0919 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019015 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1045.8619 | [19 empty] | -1.2207 | [5 empty] | 2.0917 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019014 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:13 | 20 | 20 |  | 1045.7929 | [19 empty] | -1.2207 | [5 empty] | 2.0916 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019012 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.7929 | [19 empty] | -1.2207 | [5 empty] | 2.0916 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019012 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.7239 | [19 empty] | -1.2207 | [5 empty] | 2.0914 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019011 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.7239 | [19 empty] | -1.2207 | [5 empty] | 2.0914 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019011 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.6549 | [19 empty] | -1.2207 | [5 empty] | 2.0913 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001901 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.586 | [19 empty] | -1.2207 | [5 empty] | 2.0912 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019009 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.586 | [19 empty] | -1.2207 | [5 empty] | 2.0912 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019009 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.586 | [19 empty] | -1.2207 | [5 empty] | 2.0912 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019009 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.517 | [19 empty] | -1.2207 | [5 empty] | 2.091 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019007 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.448 | [19 empty] | -1.2207 | [5 empty] | 2.0909 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019006 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.448 | [19 empty] | -1.2207 | [5 empty] | 2.0909 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019006 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:14 | 20 | 20 |  | 1045.379 | [19 empty] | -1.2207 | [5 empty] | 2.0908 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019005 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.379 | [19 empty] | -1.2207 | [5 empty] | 2.0908 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019005 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.3101 | [19 empty] | -1.2207 | [5 empty] | 2.0906 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019004 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.2411 | [19 empty] | -1.2207 | [5 empty] | 2.0905 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019002 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.2411 | [19 empty] | -1.2207 | [5 empty] | 2.0905 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019002 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.1721 | [19 empty] | -1.2207 | [5 empty] | 2.0903 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0019001 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.1031 | [19 empty] | -1.2207 | [5 empty] | 2.0902 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.1031 | [19 empty] | -1.2208 | [5 empty] | 2.0902 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0019 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.0342 | [19 empty] | -1.2207 | [5 empty] | 2.0901 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018998 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1045.0342 | [19 empty] | -1.2207 | [5 empty] | 2.0901 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018998 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:15 | 20 | 20 |  | 1044.9652 | [19 empty] | -1.2207 | [5 empty] | 2.0899 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018997 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.8962 | [19 empty] | -1.2207 | [5 empty] | 2.0898 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018996 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.8962 | [19 empty] | -1.2207 | [5 empty] | 2.0898 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018996 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.8962 | [19 empty] | -1.2207 | [5 empty] | 2.0898 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018996 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.8272 | [19 empty] | -1.2207 | [5 empty] | 2.0897 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018995 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.7583 | [19 empty] | -1.2207 | [5 empty] | 2.0895 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018993 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.7583 | [19 empty] | -1.2207 | [5 empty] | 2.0895 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018993 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.6893 | [19 empty] | -1.2207 | [5 empty] | 2.0894 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018992 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.6893 | [19 empty] | -1.2207 | [5 empty] | 2.0894 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018992 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.6203 | [19 empty] | -1.2207 | [5 empty] | 2.0892 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018991 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.6203 | [19 empty] | -1.2207 | [5 empty] | 2.0892 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018991 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:16 | 20 | 20 |  | 1044.5513 | [19 empty] | -1.2207 | [5 empty] | 2.0891 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001899 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.5513 | [19 empty] | -1.2207 | [5 empty] | 2.0891 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001899 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.5513 | [19 empty] | -1.2207 | [5 empty] | 2.0891 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001899 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.4824 | [19 empty] | -1.2206 | [5 empty] | 2.089 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018988 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.4824 | [19 empty] | -1.2207 | [5 empty] | 2.089 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018988 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.4824 | [19 empty] | -1.2207 | [5 empty] | 2.089 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018988 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.4134 | [19 empty] | -1.2207 | [5 empty] | 2.0888 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018987 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.4134 | [19 empty] | -1.2207 | [5 empty] | 2.0888 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018987 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.3444 | [19 empty] | -1.2207 | [5 empty] | 2.0887 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018986 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.2754 | [19 empty] | -1.2207 | [5 empty] | 2.0886 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018985 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:17 | 20 | 20 |  | 1044.2754 | [19 empty] | -1.2207 | [5 empty] | 2.0886 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018985 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1044.2065 | [19 empty] | -1.2207 | [5 empty] | 2.0884 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018983 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1044.1375 | [19 empty] | -1.2207 | [5 empty] | 2.0883 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018982 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1044.1375 | [19 empty] | -1.2207 | [5 empty] | 2.0883 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018982 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1044.1375 | [19 empty] | -1.2207 | [5 empty] | 2.0883 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018982 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1044.0685 | [19 empty] | -1.2207 | [5 empty] | 2.0881 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018981 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1044.0685 | [19 empty] | -1.2207 | [5 empty] | 2.0881 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018981 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1043.9995 | [19 empty] | -1.2207 | [5 empty] | 2.088 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018979 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1043.9995 | [19 empty] | -1.2207 | [5 empty] | 2.088 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018979 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1043.9306 | [19 empty] | -1.2207 | [5 empty] | 2.0879 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018978 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:18 | 20 | 20 |  | 1043.9306 | [19 empty] | -1.2207 | [5 empty] | 2.0879 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018978 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.8616 | [19 empty] | -1.2207 | [5 empty] | 2.0877 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018977 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.8616 | [19 empty] | -1.2207 | [5 empty] | 2.0877 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018977 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.7926 | [19 empty] | -1.2208 | [5 empty] | 2.0876 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018976 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.7926 | [19 empty] | -1.2207 | [5 empty] | 2.0876 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018976 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.7236 | [19 empty] | -1.2207 | [5 empty] | 2.0874 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018974 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.7236 | [19 empty] | -1.2207 | [5 empty] | 2.0874 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018974 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.7236 | [19 empty] | -1.2207 | [5 empty] | 2.0874 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018974 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.6547 | [19 empty] | -1.2207 | [5 empty] | 2.0873 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018973 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.6547 | [19 empty] | -1.2207 | [5 empty] | 2.0873 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018973 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.6547 | [19 empty] | -1.2207 | [5 empty] | 2.0873 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018973 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:19 | 20 | 20 |  | 1043.5857 | [19 empty] | -1.2207 | [5 empty] | 2.0872 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018972 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.5857 | [19 empty] | -1.2207 | [5 empty] | 2.0872 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018972 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.5857 | [19 empty] | -1.2207 | [5 empty] | 2.0872 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018972 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.5857 | [19 empty] | -1.2207 | [5 empty] | 2.0872 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018972 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.5167 | [19 empty] | -1.2207 | [5 empty] | 2.087 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018971 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.5167 | [19 empty] | -1.2207 | [5 empty] | 2.087 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018971 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.5167 | [19 empty] | -1.2207 | [5 empty] | 2.087 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018971 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.4477 | [19 empty] | -1.2207 | [5 empty] | 2.0869 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018969 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.4477 | [19 empty] | -1.2207 | [5 empty] | 2.0869 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018969 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.4477 | [19 empty] | -1.2207 | [5 empty] | 2.0869 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018969 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:20 | 20 | 20 |  | 1043.3788 | [19 empty] | -1.2207 | [5 empty] | 2.0868 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018968 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.3788 | [19 empty] | -1.2207 | [5 empty] | 2.0868 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018968 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.3788 | [19 empty] | -1.2208 | [5 empty] | 2.0868 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018968 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.3098 | [19 empty] | -1.2207 | [5 empty] | 2.0866 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018967 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.3098 | [19 empty] | -1.2207 | [5 empty] | 2.0866 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018967 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.3098 | [19 empty] | -1.2207 | [5 empty] | 2.0866 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018967 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.2408 | [19 empty] | -1.2207 | [5 empty] | 2.0865 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018966 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.2408 | [19 empty] | -1.2207 | [5 empty] | 2.0865 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018966 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.2408 | [19 empty] | -1.2207 | [5 empty] | 2.0865 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018966 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.2408 | [19 empty] | -1.2207 | [5 empty] | 2.0865 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018966 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:21 | 20 | 20 |  | 1043.2408 | [19 empty] | -1.2208 | [5 empty] | 2.0865 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018966 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1719 | [19 empty] | -1.2208 | [5 empty] | 2.0863 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018964 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1719 | [19 empty] | -1.2208 | [5 empty] | 2.0863 | -1.2207 | 0.90444 | 17.6664 | 0.23773 | -0.0018964 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1719 | [19 empty] | -1.2207 | [5 empty] | 2.0863 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018964 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1719 | [19 empty] | -1.2207 | [5 empty] | 2.0863 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018964 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:22 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2208 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1043.1029 | [19 empty] | -1.2207 | [5 empty] | 2.0862 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018963 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1043.0339 | [19 empty] | -1.2207 | [5 empty] | 2.0861 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018962 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1043.0339 | [19 empty] | -1.2207 | [5 empty] | 2.0861 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018962 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1043.0339 | [19 empty] | -1.2207 | [5 empty] | 2.0861 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018962 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1043.0339 | [19 empty] | -1.2207 | [5 empty] | 2.0861 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018962 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1043.0339 | [19 empty] | -1.2207 | [5 empty] | 2.0861 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018962 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1042.9649 | [19 empty] | -1.2207 | [5 empty] | 2.0859 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001896 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1042.9649 | [19 empty] | -1.2207 | [5 empty] | 2.0859 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001896 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1042.9649 | [19 empty] | -1.2207 | [5 empty] | 2.0859 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001896 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:23 | 20 | 20 |  | 1042.896 | [19 empty] | -1.2208 | [5 empty] | 2.0858 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018959 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.896 | [19 empty] | -1.2207 | [5 empty] | 2.0858 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018959 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.896 | [19 empty] | -1.2207 | [5 empty] | 2.0858 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018959 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.896 | [19 empty] | -1.2207 | [5 empty] | 2.0858 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018959 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.896 | [19 empty] | -1.2207 | [5 empty] | 2.0858 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018959 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.896 | [19 empty] | -1.2208 | [5 empty] | 2.0858 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018959 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.827 | [19 empty] | -1.2207 | [5 empty] | 2.0857 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018958 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.827 | [19 empty] | -1.2207 | [5 empty] | 2.0857 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018958 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.827 | [19 empty] | -1.2208 | [5 empty] | 2.0857 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018958 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.827 | [19 empty] | -1.2207 | [5 empty] | 2.0857 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018958 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.827 | [19 empty] | -1.2207 | [5 empty] | 2.0857 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018958 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:24 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2208 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2207 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2207 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2207 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2208 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2208 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2208 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.758 | [19 empty] | -1.2208 | [5 empty] | 2.0855 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018957 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.689 | [19 empty] | -1.2208 | [5 empty] | 2.0854 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018955 | -1.4564 | -1.4372 | -1.4372 | -1.4372 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.689 | [19 empty] | -1.2208 | [5 empty] | 2.0854 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018955 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:25 | 20 | 20 |  | 1042.689 | [19 empty] | -1.2208 | [5 empty] | 2.0854 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018955 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.689 | [19 empty] | -1.2208 | [5 empty] | 2.0854 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018955 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.689 | [19 empty] | -1.2207 | [5 empty] | 2.0854 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018955 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.6201 | [19 empty] | -1.2207 | [5 empty] | 2.0852 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018954 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.6201 | [19 empty] | -1.2207 | [5 empty] | 2.0852 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018954 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.6201 | [19 empty] | -1.2207 | [5 empty] | 2.0852 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018954 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.6201 | [19 empty] | -1.2208 | [5 empty] | 2.0852 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018954 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.6201 | [19 empty] | -1.2208 | [5 empty] | 2.0852 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018954 | -1.4565 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.6201 | [19 empty] | -1.2207 | [5 empty] | 2.0852 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018954 | -1.4564 | -1.4372 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:26 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.5511 | [19 empty] | -1.2207 | [5 empty] | 2.0851 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018953 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2207 | [5 empty] | 2.085 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2207 | [5 empty] | 2.085 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2207 | [5 empty] | 2.085 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:27 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2207 | [5 empty] | 2.085 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2207 | [5 empty] | 2.085 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2207 | [5 empty] | 2.085 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2207 | [5 empty] | 2.0848 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4821 | [19 empty] | -1.2206 | [5 empty] | 2.085 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018952 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2207 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2207 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2206 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2206 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2206 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2207 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:28 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2206 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2207 | [5 empty] | 2.0848 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.4131 | [19 empty] | -1.2207 | [5 empty] | 2.0848 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.001895 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2206 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:29 | 20 | 20 |  | 1042.3442 | [19 empty] | -1.2207 | [5 empty] | 2.0847 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018949 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2752 | [19 empty] | -1.2207 | [5 empty] | 2.0846 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018948 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2752 | [19 empty] | -1.2206 | [5 empty] | 2.0846 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018948 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2752 | [19 empty] | -1.2207 | [5 empty] | 2.0846 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018948 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2752 | [19 empty] | -1.2207 | [5 empty] | 2.0846 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018948 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2062 | [19 empty] | -1.2207 | [5 empty] | 2.0844 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018947 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2752 | [19 empty] | -1.2207 | [5 empty] | 2.0846 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018948 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2062 | [19 empty] | -1.2207 | [5 empty] | 2.0844 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018947 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2062 | [19 empty] | -1.2207 | [5 empty] | 2.0844 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018947 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.2062 | [19 empty] | -1.2207 | [5 empty] | 2.0844 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018947 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:30 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.1372 | [19 empty] | -1.2207 | [5 empty] | 2.0843 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018945 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4563 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:31 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:32 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2207 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1042.0683 | [19 empty] | -1.2206 | [5 empty] | 2.0841 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018944 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9993 | [19 empty] | -1.2207 | [5 empty] | 2.084 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018943 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9993 | [19 empty] | -1.2207 | [5 empty] | 2.084 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018943 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9993 | [19 empty] | -1.2207 | [5 empty] | 2.084 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018943 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9993 | [19 empty] | -1.2207 | [5 empty] | 2.084 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018943 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:33 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.9303 | [19 empty] | -1.2207 | [5 empty] | 2.0839 | -1.2206 | 0.90444 | 17.6664 | 0.23773 | -0.0018941 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2207 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4564 | -1.4371 | -1.4371 | -1.4371 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:34 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2207 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2207 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2207 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2206 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2206 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4563 | -1.437 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.8613 | [19 empty] | -1.2207 | [5 empty] | 2.0837 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.001894 | -1.4564 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.7924 | [19 empty] | -1.2206 | [5 empty] | 2.0836 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018939 | -1.4563 | -1.437 | -1.437 | -1.437 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.7924 | [19 empty] | -1.2207 | [5 empty] | 2.0836 | -1.2205 | 0.90444 | 17.6664 | 0.23773 | -0.0018939 | -1.4563 | -1.4371 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1041.7924 | [19 empty] | -1.2212 | [5 empty] | 2.0836 | -1.221 | 0.90444 | 17.6664 | 0.23773 | -0.0018939 | -1.4568 | -1.4376 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1134.4942 | [19 empty] | -1.2227 | [5 empty] | 2.269 | -1.2225 | 0.90444 | 17.6664 | 0.23773 | -0.0020641 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:35 | 20 | 20 |  | 1377.2848 | [19 empty] | -1.2231 | [5 empty] | 2.7546 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025097 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1397.3564 | [19 empty] | -1.2231 | [5 empty] | 2.7947 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025465 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1397.0805 | [19 empty] | -1.2231 | [5 empty] | 2.7942 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002546 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1396.5287 | [19 empty] | -1.2231 | [5 empty] | 2.7931 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002545 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1396.1149 | [19 empty] | -1.2231 | [5 empty] | 2.7922 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025443 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1395.701 | [19 empty] | -1.2231 | [5 empty] | 2.7914 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025435 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1395.3562 | [19 empty] | -1.2231 | [5 empty] | 2.7907 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025429 | -1.4581 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1395.0803 | [19 empty] | -1.2231 | [5 empty] | 2.7902 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025424 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1394.8733 | [19 empty] | -1.2231 | [5 empty] | 2.7897 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002542 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1394.5974 | [19 empty] | -1.2231 | [5 empty] | 2.7892 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025415 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1394.3905 | [19 empty] | -1.2231 | [5 empty] | 2.7888 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025411 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:36 | 20 | 20 |  | 1394.1836 | [19 empty] | -1.2231 | [5 empty] | 2.7884 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025407 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1394.0456 | [19 empty] | -1.2231 | [5 empty] | 2.7881 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025405 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1393.8387 | [19 empty] | -1.2231 | [5 empty] | 2.7877 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025401 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1393.7008 | [19 empty] | -1.2231 | [5 empty] | 2.7874 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025398 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1393.5628 | [19 empty] | -1.2231 | [5 empty] | 2.7871 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025396 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1393.3559 | [19 empty] | -1.2231 | [5 empty] | 2.7867 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025392 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1393.2179 | [19 empty] | -1.2231 | [5 empty] | 2.7864 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002539 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1393.08 | [19 empty] | -1.2231 | [5 empty] | 2.7862 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025387 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1392.942 | [19 empty] | -1.2231 | [5 empty] | 2.7859 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025384 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1392.8041 | [19 empty] | -1.2231 | [5 empty] | 2.7856 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025382 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:37 | 20 | 20 |  | 1392.6661 | [19 empty] | -1.2231 | [5 empty] | 2.7853 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025379 | -1.4581 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1392.5282 | [19 empty] | -1.2231 | [5 empty] | 2.7851 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025377 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1392.4592 | [19 empty] | -1.2231 | [5 empty] | 2.7849 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025376 | -1.4581 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1392.3213 | [19 empty] | -1.2231 | [5 empty] | 2.7846 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025373 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1392.2523 | [19 empty] | -1.2231 | [5 empty] | 2.7845 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025372 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1392.1144 | [19 empty] | -1.2231 | [5 empty] | 2.7842 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025369 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1392.0454 | [19 empty] | -1.2231 | [5 empty] | 2.7841 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025368 | -1.4581 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1391.9074 | [19 empty] | -1.2231 | [5 empty] | 2.7838 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025365 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1391.8385 | [19 empty] | -1.2231 | [5 empty] | 2.7837 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025364 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1391.7005 | [19 empty] | -1.2231 | [5 empty] | 2.7834 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025362 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1391.6315 | [19 empty] | -1.2231 | [5 empty] | 2.7833 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002536 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:38 | 20 | 20 |  | 1391.5626 | [19 empty] | -1.2231 | [5 empty] | 2.7831 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025359 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1391.4246 | [19 empty] | -1.2231 | [5 empty] | 2.7828 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025357 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1391.3556 | [19 empty] | -1.2231 | [5 empty] | 2.7827 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025355 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1391.2867 | [19 empty] | -1.2232 | [5 empty] | 2.7826 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025354 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1391.1487 | [19 empty] | -1.2232 | [5 empty] | 2.7823 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025352 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1391.0797 | [19 empty] | -1.2231 | [5 empty] | 2.7822 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002535 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1391.0108 | [19 empty] | -1.2232 | [5 empty] | 2.782 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025349 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1390.8728 | [19 empty] | -1.2231 | [5 empty] | 2.7817 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025346 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1390.8038 | [19 empty] | -1.2231 | [5 empty] | 2.7816 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025345 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1390.7349 | [19 empty] | -1.2231 | [5 empty] | 2.7815 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025344 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:39 | 20 | 20 |  | 1390.6659 | [19 empty] | -1.2232 | [5 empty] | 2.7813 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025343 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.5969 | [19 empty] | -1.2231 | [5 empty] | 2.7812 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025341 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.5279 | [19 empty] | -1.2231 | [5 empty] | 2.7811 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002534 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.39 | [19 empty] | -1.2232 | [5 empty] | 2.7808 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025338 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.321 | [19 empty] | -1.2231 | [5 empty] | 2.7806 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025336 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.252 | [19 empty] | -1.2232 | [5 empty] | 2.7805 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025335 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.1831 | [19 empty] | -1.2231 | [5 empty] | 2.7804 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025334 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.1141 | [19 empty] | -1.2231 | [5 empty] | 2.7802 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025333 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.0451 | [19 empty] | -1.2231 | [5 empty] | 2.7801 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025331 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1390.0451 | [19 empty] | -1.2232 | [5 empty] | 2.7801 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025331 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:40 | 20 | 20 |  | 1389.9072 | [19 empty] | -1.2231 | [5 empty] | 2.7798 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025329 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.9072 | [19 empty] | -1.2232 | [5 empty] | 2.7798 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025329 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.8382 | [19 empty] | -1.2232 | [5 empty] | 2.7797 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025327 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.7692 | [19 empty] | -1.2232 | [5 empty] | 2.7795 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025326 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.7002 | [19 empty] | -1.2231 | [5 empty] | 2.7794 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025325 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.6313 | [19 empty] | -1.2231 | [5 empty] | 2.7793 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025324 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.5623 | [19 empty] | -1.2231 | [5 empty] | 2.7791 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025322 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.4933 | [19 empty] | -1.2232 | [5 empty] | 2.779 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025321 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.4243 | [19 empty] | -1.2231 | [5 empty] | 2.7788 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002532 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.3554 | [19 empty] | -1.2231 | [5 empty] | 2.7787 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025319 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.2864 | [19 empty] | -1.2232 | [5 empty] | 2.7786 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025317 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:41 | 20 | 20 |  | 1389.2174 | [19 empty] | -1.2232 | [5 empty] | 2.7784 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025316 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1389.1484 | [19 empty] | -1.2232 | [5 empty] | 2.7783 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025315 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1389.0795 | [19 empty] | -1.2232 | [5 empty] | 2.7782 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025314 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1389.0105 | [19 empty] | -1.2232 | [5 empty] | 2.778 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025312 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1389.0105 | [19 empty] | -1.2232 | [5 empty] | 2.778 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025312 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1388.9415 | [19 empty] | -1.2232 | [5 empty] | 2.7779 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025311 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1388.8725 | [19 empty] | -1.2232 | [5 empty] | 2.7777 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002531 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1388.8036 | [19 empty] | -1.2231 | [5 empty] | 2.7776 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025308 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1388.7346 | [19 empty] | -1.2232 | [5 empty] | 2.7775 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025307 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1388.6656 | [19 empty] | -1.2232 | [5 empty] | 2.7773 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025306 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:42 | 20 | 20 |  | 1388.6656 | [19 empty] | -1.2232 | [5 empty] | 2.7773 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025306 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.5966 | [19 empty] | -1.2232 | [5 empty] | 2.7772 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025305 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.5277 | [19 empty] | -1.2232 | [5 empty] | 2.7771 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025303 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.5277 | [19 empty] | -1.2232 | [5 empty] | 2.7771 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025303 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.4587 | [19 empty] | -1.2232 | [5 empty] | 2.7769 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025302 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.3897 | [19 empty] | -1.2232 | [5 empty] | 2.7768 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025301 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.3207 | [19 empty] | -1.2232 | [5 empty] | 2.7766 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.00253 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.3207 | [19 empty] | -1.2232 | [5 empty] | 2.7766 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.00253 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.2518 | [19 empty] | -1.2232 | [5 empty] | 2.7765 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025298 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.1828 | [19 empty] | -1.2232 | [5 empty] | 2.7764 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025297 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:43 | 20 | 20 |  | 1388.1138 | [19 empty] | -1.2231 | [5 empty] | 2.7762 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025296 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1388.0449 | [19 empty] | -1.2232 | [5 empty] | 2.7761 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025295 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.9759 | [19 empty] | -1.2232 | [5 empty] | 2.776 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025293 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.9759 | [19 empty] | -1.2232 | [5 empty] | 2.776 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025293 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.9069 | [19 empty] | -1.2232 | [5 empty] | 2.7758 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025292 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.8379 | [19 empty] | -1.2232 | [5 empty] | 2.7757 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025291 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.8379 | [19 empty] | -1.2232 | [5 empty] | 2.7757 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025291 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.769 | [19 empty] | -1.2232 | [5 empty] | 2.7755 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002529 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.7 | [19 empty] | -1.2232 | [5 empty] | 2.7754 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025288 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.7 | [19 empty] | -1.2232 | [5 empty] | 2.7754 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025288 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.631 | [19 empty] | -1.2232 | [5 empty] | 2.7753 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025287 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:44 | 20 | 20 |  | 1387.562 | [19 empty] | -1.2232 | [5 empty] | 2.7751 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025286 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.562 | [19 empty] | -1.2232 | [5 empty] | 2.7751 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025286 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.4931 | [19 empty] | -1.2232 | [5 empty] | 2.775 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025284 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.4931 | [19 empty] | -1.2232 | [5 empty] | 2.775 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025284 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.4241 | [19 empty] | -1.2232 | [5 empty] | 2.7748 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025283 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.4241 | [19 empty] | -1.2232 | [5 empty] | 2.7748 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025283 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.3551 | [19 empty] | -1.2232 | [5 empty] | 2.7747 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025282 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.2861 | [19 empty] | -1.2232 | [5 empty] | 2.7746 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025281 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.2861 | [19 empty] | -1.2232 | [5 empty] | 2.7746 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025281 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.2172 | [19 empty] | -1.2232 | [5 empty] | 2.7744 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025279 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:45 | 20 | 20 |  | 1387.1482 | [19 empty] | -1.2232 | [5 empty] | 2.7743 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025278 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1387.0792 | [19 empty] | -1.2232 | [5 empty] | 2.7742 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025277 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1387.0792 | [19 empty] | -1.2232 | [5 empty] | 2.7742 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025277 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1387.0102 | [19 empty] | -1.2233 | [5 empty] | 2.774 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025276 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1387.0102 | [19 empty] | -1.2232 | [5 empty] | 2.774 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025276 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1386.9413 | [19 empty] | -1.2232 | [5 empty] | 2.7739 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025274 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1386.8723 | [19 empty] | -1.2232 | [5 empty] | 2.7737 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025273 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1386.8723 | [19 empty] | -1.2232 | [5 empty] | 2.7737 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025273 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1386.8033 | [19 empty] | -1.2232 | [5 empty] | 2.7736 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025272 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1386.7343 | [19 empty] | -1.2232 | [5 empty] | 2.7735 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025271 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:46 | 20 | 20 |  | 1386.7343 | [19 empty] | -1.2232 | [5 empty] | 2.7735 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025271 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.6654 | [19 empty] | -1.2232 | [5 empty] | 2.7733 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025269 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.6654 | [19 empty] | -1.2232 | [5 empty] | 2.7733 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025269 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.5964 | [19 empty] | -1.2232 | [5 empty] | 2.7732 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025268 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.5964 | [19 empty] | -1.2232 | [5 empty] | 2.7732 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025268 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.5274 | [19 empty] | -1.2232 | [5 empty] | 2.7731 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025267 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.5274 | [19 empty] | -1.2232 | [5 empty] | 2.7731 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025267 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.4584 | [19 empty] | -1.2232 | [5 empty] | 2.7729 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025265 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.4584 | [19 empty] | -1.2232 | [5 empty] | 2.7729 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025265 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.4584 | [19 empty] | -1.2232 | [5 empty] | 2.7729 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025265 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.3895 | [19 empty] | -1.2232 | [5 empty] | 2.7728 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025264 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:47 | 20 | 20 |  | 1386.3205 | [19 empty] | -1.2232 | [5 empty] | 2.7726 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025263 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.3205 | [19 empty] | -1.2232 | [5 empty] | 2.7726 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025263 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.2515 | [19 empty] | -1.2232 | [5 empty] | 2.7725 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025262 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.2515 | [19 empty] | -1.2232 | [5 empty] | 2.7725 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025262 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.1825 | [19 empty] | -1.2232 | [5 empty] | 2.7724 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002526 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.1825 | [19 empty] | -1.2232 | [5 empty] | 2.7724 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002526 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.1136 | [19 empty] | -1.2232 | [5 empty] | 2.7722 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025259 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.1136 | [19 empty] | -1.2232 | [5 empty] | 2.7722 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025259 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.0446 | [19 empty] | -1.2232 | [5 empty] | 2.7721 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025258 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.0446 | [19 empty] | -1.2232 | [5 empty] | 2.7721 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025258 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:48 | 20 | 20 |  | 1386.0446 | [19 empty] | -1.2232 | [5 empty] | 2.7721 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025258 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.9756 | [19 empty] | -1.2232 | [5 empty] | 2.772 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025257 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.9756 | [19 empty] | -1.2232 | [5 empty] | 2.772 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025257 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.9066 | [19 empty] | -1.2232 | [5 empty] | 2.7718 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025255 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.9066 | [19 empty] | -1.2232 | [5 empty] | 2.7718 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025255 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.8377 | [19 empty] | -1.2232 | [5 empty] | 2.7717 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025254 | -1.4582 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.8377 | [19 empty] | -1.2232 | [5 empty] | 2.7717 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025254 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.8377 | [19 empty] | -1.2232 | [5 empty] | 2.7717 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025254 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.7687 | [19 empty] | -1.2232 | [5 empty] | 2.7715 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025253 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.7687 | [19 empty] | -1.2232 | [5 empty] | 2.7715 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025253 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:49 | 20 | 20 |  | 1385.7687 | [19 empty] | -1.2232 | [5 empty] | 2.7715 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025253 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.6997 | [19 empty] | -1.2232 | [5 empty] | 2.7714 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025252 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.6997 | [19 empty] | -1.2232 | [5 empty] | 2.7714 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025252 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.6997 | [19 empty] | -1.2232 | [5 empty] | 2.7714 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025252 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.6997 | [19 empty] | -1.2232 | [5 empty] | 2.7714 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025252 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.6307 | [19 empty] | -1.2232 | [5 empty] | 2.7713 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002525 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.5618 | [19 empty] | -1.2232 | [5 empty] | 2.7711 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025249 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.5618 | [19 empty] | -1.2232 | [5 empty] | 2.7711 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025249 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.5618 | [19 empty] | -1.2232 | [5 empty] | 2.7711 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025249 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.4928 | [19 empty] | -1.2232 | [5 empty] | 2.771 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025248 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.4928 | [19 empty] | -1.2232 | [5 empty] | 2.771 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025248 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:50 | 20 | 20 |  | 1385.4238 | [19 empty] | -1.2232 | [5 empty] | 2.7708 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025246 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.4238 | [19 empty] | -1.2232 | [5 empty] | 2.7708 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025246 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.3548 | [19 empty] | -1.2232 | [5 empty] | 2.7707 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025245 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.3548 | [19 empty] | -1.2232 | [5 empty] | 2.7707 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025245 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.3548 | [19 empty] | -1.2232 | [5 empty] | 2.7707 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025245 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.2859 | [19 empty] | -1.2233 | [5 empty] | 2.7706 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025244 | -1.4583 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.2859 | [19 empty] | -1.2233 | [5 empty] | 2.7706 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025244 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.2859 | [19 empty] | -1.2233 | [5 empty] | 2.7706 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025244 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.2169 | [19 empty] | -1.2233 | [5 empty] | 2.7704 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025243 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.2169 | [19 empty] | -1.2232 | [5 empty] | 2.7704 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025243 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:51 | 20 | 20 |  | 1385.1479 | [19 empty] | -1.2232 | [5 empty] | 2.7703 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025241 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.1479 | [19 empty] | -1.2233 | [5 empty] | 2.7703 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025241 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.1479 | [19 empty] | -1.2233 | [5 empty] | 2.7703 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025241 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.1479 | [19 empty] | -1.2233 | [5 empty] | 2.7703 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025241 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.0789 | [19 empty] | -1.2233 | [5 empty] | 2.7702 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002524 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.0789 | [19 empty] | -1.2233 | [5 empty] | 2.7702 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002524 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.0789 | [19 empty] | -1.2233 | [5 empty] | 2.7702 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002524 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.01 | [19 empty] | -1.2233 | [5 empty] | 2.77 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025239 | -1.4583 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.01 | [19 empty] | -1.2233 | [5 empty] | 2.77 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025239 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.01 | [19 empty] | -1.2233 | [5 empty] | 2.77 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025239 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.01 | [19 empty] | -1.2233 | [5 empty] | 2.77 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025239 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:52 | 20 | 20 |  | 1385.01 | [19 empty] | -1.2233 | [5 empty] | 2.77 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025239 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.941 | [19 empty] | -1.2232 | [5 empty] | 2.7699 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025238 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.941 | [19 empty] | -1.2233 | [5 empty] | 2.7699 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025238 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.872 | [19 empty] | -1.2233 | [5 empty] | 2.7697 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025236 | -1.4583 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.872 | [19 empty] | -1.2233 | [5 empty] | 2.7697 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025236 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.872 | [19 empty] | -1.2233 | [5 empty] | 2.7697 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025236 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.872 | [19 empty] | -1.2233 | [5 empty] | 2.7697 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025236 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.803 | [19 empty] | -1.2233 | [5 empty] | 2.7696 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025235 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.803 | [19 empty] | -1.2233 | [5 empty] | 2.7696 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025235 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.803 | [19 empty] | -1.2233 | [5 empty] | 2.7696 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025235 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:53 | 20 | 20 |  | 1384.7341 | [19 empty] | -1.2233 | [5 empty] | 2.7695 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025234 | -1.4583 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.7341 | [19 empty] | -1.2233 | [5 empty] | 2.7695 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025234 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.6651 | [19 empty] | -1.2233 | [5 empty] | 2.7693 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025233 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.6651 | [19 empty] | -1.2233 | [5 empty] | 2.7693 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025233 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.6651 | [19 empty] | -1.2233 | [5 empty] | 2.7693 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025233 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.6651 | [19 empty] | -1.2233 | [5 empty] | 2.7693 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025233 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.6651 | [19 empty] | -1.2233 | [5 empty] | 2.7693 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025233 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.5961 | [19 empty] | -1.2232 | [5 empty] | 2.7692 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025231 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.5961 | [19 empty] | -1.2233 | [5 empty] | 2.7692 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025231 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.5271 | [19 empty] | -1.2233 | [5 empty] | 2.7691 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002523 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:54 | 20 | 20 |  | 1384.5271 | [19 empty] | -1.2233 | [5 empty] | 2.7691 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002523 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.5271 | [19 empty] | -1.2233 | [5 empty] | 2.7691 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002523 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.5271 | [19 empty] | -1.2233 | [5 empty] | 2.7691 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002523 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.4582 | [19 empty] | -1.2233 | [5 empty] | 2.7689 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025229 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.4582 | [19 empty] | -1.2233 | [5 empty] | 2.7689 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025229 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.4582 | [19 empty] | -1.2233 | [5 empty] | 2.7689 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025229 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.4582 | [19 empty] | -1.2233 | [5 empty] | 2.7689 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025229 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.3892 | [19 empty] | -1.2233 | [5 empty] | 2.7688 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025227 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.3892 | [19 empty] | -1.2233 | [5 empty] | 2.7688 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025227 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.3892 | [19 empty] | -1.2233 | [5 empty] | 2.7688 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025227 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:55 | 20 | 20 |  | 1384.3202 | [19 empty] | -1.2233 | [5 empty] | 2.7686 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025226 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.3202 | [19 empty] | -1.2233 | [5 empty] | 2.7686 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025226 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.3202 | [19 empty] | -1.2233 | [5 empty] | 2.7686 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025226 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.3202 | [19 empty] | -1.2233 | [5 empty] | 2.7686 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025226 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.2512 | [19 empty] | -1.2233 | [5 empty] | 2.7685 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025225 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.2512 | [19 empty] | -1.2233 | [5 empty] | 2.7685 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025225 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.2512 | [19 empty] | -1.2233 | [5 empty] | 2.7685 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025225 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.1823 | [19 empty] | -1.2233 | [5 empty] | 2.7684 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025224 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.1823 | [19 empty] | -1.2233 | [5 empty] | 2.7684 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025224 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.1823 | [19 empty] | -1.2233 | [5 empty] | 2.7684 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025224 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.1823 | [19 empty] | -1.2233 | [5 empty] | 2.7684 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025224 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:56 | 20 | 20 |  | 1384.1133 | [19 empty] | -1.2234 | [5 empty] | 2.7682 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025222 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.1133 | [19 empty] | -1.2234 | [5 empty] | 2.7682 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025222 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.1133 | [19 empty] | -1.2233 | [5 empty] | 2.7682 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025222 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.1133 | [19 empty] | -1.2233 | [5 empty] | 2.7682 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025222 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.0443 | [19 empty] | -1.2233 | [5 empty] | 2.7681 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025221 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.0443 | [19 empty] | -1.2234 | [5 empty] | 2.7681 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025221 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.0443 | [19 empty] | -1.2233 | [5 empty] | 2.7681 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025221 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.0443 | [19 empty] | -1.2233 | [5 empty] | 2.7681 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025221 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.0443 | [19 empty] | -1.2233 | [5 empty] | 2.7681 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025221 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1384.0443 | [19 empty] | -1.2233 | [5 empty] | 2.7681 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025221 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:57 | 20 | 20 |  | 1383.9753 | [19 empty] | -1.2233 | [5 empty] | 2.768 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002522 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9753 | [19 empty] | -1.2233 | [5 empty] | 2.768 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002522 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9753 | [19 empty] | -1.2233 | [5 empty] | 2.768 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.002522 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9753 | [19 empty] | -1.2233 | [5 empty] | 2.768 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.002522 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9064 | [19 empty] | -1.2233 | [5 empty] | 2.7678 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025219 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9064 | [19 empty] | -1.2233 | [5 empty] | 2.7678 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025219 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9064 | [19 empty] | -1.2234 | [5 empty] | 2.7678 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025219 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9064 | [19 empty] | -1.2233 | [5 empty] | 2.7678 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025219 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9064 | [19 empty] | -1.2234 | [5 empty] | 2.7678 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025219 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.9064 | [19 empty] | -1.2233 | [5 empty] | 2.7678 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025219 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.8374 | [19 empty] | -1.2233 | [5 empty] | 2.7677 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025217 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:58 | 20 | 20 |  | 1383.8374 | [19 empty] | -1.2233 | [5 empty] | 2.7677 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025217 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.8374 | [19 empty] | -1.2234 | [5 empty] | 2.7677 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025217 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.8374 | [19 empty] | -1.2233 | [5 empty] | 2.7677 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025217 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.8374 | [19 empty] | -1.2233 | [5 empty] | 2.7677 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025217 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.8374 | [19 empty] | -1.2233 | [5 empty] | 2.7677 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025217 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.7684 | [19 empty] | -1.2233 | [5 empty] | 2.7675 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025216 | -1.4583 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.7684 | [19 empty] | -1.2233 | [5 empty] | 2.7675 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025216 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.7684 | [19 empty] | -1.2233 | [5 empty] | 2.7675 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025216 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.6995 | [19 empty] | -1.2233 | [5 empty] | 2.7674 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025215 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.6995 | [19 empty] | -1.2233 | [5 empty] | 2.7674 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025215 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 16:59 | 20 | 20 |  | 1383.6995 | [19 empty] | -1.2233 | [5 empty] | 2.7674 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025215 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6995 | [19 empty] | -1.2233 | [5 empty] | 2.7674 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025215 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6995 | [19 empty] | -1.2233 | [5 empty] | 2.7674 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025215 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:00 | 20 | 20 |  | 1383.6305 | [19 empty] | -1.2233 | [5 empty] | 2.7673 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025214 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.5615 | [19 empty] | -1.2233 | [5 empty] | 2.7671 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025212 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.5615 | [19 empty] | -1.2233 | [5 empty] | 2.7671 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025212 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.5615 | [19 empty] | -1.2233 | [5 empty] | 2.7671 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025212 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.5615 | [19 empty] | -1.2233 | [5 empty] | 2.7671 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025212 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.5615 | [19 empty] | -1.2233 | [5 empty] | 2.7671 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025212 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.4925 | [19 empty] | -1.2233 | [5 empty] | 2.767 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025211 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.4925 | [19 empty] | -1.2233 | [5 empty] | 2.767 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025211 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.4925 | [19 empty] | -1.2233 | [5 empty] | 2.767 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025211 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.4925 | [19 empty] | -1.2233 | [5 empty] | 2.767 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025211 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1383.4925 | [19 empty] | -1.2235 | [5 empty] | 2.767 | -1.2233 | 0.90444 | 17.6664 | 0.23773 | -0.0025211 | -1.4585 | -1.4392 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:01 | 20 | 20 |  | 1458.8818 | [19 empty] | -1.2245 | [5 empty] | 2.9178 | -1.2243 | 0.90444 | 17.6664 | 0.23773 | -0.0026595 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1709.5355 | [19 empty] | -1.2248 | [5 empty] | 3.4191 | -1.2246 | 0.90444 | 17.6664 | 0.23773 | -0.0031196 | -1.4592 | -1.4399 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1743.9538 | [19 empty] | -1.2248 | [5 empty] | 3.4879 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031827 | -1.4592 | -1.4399 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1743.1261 | [19 empty] | -1.2249 | [5 empty] | 3.4863 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031812 | -1.4592 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1742.5053 | [19 empty] | -1.2249 | [5 empty] | 3.485 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031801 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1742.0225 | [19 empty] | -1.2248 | [5 empty] | 3.484 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031792 | -1.4592 | -1.4399 | -1.4399 | -1.4399 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1741.6776 | [19 empty] | -1.2249 | [5 empty] | 3.4834 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031786 | -1.4592 | -1.4399 | -1.4399 | -1.4399 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1741.3328 | [19 empty] | -1.2248 | [5 empty] | 3.4827 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031779 | -1.4592 | -1.4399 | -1.4399 | -1.4399 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1741.0569 | [19 empty] | -1.2249 | [5 empty] | 3.4821 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031774 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1740.781 | [19 empty] | -1.2248 | [5 empty] | 3.4816 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031769 | -1.4592 | -1.4399 | -1.4399 | -1.4399 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:02 | 20 | 20 |  | 1740.574 | [19 empty] | -1.2249 | [5 empty] | 3.4811 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031765 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1740.3671 | [19 empty] | -1.2248 | [5 empty] | 3.4807 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031762 | -1.4592 | -1.4399 | -1.4399 | -1.4399 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1740.1602 | [19 empty] | -1.2249 | [5 empty] | 3.4803 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031758 | -1.4592 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1739.9533 | [19 empty] | -1.2249 | [5 empty] | 3.4799 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031754 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1739.8153 | [19 empty] | -1.2248 | [5 empty] | 3.4796 | -1.2246 | 0.90444 | 17.6664 | 0.23773 | -0.0031751 | -1.4592 | -1.4399 | -1.4399 | -1.4399 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1739.6084 | [19 empty] | -1.2249 | [5 empty] | 3.4792 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031748 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1739.4704 | [19 empty] | -1.2249 | [5 empty] | 3.4789 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031745 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1739.2635 | [19 empty] | -1.2249 | [5 empty] | 3.4785 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031741 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1739.1256 | [19 empty] | -1.2249 | [5 empty] | 3.4783 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031739 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1738.9876 | [19 empty] | -1.2249 | [5 empty] | 3.478 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031736 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:03 | 20 | 20 |  | 1738.8497 | [19 empty] | -1.2249 | [5 empty] | 3.4777 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031734 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1738.7117 | [19 empty] | -1.2249 | [5 empty] | 3.4774 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031731 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1738.5738 | [19 empty] | -1.2249 | [5 empty] | 3.4771 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031729 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1738.4358 | [19 empty] | -1.2249 | [5 empty] | 3.4769 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031726 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1738.2979 | [19 empty] | -1.2249 | [5 empty] | 3.4766 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031724 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1738.2289 | [19 empty] | -1.2249 | [5 empty] | 3.4765 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031722 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1738.0909 | [19 empty] | -1.2249 | [5 empty] | 3.4762 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.003172 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1737.953 | [19 empty] | -1.2249 | [5 empty] | 3.4759 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031717 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1737.815 | [19 empty] | -1.2249 | [5 empty] | 3.4756 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031715 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1737.7461 | [19 empty] | -1.2249 | [5 empty] | 3.4755 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031713 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1737.6771 | [19 empty] | -1.2249 | [5 empty] | 3.4754 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031712 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:04 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2249 | [5 empty] | 3.4751 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.225 | [5 empty] | 3.4749 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2249 | [5 empty] | 3.4747 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2249 | [5 empty] | 3.4744 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1738.2289 | [19 empty] | -1.2249 | [5 empty] | 3.4765 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031722 | -1.4593 | -1.44 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1738.1599 | [19 empty] | -1.2249 | [5 empty] | 3.4763 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031721 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1738.0909 | [19 empty] | -1.225 | [5 empty] | 3.4762 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003172 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1738.022 | [19 empty] | -1.225 | [5 empty] | 3.476 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031719 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1737.884 | [19 empty] | -1.2249 | [5 empty] | 3.4758 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031716 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1737.815 | [19 empty] | -1.2249 | [5 empty] | 3.4756 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031715 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:05 | 20 | 20 |  | 1737.7461 | [19 empty] | -1.2249 | [5 empty] | 3.4755 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031713 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.6771 | [19 empty] | -1.2249 | [5 empty] | 3.4754 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031712 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.225 | [5 empty] | 3.4751 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2249 | [5 empty] | 3.4749 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2249 | [5 empty] | 3.4748 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.225 | [5 empty] | 3.4747 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2249 | [5 empty] | 3.4745 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.225 | [5 empty] | 3.4744 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2249 | [5 empty] | 3.4743 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.225 | [5 empty] | 3.4741 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:06 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.225 | [5 empty] | 3.4738 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.225 | [5 empty] | 3.4737 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.225 | [5 empty] | 3.4736 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.225 | [5 empty] | 3.4734 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.225 | [5 empty] | 3.4733 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2249 | [5 empty] | 3.4731 | -1.2247 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4593 | -1.44 | -1.44 | -1.44 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.225 | [5 empty] | 3.473 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.225 | [5 empty] | 3.4729 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.225 | [5 empty] | 3.4727 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.4593 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.225 | [5 empty] | 3.4726 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.225 | [5 empty] | 3.4725 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:07 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.225 | [5 empty] | 3.4723 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.225 | [5 empty] | 3.4722 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.225 | [5 empty] | 3.4722 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.225 | [5 empty] | 3.4719 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.8838 | [19 empty] | -1.225 | [5 empty] | 3.4718 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031679 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.8838 | [19 empty] | -1.225 | [5 empty] | 3.4718 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031679 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.8148 | [19 empty] | -1.225 | [5 empty] | 3.4716 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031678 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.225 | [5 empty] | 3.4715 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.225 | [5 empty] | 3.4714 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.6079 | [19 empty] | -1.225 | [5 empty] | 3.4712 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031674 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:08 | 20 | 20 |  | 1735.5389 | [19 empty] | -1.225 | [5 empty] | 3.4711 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031673 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.225 | [5 empty] | 3.4709 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.4009 | [19 empty] | -1.225 | [5 empty] | 3.4708 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003167 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.4009 | [19 empty] | -1.225 | [5 empty] | 3.4708 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003167 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.332 | [19 empty] | -1.225 | [5 empty] | 3.4707 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031669 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.263 | [19 empty] | -1.225 | [5 empty] | 3.4705 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031668 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2251 | [5 empty] | 3.4704 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.125 | [19 empty] | -1.225 | [5 empty] | 3.4703 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031665 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.0561 | [19 empty] | -1.225 | [5 empty] | 3.4701 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031664 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1735.0561 | [19 empty] | -1.2251 | [5 empty] | 3.4701 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031664 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:09 | 20 | 20 |  | 1734.9871 | [19 empty] | -1.225 | [5 empty] | 3.47 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031663 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.9181 | [19 empty] | -1.225 | [5 empty] | 3.4698 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031662 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.8491 | [19 empty] | -1.2251 | [5 empty] | 3.4697 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003166 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.8491 | [19 empty] | -1.2251 | [5 empty] | 3.4697 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003166 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.7802 | [19 empty] | -1.225 | [5 empty] | 3.4696 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031659 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.7112 | [19 empty] | -1.2251 | [5 empty] | 3.4694 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031658 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.6422 | [19 empty] | -1.2251 | [5 empty] | 3.4693 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031657 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.6422 | [19 empty] | -1.225 | [5 empty] | 3.4693 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031657 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.5732 | [19 empty] | -1.2251 | [5 empty] | 3.4691 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031655 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.5043 | [19 empty] | -1.2251 | [5 empty] | 3.469 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031654 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.5043 | [19 empty] | -1.225 | [5 empty] | 3.469 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031654 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:10 | 20 | 20 |  | 1734.4353 | [19 empty] | -1.225 | [5 empty] | 3.4689 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031653 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.3663 | [19 empty] | -1.225 | [5 empty] | 3.4687 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031651 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.2973 | [19 empty] | -1.2251 | [5 empty] | 3.4686 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003165 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.2973 | [19 empty] | -1.2251 | [5 empty] | 3.4686 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003165 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.2284 | [19 empty] | -1.2251 | [5 empty] | 3.4685 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031649 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.1594 | [19 empty] | -1.2251 | [5 empty] | 3.4683 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031648 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.1594 | [19 empty] | -1.225 | [5 empty] | 3.4683 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031648 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.0904 | [19 empty] | -1.225 | [5 empty] | 3.4682 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031646 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.0214 | [19 empty] | -1.2251 | [5 empty] | 3.468 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031645 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1734.0214 | [19 empty] | -1.225 | [5 empty] | 3.468 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031645 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:11 | 20 | 20 |  | 1733.9525 | [19 empty] | -1.225 | [5 empty] | 3.4679 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031644 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.8835 | [19 empty] | -1.225 | [5 empty] | 3.4678 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031643 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.8835 | [19 empty] | -1.225 | [5 empty] | 3.4678 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031643 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.8145 | [19 empty] | -1.225 | [5 empty] | 3.4676 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031641 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.7455 | [19 empty] | -1.225 | [5 empty] | 3.4675 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003164 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.7455 | [19 empty] | -1.225 | [5 empty] | 3.4675 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003164 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.6766 | [19 empty] | -1.225 | [5 empty] | 3.4674 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031639 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.6766 | [19 empty] | -1.225 | [5 empty] | 3.4674 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031639 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.6076 | [19 empty] | -1.225 | [5 empty] | 3.4672 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031638 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.6076 | [19 empty] | -1.225 | [5 empty] | 3.4672 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031638 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.5386 | [19 empty] | -1.225 | [5 empty] | 3.4671 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031636 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:12 | 20 | 20 |  | 1733.5386 | [19 empty] | -1.225 | [5 empty] | 3.4671 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031636 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.4696 | [19 empty] | -1.2251 | [5 empty] | 3.4669 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031635 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.4696 | [19 empty] | -1.2251 | [5 empty] | 3.4669 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031635 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.4007 | [19 empty] | -1.225 | [5 empty] | 3.4668 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031634 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.3317 | [19 empty] | -1.225 | [5 empty] | 3.4667 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031632 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.3317 | [19 empty] | -1.225 | [5 empty] | 3.4667 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031632 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.2627 | [19 empty] | -1.225 | [5 empty] | 3.4665 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031631 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.2627 | [19 empty] | -1.2251 | [5 empty] | 3.4665 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031631 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.1937 | [19 empty] | -1.2251 | [5 empty] | 3.4664 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003163 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.1937 | [19 empty] | -1.2251 | [5 empty] | 3.4664 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003163 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:13 | 20 | 20 |  | 1733.1248 | [19 empty] | -1.225 | [5 empty] | 3.4662 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031629 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1733.1248 | [19 empty] | -1.225 | [5 empty] | 3.4662 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031629 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1733.0558 | [19 empty] | -1.225 | [5 empty] | 3.4661 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031627 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.9868 | [19 empty] | -1.225 | [5 empty] | 3.466 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031626 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.9868 | [19 empty] | -1.2251 | [5 empty] | 3.466 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031626 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.9179 | [19 empty] | -1.2251 | [5 empty] | 3.4658 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031625 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.9179 | [19 empty] | -1.225 | [5 empty] | 3.4658 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031625 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.8489 | [19 empty] | -1.225 | [5 empty] | 3.4657 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031624 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.8489 | [19 empty] | -1.2251 | [5 empty] | 3.4657 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031624 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.8489 | [19 empty] | -1.225 | [5 empty] | 3.4657 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031624 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:14 | 20 | 20 |  | 1732.7799 | [19 empty] | -1.2251 | [5 empty] | 3.4656 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031622 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.7799 | [19 empty] | -1.2251 | [5 empty] | 3.4656 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031622 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.7109 | [19 empty] | -1.2251 | [5 empty] | 3.4654 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031621 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.642 | [19 empty] | -1.2251 | [5 empty] | 3.4653 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003162 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.642 | [19 empty] | -1.2251 | [5 empty] | 3.4653 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003162 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.642 | [19 empty] | -1.2251 | [5 empty] | 3.4653 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003162 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.573 | [19 empty] | -1.2251 | [5 empty] | 3.4651 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031619 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.573 | [19 empty] | -1.2251 | [5 empty] | 3.4651 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031619 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.504 | [19 empty] | -1.2251 | [5 empty] | 3.465 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.504 | [19 empty] | -1.2251 | [5 empty] | 3.465 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:15 | 20 | 20 |  | 1732.504 | [19 empty] | -1.2251 | [5 empty] | 3.465 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.504 | [19 empty] | -1.225 | [5 empty] | 3.465 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.435 | [19 empty] | -1.225 | [5 empty] | 3.4649 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031616 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.435 | [19 empty] | -1.225 | [5 empty] | 3.4649 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031616 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.435 | [19 empty] | -1.225 | [5 empty] | 3.4649 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031616 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.3661 | [19 empty] | -1.225 | [5 empty] | 3.4647 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031615 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.3661 | [19 empty] | -1.225 | [5 empty] | 3.4647 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031615 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.2971 | [19 empty] | -1.2251 | [5 empty] | 3.4646 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031613 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.2971 | [19 empty] | -1.225 | [5 empty] | 3.4646 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031613 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.2281 | [19 empty] | -1.2251 | [5 empty] | 3.4645 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031612 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.2281 | [19 empty] | -1.2251 | [5 empty] | 3.4645 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031612 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:16 | 20 | 20 |  | 1732.1591 | [19 empty] | -1.2251 | [5 empty] | 3.4643 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031611 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1732.1591 | [19 empty] | -1.2251 | [5 empty] | 3.4643 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031611 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1732.1591 | [19 empty] | -1.225 | [5 empty] | 3.4643 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.0031611 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1732.0902 | [19 empty] | -1.2251 | [5 empty] | 3.4642 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003161 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1732.0902 | [19 empty] | -1.225 | [5 empty] | 3.4642 | -1.2248 | 0.90444 | 17.6664 | 0.23773 | -0.003161 | -1.4594 | -1.4401 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1732.0212 | [19 empty] | -1.2251 | [5 empty] | 3.464 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031608 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1732.0212 | [19 empty] | -1.2251 | [5 empty] | 3.464 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031608 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1731.9522 | [19 empty] | -1.2251 | [5 empty] | 3.4639 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031607 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1731.9522 | [19 empty] | -1.2251 | [5 empty] | 3.4639 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031607 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1731.8832 | [19 empty] | -1.2251 | [5 empty] | 3.4638 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031606 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:17 | 20 | 20 |  | 1731.8832 | [19 empty] | -1.2251 | [5 empty] | 3.4638 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031606 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.8832 | [19 empty] | -1.2251 | [5 empty] | 3.4638 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031606 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.8143 | [19 empty] | -1.2251 | [5 empty] | 3.4636 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031605 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.8143 | [19 empty] | -1.2251 | [5 empty] | 3.4636 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031605 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.8143 | [19 empty] | -1.2251 | [5 empty] | 3.4636 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031605 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.7453 | [19 empty] | -1.2251 | [5 empty] | 3.4635 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031603 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.7453 | [19 empty] | -1.2251 | [5 empty] | 3.4635 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031603 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.7453 | [19 empty] | -1.225 | [5 empty] | 3.4635 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031603 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.6763 | [19 empty] | -1.2251 | [5 empty] | 3.4634 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031602 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.6763 | [19 empty] | -1.2251 | [5 empty] | 3.4634 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031602 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.6763 | [19 empty] | -1.2251 | [5 empty] | 3.4634 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031602 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:18 | 20 | 20 |  | 1731.6073 | [19 empty] | -1.2251 | [5 empty] | 3.4632 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031601 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.6073 | [19 empty] | -1.2251 | [5 empty] | 3.4632 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031601 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.5384 | [19 empty] | -1.2251 | [5 empty] | 3.4631 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.00316 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.5384 | [19 empty] | -1.2251 | [5 empty] | 3.4631 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.00316 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.5384 | [19 empty] | -1.2251 | [5 empty] | 3.4631 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.00316 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.4694 | [19 empty] | -1.225 | [5 empty] | 3.4629 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031598 | -1.4594 | -1.4401 | -1.4401 | -1.4401 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.4694 | [19 empty] | -1.2251 | [5 empty] | 3.4629 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031598 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.4694 | [19 empty] | -1.225 | [5 empty] | 3.4629 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031598 | -1.4594 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.4004 | [19 empty] | -1.2251 | [5 empty] | 3.4628 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031597 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.4004 | [19 empty] | -1.2251 | [5 empty] | 3.4628 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031597 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:19 | 20 | 20 |  | 1731.3314 | [19 empty] | -1.2251 | [5 empty] | 3.4627 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031596 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.3314 | [19 empty] | -1.2251 | [5 empty] | 3.4627 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031596 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.3314 | [19 empty] | -1.2251 | [5 empty] | 3.4627 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031596 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.3314 | [19 empty] | -1.2251 | [5 empty] | 3.4627 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031596 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.2625 | [19 empty] | -1.2251 | [5 empty] | 3.4625 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031594 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.2625 | [19 empty] | -1.2251 | [5 empty] | 3.4625 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031594 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.1935 | [19 empty] | -1.2251 | [5 empty] | 3.4624 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031593 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.1935 | [19 empty] | -1.2251 | [5 empty] | 3.4624 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031593 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2251 | [5 empty] | 3.4622 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2251 | [5 empty] | 3.4622 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:20 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2251 | [5 empty] | 3.4622 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2251 | [5 empty] | 3.4622 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2251 | [5 empty] | 3.4622 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1731.0555 | [19 empty] | -1.2251 | [5 empty] | 3.4621 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031591 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1731.0555 | [19 empty] | -1.2251 | [5 empty] | 3.4621 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031591 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1731.0555 | [19 empty] | -1.2251 | [5 empty] | 3.4621 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031591 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1730.9866 | [19 empty] | -1.2251 | [5 empty] | 3.462 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031589 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1730.9866 | [19 empty] | -1.2251 | [5 empty] | 3.462 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031589 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1730.9866 | [19 empty] | -1.2251 | [5 empty] | 3.462 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031589 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1730.9866 | [19 empty] | -1.2251 | [5 empty] | 3.462 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031589 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1730.9176 | [19 empty] | -1.2251 | [5 empty] | 3.4618 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031588 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:21 | 20 | 20 |  | 1730.9176 | [19 empty] | -1.2251 | [5 empty] | 3.4618 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031588 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.9176 | [19 empty] | -1.2251 | [5 empty] | 3.4618 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031588 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.8486 | [19 empty] | -1.2251 | [5 empty] | 3.4617 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031587 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.8486 | [19 empty] | -1.2251 | [5 empty] | 3.4617 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031587 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.8486 | [19 empty] | -1.2251 | [5 empty] | 3.4617 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031587 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.7796 | [19 empty] | -1.2251 | [5 empty] | 3.4616 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031586 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.7796 | [19 empty] | -1.2251 | [5 empty] | 3.4616 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031586 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.7796 | [19 empty] | -1.2251 | [5 empty] | 3.4616 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031586 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.7796 | [19 empty] | -1.2251 | [5 empty] | 3.4616 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031586 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.7107 | [19 empty] | -1.2251 | [5 empty] | 3.4614 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031584 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:22 | 20 | 20 |  | 1730.7107 | [19 empty] | -1.2251 | [5 empty] | 3.4614 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031584 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.7107 | [19 empty] | -1.2251 | [5 empty] | 3.4614 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031584 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.6417 | [19 empty] | -1.2251 | [5 empty] | 3.4613 | -1.225 | 0.90444 | 17.6664 | 0.23773 | -0.0031583 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.6417 | [19 empty] | -1.2251 | [5 empty] | 3.4613 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031583 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.6417 | [19 empty] | -1.2251 | [5 empty] | 3.4613 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031583 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.6417 | [19 empty] | -1.2251 | [5 empty] | 3.4613 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031583 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.5727 | [19 empty] | -1.2251 | [5 empty] | 3.4611 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031582 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.5727 | [19 empty] | -1.2251 | [5 empty] | 3.4611 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031582 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.5727 | [19 empty] | -1.2251 | [5 empty] | 3.4611 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031582 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.5727 | [19 empty] | -1.2251 | [5 empty] | 3.4611 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031582 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:23 | 20 | 20 |  | 1730.5037 | [19 empty] | -1.2251 | [5 empty] | 3.461 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031581 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.5037 | [19 empty] | -1.2251 | [5 empty] | 3.461 | -1.225 | 0.90444 | 17.6664 | 0.23773 | -0.0031581 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.5037 | [19 empty] | -1.2252 | [5 empty] | 3.461 | -1.225 | 0.90444 | 17.6664 | 0.23773 | -0.0031581 | -1.4596 | -1.4403 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.5037 | [19 empty] | -1.2251 | [5 empty] | 3.461 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031581 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.4348 | [19 empty] | -1.2251 | [5 empty] | 3.4609 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031579 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.4348 | [19 empty] | -1.2251 | [5 empty] | 3.4609 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031579 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.4348 | [19 empty] | -1.2251 | [5 empty] | 3.4609 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031579 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.4348 | [19 empty] | -1.2251 | [5 empty] | 3.4609 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031579 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.4348 | [19 empty] | -1.2251 | [5 empty] | 3.4609 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031579 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.3658 | [19 empty] | -1.2251 | [5 empty] | 3.4607 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031578 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2251 | [5 empty] | 3.4606 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:24 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2251 | [5 empty] | 3.4606 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2251 | [5 empty] | 3.4606 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2251 | [5 empty] | 3.4606 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2251 | [5 empty] | 3.4606 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2278 | [19 empty] | -1.2251 | [5 empty] | 3.4605 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031575 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2278 | [19 empty] | -1.2251 | [5 empty] | 3.4605 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031575 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2278 | [19 empty] | -1.2251 | [5 empty] | 3.4605 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031575 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.2278 | [19 empty] | -1.2251 | [5 empty] | 3.4605 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031575 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.1589 | [19 empty] | -1.2251 | [5 empty] | 3.4603 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031574 | -1.4594 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.1589 | [19 empty] | -1.2251 | [5 empty] | 3.4603 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031574 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:25 | 20 | 20 |  | 1730.1589 | [19 empty] | -1.2251 | [5 empty] | 3.4603 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031574 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.1589 | [19 empty] | -1.2251 | [5 empty] | 3.4603 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031574 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0899 | [19 empty] | -1.2251 | [5 empty] | 3.4602 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031573 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0899 | [19 empty] | -1.2251 | [5 empty] | 3.4602 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031573 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0899 | [19 empty] | -1.2251 | [5 empty] | 3.4602 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031573 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0899 | [19 empty] | -1.2251 | [5 empty] | 3.4602 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031573 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0899 | [19 empty] | -1.2251 | [5 empty] | 3.4602 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031573 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0209 | [19 empty] | -1.2251 | [5 empty] | 3.46 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031572 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1730.0209 | [19 empty] | -1.2251 | [5 empty] | 3.46 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031572 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2251 | [5 empty] | 3.4599 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2251 | [5 empty] | 3.4599 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:26 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2251 | [5 empty] | 3.4599 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2251 | [5 empty] | 3.4599 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2252 | [5 empty] | 3.4599 | -1.225 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4595 | -1.4403 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2251 | [5 empty] | 3.4599 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.883 | [19 empty] | -1.2251 | [5 empty] | 3.4598 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031569 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.883 | [19 empty] | -1.2251 | [5 empty] | 3.4598 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031569 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.883 | [19 empty] | -1.2251 | [5 empty] | 3.4598 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031569 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.883 | [19 empty] | -1.2251 | [5 empty] | 3.4598 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031569 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.883 | [19 empty] | -1.2251 | [5 empty] | 3.4598 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031569 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.814 | [19 empty] | -1.2251 | [5 empty] | 3.4596 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031568 | -1.4595 | -1.4402 | -1.4402 | -1.4402 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:27 | 20 | 20 |  | 1729.814 | [19 empty] | -1.2251 | [5 empty] | 3.4596 | -1.2249 | 0.90444 | 17.6664 | 0.23773 | -0.0031568 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 1729.814 | [19 empty] | -1.2251 | [5 empty] | 3.4596 | -1.225 | 0.90444 | 17.6664 | 0.23773 | -0.0031568 | -1.4595 | -1.4402 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 1729.814 | [19 empty] | -1.2255 | [5 empty] | 3.4596 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031568 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 1854.5201 | [19 empty] | -1.2262 | [5 empty] | 3.709 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0033857 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2262 | [5 empty] | 4.1641 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2084.4124 | [19 empty] | -1.2262 | [5 empty] | 4.1688 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038077 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2083.6537 | [19 empty] | -1.2262 | [5 empty] | 4.1673 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038063 | -1.4599 | -1.4406 | -1.4406 | -1.4406 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2262 | [5 empty] | 4.1719 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2263 | [5 empty] | 4.1719 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2085.516 | [19 empty] | -1.2262 | [5 empty] | 4.171 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038097 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:28 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2262 | [5 empty] | 4.1705 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2262 | [5 empty] | 4.1698 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2084.6883 | [19 empty] | -1.2262 | [5 empty] | 4.1694 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038082 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2084.4124 | [19 empty] | -1.2262 | [5 empty] | 4.1688 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038077 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2084.2055 | [19 empty] | -1.2262 | [5 empty] | 4.1684 | -1.226 | 0.90444 | 17.6664 | 0.23773 | -0.0038073 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2083.9986 | [19 empty] | -1.2262 | [5 empty] | 4.168 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038069 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2083.7916 | [19 empty] | -1.2263 | [5 empty] | 4.1676 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038065 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2083.5847 | [19 empty] | -1.2263 | [5 empty] | 4.1672 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038061 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2262 | [5 empty] | 4.1669 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2263 | [5 empty] | 4.1665 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2083.1019 | [19 empty] | -1.2263 | [5 empty] | 4.1662 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038053 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:29 | 20 | 20 |  | 2082.9639 | [19 empty] | -1.2263 | [5 empty] | 4.1659 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.003805 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2263 | [5 empty] | 4.1658 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.757 | [19 empty] | -1.2263 | [5 empty] | 4.1655 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038046 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.6191 | [19 empty] | -1.2263 | [5 empty] | 4.1652 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038044 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.4811 | [19 empty] | -1.2263 | [5 empty] | 4.165 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038041 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.3432 | [19 empty] | -1.2263 | [5 empty] | 4.1647 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038039 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.2052 | [19 empty] | -1.2263 | [5 empty] | 4.1644 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038036 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2263 | [5 empty] | 4.1641 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2081.9983 | [19 empty] | -1.2263 | [5 empty] | 4.164 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038032 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2081.8604 | [19 empty] | -1.2263 | [5 empty] | 4.1637 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.003803 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:30 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2263 | [5 empty] | 4.1636 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2263 | [5 empty] | 4.1633 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.46 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2263 | [5 empty] | 4.1632 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.46 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2263 | [5 empty] | 4.163 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.4465 | [19 empty] | -1.2263 | [5 empty] | 4.1629 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038022 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.3086 | [19 empty] | -1.2264 | [5 empty] | 4.1626 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003802 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.2396 | [19 empty] | -1.2263 | [5 empty] | 4.1625 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038018 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.1016 | [19 empty] | -1.2263 | [5 empty] | 4.1622 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038016 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2263 | [5 empty] | 4.1621 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2080.9637 | [19 empty] | -1.2263 | [5 empty] | 4.1619 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0038013 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:31 | 20 | 20 |  | 2080.8257 | [19 empty] | -1.2264 | [5 empty] | 4.1617 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038011 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.7568 | [19 empty] | -1.2263 | [5 empty] | 4.1615 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.003801 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2264 | [5 empty] | 4.1614 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.5498 | [19 empty] | -1.2264 | [5 empty] | 4.1611 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038006 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.4809 | [19 empty] | -1.2264 | [5 empty] | 4.161 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038005 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.4119 | [19 empty] | -1.2264 | [5 empty] | 4.1608 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038003 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.3429 | [19 empty] | -1.2264 | [5 empty] | 4.1607 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0038002 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.205 | [19 empty] | -1.2263 | [5 empty] | 4.1604 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037999 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2263 | [5 empty] | 4.1603 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2080.067 | [19 empty] | -1.2264 | [5 empty] | 4.1601 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037997 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2079.998 | [19 empty] | -1.2264 | [5 empty] | 4.16 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037996 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:32 | 20 | 20 |  | 2079.9291 | [19 empty] | -1.2264 | [5 empty] | 4.1599 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037994 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2263 | [5 empty] | 4.1597 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.7911 | [19 empty] | -1.2264 | [5 empty] | 4.1596 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037992 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.6532 | [19 empty] | -1.2264 | [5 empty] | 4.1593 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037989 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.5842 | [19 empty] | -1.2263 | [5 empty] | 4.1592 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037988 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.5842 | [19 empty] | -1.2264 | [5 empty] | 4.1592 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037988 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.4462 | [19 empty] | -1.2264 | [5 empty] | 4.1589 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037986 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2264 | [5 empty] | 4.1588 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.3083 | [19 empty] | -1.2264 | [5 empty] | 4.1586 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037983 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2264 | [5 empty] | 4.1585 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:33 | 20 | 20 |  | 2079.1703 | [19 empty] | -1.2263 | [5 empty] | 4.1583 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003798 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2079.1014 | [19 empty] | -1.2264 | [5 empty] | 4.1582 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037979 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2264 | [5 empty] | 4.1581 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2263 | [5 empty] | 4.1581 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.8944 | [19 empty] | -1.2264 | [5 empty] | 4.1578 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037975 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.8944 | [19 empty] | -1.2264 | [5 empty] | 4.1578 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037975 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.8255 | [19 empty] | -1.2264 | [5 empty] | 4.1577 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037974 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.7565 | [19 empty] | -1.2264 | [5 empty] | 4.1575 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037973 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.6875 | [19 empty] | -1.2263 | [5 empty] | 4.1574 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037972 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2264 | [5 empty] | 4.1572 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:34 | 20 | 20 |  | 2078.5496 | [19 empty] | -1.2264 | [5 empty] | 4.1571 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037969 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.5496 | [19 empty] | -1.2263 | [5 empty] | 4.1571 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037969 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.4116 | [19 empty] | -1.2264 | [5 empty] | 4.1568 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037967 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.4116 | [19 empty] | -1.2264 | [5 empty] | 4.1568 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037967 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.3426 | [19 empty] | -1.2264 | [5 empty] | 4.1567 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037965 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2264 | [5 empty] | 4.1565 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.2047 | [19 empty] | -1.2264 | [5 empty] | 4.1564 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037963 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2264 | [5 empty] | 4.1563 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2078.0667 | [19 empty] | -1.2264 | [5 empty] | 4.1561 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003796 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2077.9978 | [19 empty] | -1.2264 | [5 empty] | 4.156 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037959 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2077.9978 | [19 empty] | -1.2264 | [5 empty] | 4.156 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037959 | -1.4601 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:35 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2264 | [5 empty] | 4.1559 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.8598 | [19 empty] | -1.2264 | [5 empty] | 4.1557 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037956 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2264 | [5 empty] | 4.1556 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2264 | [5 empty] | 4.1556 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.7219 | [19 empty] | -1.2264 | [5 empty] | 4.1554 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037954 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.6529 | [19 empty] | -1.2264 | [5 empty] | 4.1553 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037953 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.5839 | [19 empty] | -1.2264 | [5 empty] | 4.1552 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037951 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.515 | [19 empty] | -1.2264 | [5 empty] | 4.155 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003795 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.446 | [19 empty] | -1.2263 | [5 empty] | 4.1549 | -1.2261 | 0.90444 | 17.6664 | 0.23773 | -0.0037949 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.446 | [19 empty] | -1.2264 | [5 empty] | 4.1549 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037949 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:36 | 20 | 20 |  | 2077.377 | [19 empty] | -1.2264 | [5 empty] | 4.1548 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037948 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.308 | [19 empty] | -1.2264 | [5 empty] | 4.1546 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037946 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.2391 | [19 empty] | -1.2265 | [5 empty] | 4.1545 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037945 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.2391 | [19 empty] | -1.2264 | [5 empty] | 4.1545 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037945 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.1701 | [19 empty] | -1.2264 | [5 empty] | 4.1543 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037944 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.1011 | [19 empty] | -1.2264 | [5 empty] | 4.1542 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037942 | -1.4601 | -1.4408 | -1.4408 | -1.4408 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.0321 | [19 empty] | -1.2264 | [5 empty] | 4.1541 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037941 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2077.0321 | [19 empty] | -1.2264 | [5 empty] | 4.1541 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037941 | -1.4601 | -1.4408 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2076.9632 | [19 empty] | -1.2264 | [5 empty] | 4.1539 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003794 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2076.8942 | [19 empty] | -1.2264 | [5 empty] | 4.1538 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037939 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:37 | 20 | 20 |  | 2076.8942 | [19 empty] | -1.2264 | [5 empty] | 4.1538 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037939 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.8252 | [19 empty] | -1.2264 | [5 empty] | 4.1537 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037937 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.7562 | [19 empty] | -1.2264 | [5 empty] | 4.1535 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037936 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.6873 | [19 empty] | -1.2264 | [5 empty] | 4.1534 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037935 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.6873 | [19 empty] | -1.2264 | [5 empty] | 4.1534 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037935 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.6183 | [19 empty] | -1.2264 | [5 empty] | 4.1532 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037934 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.5493 | [19 empty] | -1.2264 | [5 empty] | 4.1531 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037932 | -1.4601 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.5493 | [19 empty] | -1.2264 | [5 empty] | 4.1531 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037932 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.4803 | [19 empty] | -1.2264 | [5 empty] | 4.153 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037931 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2264 | [5 empty] | 4.1528 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2264 | [5 empty] | 4.1528 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:38 | 20 | 20 |  | 2076.3424 | [19 empty] | -1.2264 | [5 empty] | 4.1527 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037929 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.3424 | [19 empty] | -1.2264 | [5 empty] | 4.1527 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037929 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.2734 | [19 empty] | -1.2264 | [5 empty] | 4.1525 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037927 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.2044 | [19 empty] | -1.2264 | [5 empty] | 4.1524 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037926 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.2044 | [19 empty] | -1.2264 | [5 empty] | 4.1524 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037926 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2264 | [5 empty] | 4.1523 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2265 | [5 empty] | 4.1523 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2265 | [5 empty] | 4.1521 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2265 | [5 empty] | 4.1521 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2265 | [5 empty] | 4.152 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:39 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2264 | [5 empty] | 4.152 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.9285 | [19 empty] | -1.2265 | [5 empty] | 4.1519 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037921 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.9285 | [19 empty] | -1.2265 | [5 empty] | 4.1519 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037921 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2264 | [5 empty] | 4.1517 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2265 | [5 empty] | 4.1517 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.7906 | [19 empty] | -1.2265 | [5 empty] | 4.1516 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037918 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.7906 | [19 empty] | -1.2265 | [5 empty] | 4.1516 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037918 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2264 | [5 empty] | 4.1514 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2265 | [5 empty] | 4.1514 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2265 | [5 empty] | 4.1513 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2264 | [5 empty] | 4.1513 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:40 | 20 | 20 |  | 2075.5837 | [19 empty] | -1.2265 | [5 empty] | 4.1512 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037915 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.5837 | [19 empty] | -1.2265 | [5 empty] | 4.1512 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037915 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2265 | [5 empty] | 4.151 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2265 | [5 empty] | 4.151 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.4457 | [19 empty] | -1.2265 | [5 empty] | 4.1509 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037912 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.4457 | [19 empty] | -1.2265 | [5 empty] | 4.1509 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037912 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2265 | [5 empty] | 4.1508 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2265 | [5 empty] | 4.1508 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2265 | [5 empty] | 4.1508 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2265 | [5 empty] | 4.1506 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:41 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2265 | [5 empty] | 4.1506 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.2388 | [19 empty] | -1.2265 | [5 empty] | 4.1505 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037908 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.2388 | [19 empty] | -1.2265 | [5 empty] | 4.1505 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037908 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2265 | [5 empty] | 4.1503 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2265 | [5 empty] | 4.1503 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2265 | [5 empty] | 4.1502 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2265 | [5 empty] | 4.1502 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.0319 | [19 empty] | -1.2265 | [5 empty] | 4.1501 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037904 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2075.0319 | [19 empty] | -1.2264 | [5 empty] | 4.1501 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037904 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2265 | [5 empty] | 4.1499 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:42 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2265 | [5 empty] | 4.1499 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.8939 | [19 empty] | -1.2265 | [5 empty] | 4.1498 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037902 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.8939 | [19 empty] | -1.2265 | [5 empty] | 4.1498 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037902 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2265 | [5 empty] | 4.1496 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2264 | [5 empty] | 4.1496 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2265 | [5 empty] | 4.1496 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2265 | [5 empty] | 4.1496 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.756 | [19 empty] | -1.2265 | [5 empty] | 4.1495 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037899 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.756 | [19 empty] | -1.2265 | [5 empty] | 4.1495 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037899 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2265 | [5 empty] | 4.1494 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:43 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2265 | [5 empty] | 4.1494 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2265 | [5 empty] | 4.1494 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2265 | [5 empty] | 4.1492 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2265 | [5 empty] | 4.1492 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2265 | [5 empty] | 4.1492 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2265 | [5 empty] | 4.1491 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2265 | [5 empty] | 4.1491 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2265 | [5 empty] | 4.1491 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2265 | [5 empty] | 4.149 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2265 | [5 empty] | 4.149 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2265 | [5 empty] | 4.149 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:44 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2265 | [5 empty] | 4.1488 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2265 | [5 empty] | 4.1488 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2265 | [5 empty] | 4.1487 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2265 | [5 empty] | 4.1487 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2265 | [5 empty] | 4.1487 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2265 | [5 empty] | 4.1485 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2265 | [5 empty] | 4.1485 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2265 | [5 empty] | 4.1485 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2265 | [5 empty] | 4.1484 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2265 | [5 empty] | 4.1484 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:45 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2265 | [5 empty] | 4.1484 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2265 | [5 empty] | 4.1484 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2265 | [5 empty] | 4.1483 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2265 | [5 empty] | 4.1483 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2074.0662 | [19 empty] | -1.2265 | [5 empty] | 4.1481 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037887 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2074.0662 | [19 empty] | -1.2265 | [5 empty] | 4.1481 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037887 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2265 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2265 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2265 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2265 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2265 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:46 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2265 | [5 empty] | 4.1479 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2265 | [5 empty] | 4.1479 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2265 | [5 empty] | 4.1479 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2265 | [5 empty] | 4.1479 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2265 | [5 empty] | 4.1477 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2265 | [5 empty] | 4.1477 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2265 | [5 empty] | 4.1477 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2265 | [5 empty] | 4.1477 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2265 | [5 empty] | 4.1476 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2265 | [5 empty] | 4.1476 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:47 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2265 | [5 empty] | 4.1476 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2265 | [5 empty] | 4.1476 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2265 | [5 empty] | 4.1474 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2265 | [5 empty] | 4.1474 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2265 | [5 empty] | 4.1474 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2265 | [5 empty] | 4.1474 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2265 | [5 empty] | 4.1473 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2265 | [5 empty] | 4.1473 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2265 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2265 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:48 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2265 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2265 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2265 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2265 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2265 | [5 empty] | 4.147 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2265 | [5 empty] | 4.147 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2265 | [5 empty] | 4.147 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2265 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2265 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2265 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2265 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:49 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2265 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2265 | [5 empty] | 4.1468 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2265 | [5 empty] | 4.1468 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2265 | [5 empty] | 4.1468 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2265 | [5 empty] | 4.1468 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2265 | [5 empty] | 4.1466 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2265 | [5 empty] | 4.1466 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2265 | [5 empty] | 4.1466 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2265 | [5 empty] | 4.1465 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2266 | [5 empty] | 4.1465 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:50 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2265 | [5 empty] | 4.1465 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2265 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2265 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2266 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2265 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2265 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2265 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:51 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2266 | [5 empty] | 4.1461 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2266 | [5 empty] | 4.1459 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2266 | [5 empty] | 4.1459 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2265 | [5 empty] | 4.1459 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:52 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2265 | [5 empty] | 4.1456 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2265 | [5 empty] | 4.1456 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2265 | [5 empty] | 4.1456 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2265 | [5 empty] | 4.1455 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4602 | -1.4409 | -1.4409 | -1.4409 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2265 | [5 empty] | 4.1455 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2265 | [5 empty] | 4.1455 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2265 | [5 empty] | 4.1454 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:53 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2265 | [5 empty] | 4.1454 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2265 | [5 empty] | 4.1454 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2265 | [5 empty] | 4.1454 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2265 | [5 empty] | 4.1454 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2266 | [5 empty] | 4.1452 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2265 | [5 empty] | 4.1452 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2265 | [5 empty] | 4.1452 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2265 | [5 empty] | 4.1451 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2265 | [5 empty] | 4.1451 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2265 | [5 empty] | 4.1451 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2265 | [5 empty] | 4.1451 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:54 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2265 | [5 empty] | 4.1451 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2265 | [5 empty] | 4.1451 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4602 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:55 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2265 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2265 | [5 empty] | 4.1448 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2265 | [5 empty] | 4.1448 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2265 | [5 empty] | 4.1448 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2266 | [5 empty] | 4.1448 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:56 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2266 | [5 empty] | 4.1448 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2265 | [5 empty] | 4.1447 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2265 | [5 empty] | 4.1447 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2265 | [5 empty] | 4.1447 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2266 | [5 empty] | 4.1445 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2266 | [5 empty] | 4.1445 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2265 | [5 empty] | 4.1445 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:57 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2265 | [5 empty] | 4.1445 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2265 | [5 empty] | 4.1445 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2266 | [5 empty] | 4.1445 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2265 | [5 empty] | 4.1444 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2265 | [5 empty] | 4.1444 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2265 | [5 empty] | 4.1444 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2266 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:58 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2265 | [5 empty] | 4.1443 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2266 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2265 | [5 empty] | 4.1443 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2265 | [5 empty] | 4.1443 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2265 | [5 empty] | 4.1443 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2265 | [5 empty] | 4.1443 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2265 | [5 empty] | 4.1441 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2266 | [5 empty] | 4.1441 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2266 | [5 empty] | 4.1441 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2265 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 17:59 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2265 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2265 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2265 | [5 empty] | 4.1439 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:00 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2265 | [5 empty] | 4.1439 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2265 | [5 empty] | 4.1437 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2266 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2266 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2266 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:01 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2265 | [5 empty] | 4.1436 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2265 | [5 empty] | 4.1436 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:02 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2265 | [5 empty] | 4.1434 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2265 | [5 empty] | 4.1434 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2265 | [5 empty] | 4.1434 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2265 | [5 empty] | 4.1433 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:03 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:04 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:05 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2265 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2265 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2265 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2265 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:06 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:07 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2265 | [5 empty] | 4.1426 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:08 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2265 | [5 empty] | 4.1425 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:09 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2267 | [5 empty] | 4.1422 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:10 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:11 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:12 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2267 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:13 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2267 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2266 | [5 empty] | 4.1419 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:14 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.0313 | [19 empty] | -1.2266 | [5 empty] | 4.1421 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037831 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:15 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2266 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2267 | [5 empty] | 4.1422 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:16 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2266 | [5 empty] | 4.1423 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2266 | [5 empty] | 4.1425 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2267 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2267 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2267 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:17 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2266 | [5 empty] | 4.1426 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2267 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2266 | [5 empty] | 4.1428 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2267 | [5 empty] | 4.1429 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2266 | [5 empty] | 4.1429 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2267 | [5 empty] | 4.1429 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:18 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2267 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2267 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2266 | [5 empty] | 4.143 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2267 | [5 empty] | 4.1432 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2267 | [5 empty] | 4.1432 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2266 | [5 empty] | 4.1432 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:19 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2267 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2266 | [5 empty] | 4.1433 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2266 | [5 empty] | 4.1434 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:20 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2266 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2267 | [5 empty] | 4.1436 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2266 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2267 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2267 | [5 empty] | 4.1437 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2267 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2267 | [5 empty] | 4.1437 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2267 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:21 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2266 | [5 empty] | 4.1439 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2267 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2266 | [5 empty] | 4.144 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2267 | [5 empty] | 4.1441 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2267 | [5 empty] | 4.1441 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2266 | [5 empty] | 4.1441 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:22 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2267 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2266 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2265 | [5 empty] | 4.1443 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2266 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2266 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2266 | [5 empty] | 4.1443 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2267 | [5 empty] | 4.1444 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2266 | [5 empty] | 4.1444 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2266 | [5 empty] | 4.1445 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:23 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2267 | [5 empty] | 4.1445 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2266 | [5 empty] | 4.1445 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2266 | [5 empty] | 4.1445 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2267 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2266 | [5 empty] | 4.1447 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2267 | [5 empty] | 4.1448 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:24 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2266 | [5 empty] | 4.1448 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2266 | [5 empty] | 4.1448 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2267 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2266 | [5 empty] | 4.145 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2266 | [5 empty] | 4.1451 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2267 | [5 empty] | 4.1451 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2267 | [5 empty] | 4.1451 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:25 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2267 | [5 empty] | 4.1451 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2266 | [5 empty] | 4.1451 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2266 | [5 empty] | 4.1451 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2266 | [5 empty] | 4.1452 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2266 | [5 empty] | 4.1452 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2266 | [5 empty] | 4.1452 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2266 | [5 empty] | 4.1452 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2266 | [5 empty] | 4.1454 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2267 | [5 empty] | 4.1454 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2267 | [5 empty] | 4.1454 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2267 | [5 empty] | 4.1454 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:26 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2266 | [5 empty] | 4.1455 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2266 | [5 empty] | 4.1455 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2267 | [5 empty] | 4.1455 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2266 | [5 empty] | 4.1455 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2267 | [5 empty] | 4.1455 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2267 | [5 empty] | 4.1456 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2267 | [5 empty] | 4.1456 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2266 | [5 empty] | 4.1456 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2266 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2266 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:27 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2265 | [5 empty] | 4.1458 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2265 | [5 empty] | 4.1459 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2265 | [5 empty] | 4.1459 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2266 | [5 empty] | 4.1459 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2266 | [5 empty] | 4.1459 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2266 | [5 empty] | 4.1459 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2266 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2266 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2265 | [5 empty] | 4.1461 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:28 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2265 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2265 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4602 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2266 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2266 | [5 empty] | 4.1462 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2265 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2266 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2265 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2266 | [5 empty] | 4.1463 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2266 | [5 empty] | 4.1465 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2266 | [5 empty] | 4.1465 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:29 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2266 | [5 empty] | 4.1465 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2267 | [5 empty] | 4.1465 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2266 | [5 empty] | 4.1466 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2266 | [5 empty] | 4.1466 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2266 | [5 empty] | 4.1466 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2266 | [5 empty] | 4.1468 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2266 | [5 empty] | 4.1468 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2266 | [5 empty] | 4.1468 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2266 | [5 empty] | 4.1469 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2266 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:30 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2266 | [5 empty] | 4.1469 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2266 | [5 empty] | 4.1469 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2266 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2266 | [5 empty] | 4.1469 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2266 | [5 empty] | 4.147 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2266 | [5 empty] | 4.147 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2266 | [5 empty] | 4.147 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2266 | [5 empty] | 4.1472 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2266 | [5 empty] | 4.1472 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2266 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:31 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2266 | [5 empty] | 4.1472 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2266 | [5 empty] | 4.1473 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2266 | [5 empty] | 4.1473 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2266 | [5 empty] | 4.1473 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2266 | [5 empty] | 4.1473 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2266 | [5 empty] | 4.1473 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2266 | [5 empty] | 4.1474 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2266 | [5 empty] | 4.1474 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2266 | [5 empty] | 4.1474 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2266 | [5 empty] | 4.1476 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2266 | [5 empty] | 4.1476 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:32 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2266 | [5 empty] | 4.1476 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2266 | [5 empty] | 4.1477 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2266 | [5 empty] | 4.1476 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2266 | [5 empty] | 4.1477 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2266 | [5 empty] | 4.1477 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2266 | [5 empty] | 4.1479 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2266 | [5 empty] | 4.1479 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2266 | [5 empty] | 4.1479 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2266 | [5 empty] | 4.1479 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2266 | [5 empty] | 4.1479 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:33 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2266 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2266 | [5 empty] | 4.148 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2266 | [5 empty] | 4.148 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.0662 | [19 empty] | -1.2266 | [5 empty] | 4.1481 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037887 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.0662 | [19 empty] | -1.2266 | [5 empty] | 4.1481 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037887 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.0662 | [19 empty] | -1.2266 | [5 empty] | 4.1481 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037887 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2266 | [5 empty] | 4.1483 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2266 | [5 empty] | 4.1483 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2266 | [5 empty] | 4.1483 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2266 | [5 empty] | 4.1483 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2266 | [5 empty] | 4.1483 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:34 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2266 | [5 empty] | 4.1484 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2266 | [5 empty] | 4.1484 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2266 | [5 empty] | 4.1484 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2266 | [5 empty] | 4.1485 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2266 | [5 empty] | 4.1485 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2266 | [5 empty] | 4.1485 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2266 | [5 empty] | 4.1485 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2266 | [5 empty] | 4.1487 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2266 | [5 empty] | 4.1487 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2266 | [5 empty] | 4.1487 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:35 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2266 | [5 empty] | 4.1487 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2266 | [5 empty] | 4.1488 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2266 | [5 empty] | 4.1488 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2266 | [5 empty] | 4.1488 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2266 | [5 empty] | 4.1488 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2266 | [5 empty] | 4.1488 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2266 | [5 empty] | 4.149 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2266 | [5 empty] | 4.149 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2266 | [5 empty] | 4.149 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2266 | [5 empty] | 4.1491 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:36 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2266 | [5 empty] | 4.1491 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2266 | [5 empty] | 4.1491 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2266 | [5 empty] | 4.1491 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2266 | [5 empty] | 4.1492 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2267 | [5 empty] | 4.1492 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2266 | [5 empty] | 4.1492 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2266 | [5 empty] | 4.1492 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2266 | [5 empty] | 4.1492 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2266 | [5 empty] | 4.1494 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2266 | [5 empty] | 4.1494 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2266 | [5 empty] | 4.1494 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:37 | 20 | 20 |  | 2074.756 | [19 empty] | -1.2266 | [5 empty] | 4.1495 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037899 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.756 | [19 empty] | -1.2267 | [5 empty] | 4.1495 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037899 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.756 | [19 empty] | -1.2267 | [5 empty] | 4.1495 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037899 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2266 | [5 empty] | 4.1496 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2266 | [5 empty] | 4.1496 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2266 | [5 empty] | 4.1496 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2266 | [5 empty] | 4.1496 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8939 | [19 empty] | -1.2266 | [5 empty] | 4.1498 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037902 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8939 | [19 empty] | -1.2266 | [5 empty] | 4.1498 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037902 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.8939 | [19 empty] | -1.2266 | [5 empty] | 4.1498 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037902 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:38 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2266 | [5 empty] | 4.1499 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2266 | [5 empty] | 4.1499 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2266 | [5 empty] | 4.1499 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2266 | [5 empty] | 4.1499 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2266 | [5 empty] | 4.1499 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2075.0319 | [19 empty] | -1.2266 | [5 empty] | 4.1501 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037904 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2075.0319 | [19 empty] | -1.2266 | [5 empty] | 4.1501 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037904 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2075.0319 | [19 empty] | -1.2266 | [5 empty] | 4.1501 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037904 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2266 | [5 empty] | 4.1502 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2266 | [5 empty] | 4.1502 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:39 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2266 | [5 empty] | 4.1502 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2266 | [5 empty] | 4.1502 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2266 | [5 empty] | 4.1503 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2267 | [5 empty] | 4.1503 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2266 | [5 empty] | 4.1503 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.2388 | [19 empty] | -1.2266 | [5 empty] | 4.1505 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037908 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.2388 | [19 empty] | -1.2266 | [5 empty] | 4.1505 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037908 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2267 | [5 empty] | 4.1506 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2267 | [5 empty] | 4.1506 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2266 | [5 empty] | 4.1506 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2266 | [5 empty] | 4.1506 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:40 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2267 | [5 empty] | 4.1508 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2266 | [5 empty] | 4.1508 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2267 | [5 empty] | 4.1508 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2266 | [5 empty] | 4.1508 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.4457 | [19 empty] | -1.2266 | [5 empty] | 4.1509 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037912 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.4457 | [19 empty] | -1.2267 | [5 empty] | 4.1509 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037912 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2266 | [5 empty] | 4.151 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2266 | [5 empty] | 4.151 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2266 | [5 empty] | 4.151 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2266 | [5 empty] | 4.151 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:41 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2266 | [5 empty] | 4.151 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.5837 | [19 empty] | -1.2266 | [5 empty] | 4.1512 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037915 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.5837 | [19 empty] | -1.2267 | [5 empty] | 4.1512 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037915 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.5837 | [19 empty] | -1.2267 | [5 empty] | 4.1512 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037915 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2266 | [5 empty] | 4.1513 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2266 | [5 empty] | 4.1513 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2266 | [5 empty] | 4.1513 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2267 | [5 empty] | 4.1514 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2266 | [5 empty] | 4.1514 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2266 | [5 empty] | 4.1514 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2266 | [5 empty] | 4.1514 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:42 | 20 | 20 |  | 2075.7906 | [19 empty] | -1.2266 | [5 empty] | 4.1516 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037918 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.7906 | [19 empty] | -1.2266 | [5 empty] | 4.1516 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037918 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.7906 | [19 empty] | -1.2266 | [5 empty] | 4.1516 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037918 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2266 | [5 empty] | 4.1517 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2267 | [5 empty] | 4.1517 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2267 | [5 empty] | 4.1517 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2266 | [5 empty] | 4.1517 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.9285 | [19 empty] | -1.2266 | [5 empty] | 4.1519 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037921 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.9285 | [19 empty] | -1.2266 | [5 empty] | 4.1519 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037921 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.9285 | [19 empty] | -1.2267 | [5 empty] | 4.1519 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037921 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:43 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2267 | [5 empty] | 4.152 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2267 | [5 empty] | 4.152 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2266 | [5 empty] | 4.152 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2267 | [5 empty] | 4.152 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2266 | [5 empty] | 4.1521 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2267 | [5 empty] | 4.1521 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2266 | [5 empty] | 4.1521 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2267 | [5 empty] | 4.1523 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2266 | [5 empty] | 4.1523 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2266 | [5 empty] | 4.1523 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:44 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2267 | [5 empty] | 4.1523 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.2044 | [19 empty] | -1.2266 | [5 empty] | 4.1524 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037926 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.2044 | [19 empty] | -1.2266 | [5 empty] | 4.1524 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037926 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.2044 | [19 empty] | -1.2266 | [5 empty] | 4.1524 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037926 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.2734 | [19 empty] | -1.2267 | [5 empty] | 4.1525 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037927 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.2734 | [19 empty] | -1.2266 | [5 empty] | 4.1525 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037927 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.2734 | [19 empty] | -1.2266 | [5 empty] | 4.1525 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037927 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.3424 | [19 empty] | -1.2266 | [5 empty] | 4.1527 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037929 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.3424 | [19 empty] | -1.2266 | [5 empty] | 4.1527 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037929 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.3424 | [19 empty] | -1.2266 | [5 empty] | 4.1527 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037929 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:45 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2267 | [5 empty] | 4.1528 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2267 | [5 empty] | 4.1528 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2267 | [5 empty] | 4.1528 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2266 | [5 empty] | 4.1528 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2267 | [5 empty] | 4.1528 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4803 | [19 empty] | -1.2266 | [5 empty] | 4.153 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037931 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4803 | [19 empty] | -1.2266 | [5 empty] | 4.153 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037931 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4803 | [19 empty] | -1.2266 | [5 empty] | 4.153 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037931 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.4803 | [19 empty] | -1.2266 | [5 empty] | 4.153 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037931 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.5493 | [19 empty] | -1.2267 | [5 empty] | 4.1531 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037932 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.5493 | [19 empty] | -1.2266 | [5 empty] | 4.1531 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037932 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:46 | 20 | 20 |  | 2076.5493 | [19 empty] | -1.2267 | [5 empty] | 4.1531 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037932 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.6183 | [19 empty] | -1.2267 | [5 empty] | 4.1532 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037934 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.6183 | [19 empty] | -1.2266 | [5 empty] | 4.1532 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037934 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.6183 | [19 empty] | -1.2267 | [5 empty] | 4.1532 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037934 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.6873 | [19 empty] | -1.2267 | [5 empty] | 4.1534 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037935 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.6873 | [19 empty] | -1.2266 | [5 empty] | 4.1534 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037935 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.6873 | [19 empty] | -1.2266 | [5 empty] | 4.1534 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037935 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.7562 | [19 empty] | -1.2266 | [5 empty] | 4.1535 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037936 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.7562 | [19 empty] | -1.2266 | [5 empty] | 4.1535 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037936 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.7562 | [19 empty] | -1.2267 | [5 empty] | 4.1535 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037936 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:47 | 20 | 20 |  | 2076.7562 | [19 empty] | -1.2267 | [5 empty] | 4.1535 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037936 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.8252 | [19 empty] | -1.2267 | [5 empty] | 4.1537 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037937 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.8252 | [19 empty] | -1.2267 | [5 empty] | 4.1537 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037937 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.8252 | [19 empty] | -1.2267 | [5 empty] | 4.1537 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037937 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.8942 | [19 empty] | -1.2266 | [5 empty] | 4.1538 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037939 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.8942 | [19 empty] | -1.2267 | [5 empty] | 4.1538 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037939 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.8942 | [19 empty] | -1.2267 | [5 empty] | 4.1538 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037939 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.9632 | [19 empty] | -1.2266 | [5 empty] | 4.1539 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003794 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.9632 | [19 empty] | -1.2267 | [5 empty] | 4.1539 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003794 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.9632 | [19 empty] | -1.2267 | [5 empty] | 4.1539 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003794 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2076.9632 | [19 empty] | -1.2267 | [5 empty] | 4.1539 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003794 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:48 | 20 | 20 |  | 2077.0321 | [19 empty] | -1.2266 | [5 empty] | 4.1541 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037941 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.0321 | [19 empty] | -1.2267 | [5 empty] | 4.1541 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037941 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.0321 | [19 empty] | -1.2267 | [5 empty] | 4.1541 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037941 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1011 | [19 empty] | -1.2267 | [5 empty] | 4.1542 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037942 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1011 | [19 empty] | -1.2267 | [5 empty] | 4.1542 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037942 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1011 | [19 empty] | -1.2266 | [5 empty] | 4.1542 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037942 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1701 | [19 empty] | -1.2266 | [5 empty] | 4.1543 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037944 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1701 | [19 empty] | -1.2267 | [5 empty] | 4.1543 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037944 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1701 | [19 empty] | -1.2267 | [5 empty] | 4.1543 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037944 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.1701 | [19 empty] | -1.2266 | [5 empty] | 4.1543 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037944 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:49 | 20 | 20 |  | 2077.2391 | [19 empty] | -1.2267 | [5 empty] | 4.1545 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037945 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.2391 | [19 empty] | -1.2267 | [5 empty] | 4.1545 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037945 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.2391 | [19 empty] | -1.2267 | [5 empty] | 4.1545 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037945 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.308 | [19 empty] | -1.2266 | [5 empty] | 4.1546 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037946 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.308 | [19 empty] | -1.2267 | [5 empty] | 4.1546 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037946 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.377 | [19 empty] | -1.2266 | [5 empty] | 4.1548 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037948 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.377 | [19 empty] | -1.2267 | [5 empty] | 4.1548 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037948 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.377 | [19 empty] | -1.2267 | [5 empty] | 4.1548 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037948 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.377 | [19 empty] | -1.2266 | [5 empty] | 4.1548 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037948 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.446 | [19 empty] | -1.2267 | [5 empty] | 4.1549 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037949 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:50 | 20 | 20 |  | 2077.446 | [19 empty] | -1.2266 | [5 empty] | 4.1549 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037949 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.446 | [19 empty] | -1.2267 | [5 empty] | 4.1549 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037949 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.515 | [19 empty] | -1.2267 | [5 empty] | 4.155 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003795 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.515 | [19 empty] | -1.2267 | [5 empty] | 4.155 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003795 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.515 | [19 empty] | -1.2266 | [5 empty] | 4.155 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003795 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.5839 | [19 empty] | -1.2267 | [5 empty] | 4.1552 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037951 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.5839 | [19 empty] | -1.2267 | [5 empty] | 4.1552 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037951 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.5839 | [19 empty] | -1.2267 | [5 empty] | 4.1552 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037951 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.5839 | [19 empty] | -1.2267 | [5 empty] | 4.1552 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037951 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.6529 | [19 empty] | -1.2267 | [5 empty] | 4.1553 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037953 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.6529 | [19 empty] | -1.2267 | [5 empty] | 4.1553 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037953 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:51 | 20 | 20 |  | 2077.6529 | [19 empty] | -1.2267 | [5 empty] | 4.1553 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037953 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7219 | [19 empty] | -1.2267 | [5 empty] | 4.1554 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037954 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7219 | [19 empty] | -1.2267 | [5 empty] | 4.1554 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037954 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7219 | [19 empty] | -1.2267 | [5 empty] | 4.1554 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037954 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7219 | [19 empty] | -1.2268 | [5 empty] | 4.1554 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037954 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2267 | [5 empty] | 4.1556 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2267 | [5 empty] | 4.1556 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2268 | [5 empty] | 4.1556 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2267 | [5 empty] | 4.1556 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.8598 | [19 empty] | -1.2268 | [5 empty] | 4.1557 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037956 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:52 | 20 | 20 |  | 2077.8598 | [19 empty] | -1.2268 | [5 empty] | 4.1557 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037956 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.8598 | [19 empty] | -1.2268 | [5 empty] | 4.1557 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037956 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2268 | [5 empty] | 4.1559 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2268 | [5 empty] | 4.1559 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2268 | [5 empty] | 4.1559 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2268 | [5 empty] | 4.1559 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9978 | [19 empty] | -1.2267 | [5 empty] | 4.156 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037959 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9978 | [19 empty] | -1.2267 | [5 empty] | 4.156 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037959 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2077.9978 | [19 empty] | -1.2268 | [5 empty] | 4.156 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037959 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2078.0667 | [19 empty] | -1.2268 | [5 empty] | 4.1561 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003796 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:53 | 20 | 20 |  | 2078.0667 | [19 empty] | -1.2268 | [5 empty] | 4.1561 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003796 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.0667 | [19 empty] | -1.2268 | [5 empty] | 4.1561 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003796 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2267 | [5 empty] | 4.1563 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2267 | [5 empty] | 4.1563 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2267 | [5 empty] | 4.1563 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2267 | [5 empty] | 4.1563 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2267 | [5 empty] | 4.1563 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.2047 | [19 empty] | -1.2268 | [5 empty] | 4.1564 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037963 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.2047 | [19 empty] | -1.2268 | [5 empty] | 4.1564 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037963 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2268 | [5 empty] | 4.1565 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2268 | [5 empty] | 4.1565 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:54 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2268 | [5 empty] | 4.1565 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2267 | [5 empty] | 4.1565 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.3426 | [19 empty] | -1.2267 | [5 empty] | 4.1567 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037965 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.3426 | [19 empty] | -1.2268 | [5 empty] | 4.1567 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037965 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.3426 | [19 empty] | -1.2267 | [5 empty] | 4.1567 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037965 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.3426 | [19 empty] | -1.2268 | [5 empty] | 4.1567 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037965 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.4116 | [19 empty] | -1.2268 | [5 empty] | 4.1568 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037967 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.4116 | [19 empty] | -1.2268 | [5 empty] | 4.1568 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037967 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.4116 | [19 empty] | -1.2268 | [5 empty] | 4.1568 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037967 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.4806 | [19 empty] | -1.2268 | [5 empty] | 4.157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037968 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:55 | 20 | 20 |  | 2078.4806 | [19 empty] | -1.2267 | [5 empty] | 4.157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037968 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.4806 | [19 empty] | -1.2268 | [5 empty] | 4.157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037968 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.4806 | [19 empty] | -1.2267 | [5 empty] | 4.157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037968 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.5496 | [19 empty] | -1.2267 | [5 empty] | 4.1571 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037969 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.5496 | [19 empty] | -1.2267 | [5 empty] | 4.1571 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037969 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2267 | [5 empty] | 4.1572 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2268 | [5 empty] | 4.1572 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2267 | [5 empty] | 4.1572 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2267 | [5 empty] | 4.1572 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6875 | [19 empty] | -1.2268 | [5 empty] | 4.1574 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037972 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6875 | [19 empty] | -1.2268 | [5 empty] | 4.1574 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037972 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:56 | 20 | 20 |  | 2078.6875 | [19 empty] | -1.2268 | [5 empty] | 4.1574 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037972 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.6875 | [19 empty] | -1.2267 | [5 empty] | 4.1574 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037972 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.7565 | [19 empty] | -1.2268 | [5 empty] | 4.1575 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037973 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.7565 | [19 empty] | -1.2268 | [5 empty] | 4.1575 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037973 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.7565 | [19 empty] | -1.2268 | [5 empty] | 4.1575 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037973 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.7565 | [19 empty] | -1.2268 | [5 empty] | 4.1575 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037973 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.8255 | [19 empty] | -1.2268 | [5 empty] | 4.1577 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037974 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.8255 | [19 empty] | -1.2268 | [5 empty] | 4.1577 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037974 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.8255 | [19 empty] | -1.2268 | [5 empty] | 4.1577 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037974 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.8944 | [19 empty] | -1.2268 | [5 empty] | 4.1578 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037975 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:57 | 20 | 20 |  | 2078.8944 | [19 empty] | -1.2267 | [5 empty] | 4.1578 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037975 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2078.8944 | [19 empty] | -1.2268 | [5 empty] | 4.1578 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037975 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2078.9634 | [19 empty] | -1.2268 | [5 empty] | 4.1579 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037977 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2078.9634 | [19 empty] | -1.2267 | [5 empty] | 4.1579 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037977 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2078.9634 | [19 empty] | -1.2268 | [5 empty] | 4.1579 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037977 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2078.9634 | [19 empty] | -1.2268 | [5 empty] | 4.1579 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037977 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2268 | [5 empty] | 4.1581 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2267 | [5 empty] | 4.1581 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2268 | [5 empty] | 4.1581 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2268 | [5 empty] | 4.1581 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:58 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2268 | [5 empty] | 4.1581 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.1014 | [19 empty] | -1.2267 | [5 empty] | 4.1582 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037979 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.1014 | [19 empty] | -1.2267 | [5 empty] | 4.1582 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037979 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.1703 | [19 empty] | -1.2267 | [5 empty] | 4.1583 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003798 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.1703 | [19 empty] | -1.2268 | [5 empty] | 4.1583 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003798 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.1703 | [19 empty] | -1.2268 | [5 empty] | 4.1583 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003798 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.1703 | [19 empty] | -1.2268 | [5 empty] | 4.1583 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003798 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2268 | [5 empty] | 4.1585 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2268 | [5 empty] | 4.1585 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2267 | [5 empty] | 4.1585 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2267 | [5 empty] | 4.1585 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 18:59 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2268 | [5 empty] | 4.1585 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.3083 | [19 empty] | -1.2267 | [5 empty] | 4.1586 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037983 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.3083 | [19 empty] | -1.2268 | [5 empty] | 4.1586 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037983 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2268 | [5 empty] | 4.1588 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2268 | [5 empty] | 4.1588 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2268 | [5 empty] | 4.1588 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2267 | [5 empty] | 4.1588 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.4462 | [19 empty] | -1.2268 | [5 empty] | 4.1589 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037986 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.4462 | [19 empty] | -1.2268 | [5 empty] | 4.1589 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037986 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.4462 | [19 empty] | -1.2268 | [5 empty] | 4.1589 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037986 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:00 | 20 | 20 |  | 2079.5152 | [19 empty] | -1.2268 | [5 empty] | 4.159 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037987 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.5152 | [19 empty] | -1.2268 | [5 empty] | 4.159 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037987 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.5152 | [19 empty] | -1.2268 | [5 empty] | 4.159 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037987 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.5152 | [19 empty] | -1.2268 | [5 empty] | 4.159 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037987 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.5842 | [19 empty] | -1.2268 | [5 empty] | 4.1592 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037988 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.5842 | [19 empty] | -1.2268 | [5 empty] | 4.1592 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037988 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.5842 | [19 empty] | -1.2268 | [5 empty] | 4.1592 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037988 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.6532 | [19 empty] | -1.2269 | [5 empty] | 4.1593 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037989 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.6532 | [19 empty] | -1.2269 | [5 empty] | 4.1593 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037989 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.6532 | [19 empty] | -1.2268 | [5 empty] | 4.1593 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037989 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:01 | 20 | 20 |  | 2079.7221 | [19 empty] | -1.2268 | [5 empty] | 4.1594 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037991 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.7221 | [19 empty] | -1.2269 | [5 empty] | 4.1594 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037991 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.7221 | [19 empty] | -1.2268 | [5 empty] | 4.1594 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037991 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.7221 | [19 empty] | -1.2268 | [5 empty] | 4.1594 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037991 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.7911 | [19 empty] | -1.2268 | [5 empty] | 4.1596 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037992 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.7911 | [19 empty] | -1.2268 | [5 empty] | 4.1596 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037992 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.7911 | [19 empty] | -1.2268 | [5 empty] | 4.1596 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037992 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2267 | [5 empty] | 4.1597 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2268 | [5 empty] | 4.1597 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2268 | [5 empty] | 4.1597 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2268 | [5 empty] | 4.1597 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:02 | 20 | 20 |  | 2079.9291 | [19 empty] | -1.2268 | [5 empty] | 4.1599 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037994 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2079.9291 | [19 empty] | -1.2268 | [5 empty] | 4.1599 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037994 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2079.9291 | [19 empty] | -1.2268 | [5 empty] | 4.1599 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037994 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2079.9291 | [19 empty] | -1.2268 | [5 empty] | 4.1599 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037994 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2079.998 | [19 empty] | -1.2268 | [5 empty] | 4.16 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037996 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2079.998 | [19 empty] | -1.2268 | [5 empty] | 4.16 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037996 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2079.998 | [19 empty] | -1.2268 | [5 empty] | 4.16 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037996 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2080.067 | [19 empty] | -1.2268 | [5 empty] | 4.1601 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037997 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2080.067 | [19 empty] | -1.2268 | [5 empty] | 4.1601 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037997 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2080.067 | [19 empty] | -1.2268 | [5 empty] | 4.1601 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037997 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:03 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2268 | [5 empty] | 4.1603 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2268 | [5 empty] | 4.1603 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2267 | [5 empty] | 4.1603 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2268 | [5 empty] | 4.1603 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.205 | [19 empty] | -1.2268 | [5 empty] | 4.1604 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037999 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.205 | [19 empty] | -1.2268 | [5 empty] | 4.1604 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037999 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.205 | [19 empty] | -1.2268 | [5 empty] | 4.1604 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037999 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.205 | [19 empty] | -1.2268 | [5 empty] | 4.1604 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037999 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.2739 | [19 empty] | -1.2268 | [5 empty] | 4.1605 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038001 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.2739 | [19 empty] | -1.2268 | [5 empty] | 4.1605 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038001 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:04 | 20 | 20 |  | 2080.2739 | [19 empty] | -1.2268 | [5 empty] | 4.1605 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038001 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.2739 | [19 empty] | -1.2268 | [5 empty] | 4.1605 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038001 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.3429 | [19 empty] | -1.2268 | [5 empty] | 4.1607 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038002 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.3429 | [19 empty] | -1.2268 | [5 empty] | 4.1607 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038002 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.3429 | [19 empty] | -1.2267 | [5 empty] | 4.1607 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038002 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.3429 | [19 empty] | -1.2267 | [5 empty] | 4.1607 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038002 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.4119 | [19 empty] | -1.2268 | [5 empty] | 4.1608 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038003 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.4119 | [19 empty] | -1.2267 | [5 empty] | 4.1608 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038003 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.4119 | [19 empty] | -1.2268 | [5 empty] | 4.1608 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038003 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.4809 | [19 empty] | -1.2268 | [5 empty] | 4.161 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038005 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.4809 | [19 empty] | -1.2268 | [5 empty] | 4.161 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038005 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:05 | 20 | 20 |  | 2080.4809 | [19 empty] | -1.2267 | [5 empty] | 4.161 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038005 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.4809 | [19 empty] | -1.2268 | [5 empty] | 4.161 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038005 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.5498 | [19 empty] | -1.2268 | [5 empty] | 4.1611 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038006 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.5498 | [19 empty] | -1.2268 | [5 empty] | 4.1611 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038006 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.5498 | [19 empty] | -1.2268 | [5 empty] | 4.1611 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038006 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.5498 | [19 empty] | -1.2268 | [5 empty] | 4.1611 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038006 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.6188 | [19 empty] | -1.2268 | [5 empty] | 4.1612 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038007 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.6188 | [19 empty] | -1.2268 | [5 empty] | 4.1612 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038007 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.6188 | [19 empty] | -1.2268 | [5 empty] | 4.1612 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038007 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2269 | [5 empty] | 4.1614 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:06 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2268 | [5 empty] | 4.1614 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2268 | [5 empty] | 4.1614 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2268 | [5 empty] | 4.1614 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2268 | [5 empty] | 4.1614 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.7568 | [19 empty] | -1.2269 | [5 empty] | 4.1615 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003801 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.7568 | [19 empty] | -1.2268 | [5 empty] | 4.1615 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003801 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.8257 | [19 empty] | -1.2268 | [5 empty] | 4.1617 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038011 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.8257 | [19 empty] | -1.2267 | [5 empty] | 4.1617 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038011 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.8257 | [19 empty] | -1.2268 | [5 empty] | 4.1617 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038011 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.8257 | [19 empty] | -1.2268 | [5 empty] | 4.1617 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038011 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:07 | 20 | 20 |  | 2080.8947 | [19 empty] | -1.2268 | [5 empty] | 4.1618 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038012 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2080.8947 | [19 empty] | -1.2268 | [5 empty] | 4.1618 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038012 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2080.8947 | [19 empty] | -1.2267 | [5 empty] | 4.1618 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038012 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2080.9637 | [19 empty] | -1.2268 | [5 empty] | 4.1619 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038013 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2080.9637 | [19 empty] | -1.2268 | [5 empty] | 4.1619 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038013 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2080.9637 | [19 empty] | -1.2268 | [5 empty] | 4.1619 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038013 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2080.9637 | [19 empty] | -1.2268 | [5 empty] | 4.1619 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038013 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2267 | [5 empty] | 4.1621 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2267 | [5 empty] | 4.1621 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2268 | [5 empty] | 4.1621 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2268 | [5 empty] | 4.1621 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:08 | 20 | 20 |  | 2081.1016 | [19 empty] | -1.2268 | [5 empty] | 4.1622 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038016 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.1016 | [19 empty] | -1.2268 | [5 empty] | 4.1622 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038016 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.1016 | [19 empty] | -1.2268 | [5 empty] | 4.1622 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038016 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.1706 | [19 empty] | -1.2268 | [5 empty] | 4.1623 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038017 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.1706 | [19 empty] | -1.2268 | [5 empty] | 4.1623 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038017 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.1706 | [19 empty] | -1.2268 | [5 empty] | 4.1623 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038017 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.1706 | [19 empty] | -1.2268 | [5 empty] | 4.1623 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038017 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.2396 | [19 empty] | -1.2268 | [5 empty] | 4.1625 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038018 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.2396 | [19 empty] | -1.2267 | [5 empty] | 4.1625 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038018 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.2396 | [19 empty] | -1.2268 | [5 empty] | 4.1625 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038018 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:09 | 20 | 20 |  | 2081.2396 | [19 empty] | -1.2268 | [5 empty] | 4.1625 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038018 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3086 | [19 empty] | -1.2268 | [5 empty] | 4.1626 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003802 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3086 | [19 empty] | -1.2268 | [5 empty] | 4.1626 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003802 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3775 | [19 empty] | -1.2268 | [5 empty] | 4.1628 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038021 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3086 | [19 empty] | -1.2268 | [5 empty] | 4.1626 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003802 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3775 | [19 empty] | -1.2268 | [5 empty] | 4.1628 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038021 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3775 | [19 empty] | -1.2268 | [5 empty] | 4.1628 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038021 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.3775 | [19 empty] | -1.2268 | [5 empty] | 4.1628 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038021 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.4465 | [19 empty] | -1.2268 | [5 empty] | 4.1629 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038022 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.4465 | [19 empty] | -1.2268 | [5 empty] | 4.1629 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038022 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.4465 | [19 empty] | -1.2268 | [5 empty] | 4.1629 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038022 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:10 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2268 | [5 empty] | 4.163 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2268 | [5 empty] | 4.163 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2268 | [5 empty] | 4.163 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2268 | [5 empty] | 4.163 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2268 | [5 empty] | 4.1632 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2268 | [5 empty] | 4.1632 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2269 | [5 empty] | 4.1632 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2268 | [5 empty] | 4.1632 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2269 | [5 empty] | 4.1633 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2268 | [5 empty] | 4.1633 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:11 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2268 | [5 empty] | 4.1633 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2268 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2268 | [5 empty] | 4.1633 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2268 | [5 empty] | 4.1634 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2269 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2268 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2268 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2269 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2269 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7224 | [19 empty] | -1.2269 | [5 empty] | 4.1634 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038027 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:12 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2269 | [5 empty] | 4.1636 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2268 | [5 empty] | 4.1636 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2269 | [5 empty] | 4.1636 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2268 | [5 empty] | 4.1636 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.8604 | [19 empty] | -1.2268 | [5 empty] | 4.1637 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003803 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.8604 | [19 empty] | -1.2268 | [5 empty] | 4.1637 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003803 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.8604 | [19 empty] | -1.2268 | [5 empty] | 4.1637 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003803 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.9293 | [19 empty] | -1.2268 | [5 empty] | 4.1639 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038031 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.9293 | [19 empty] | -1.2268 | [5 empty] | 4.1639 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038031 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.9293 | [19 empty] | -1.2268 | [5 empty] | 4.1639 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038031 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:13 | 20 | 20 |  | 2081.9293 | [19 empty] | -1.2268 | [5 empty] | 4.1639 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038031 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2081.9983 | [19 empty] | -1.2268 | [5 empty] | 4.164 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038032 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2081.9983 | [19 empty] | -1.2268 | [5 empty] | 4.164 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038032 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2081.9983 | [19 empty] | -1.2268 | [5 empty] | 4.164 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038032 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2268 | [5 empty] | 4.1641 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2269 | [5 empty] | 4.1641 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2269 | [5 empty] | 4.1641 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2268 | [5 empty] | 4.1641 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.1363 | [19 empty] | -1.2267 | [5 empty] | 4.1643 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038035 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.1363 | [19 empty] | -1.2268 | [5 empty] | 4.1643 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038035 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.1363 | [19 empty] | -1.2268 | [5 empty] | 4.1643 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038035 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:14 | 20 | 20 |  | 2082.2052 | [19 empty] | -1.2269 | [5 empty] | 4.1644 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038036 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.2052 | [19 empty] | -1.2269 | [5 empty] | 4.1644 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038036 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.2052 | [19 empty] | -1.2269 | [5 empty] | 4.1644 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038036 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.2052 | [19 empty] | -1.2269 | [5 empty] | 4.1644 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038036 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.2742 | [19 empty] | -1.2268 | [5 empty] | 4.1645 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038037 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.2742 | [19 empty] | -1.2268 | [5 empty] | 4.1645 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038037 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.2742 | [19 empty] | -1.2269 | [5 empty] | 4.1645 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038037 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.3432 | [19 empty] | -1.2268 | [5 empty] | 4.1647 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038039 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.3432 | [19 empty] | -1.2269 | [5 empty] | 4.1647 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038039 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.3432 | [19 empty] | -1.2269 | [5 empty] | 4.1647 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038039 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:15 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2269 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2269 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2268 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2269 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2269 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4811 | [19 empty] | -1.2268 | [5 empty] | 4.165 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038041 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4811 | [19 empty] | -1.2268 | [5 empty] | 4.165 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038041 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.4811 | [19 empty] | -1.2268 | [5 empty] | 4.165 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038041 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.5501 | [19 empty] | -1.2269 | [5 empty] | 4.1651 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038042 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.5501 | [19 empty] | -1.2268 | [5 empty] | 4.1651 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038042 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.5501 | [19 empty] | -1.2268 | [5 empty] | 4.1651 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038042 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:16 | 20 | 20 |  | 2082.5501 | [19 empty] | -1.2268 | [5 empty] | 4.1651 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038042 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.6191 | [19 empty] | -1.2268 | [5 empty] | 4.1652 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038044 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.6191 | [19 empty] | -1.2268 | [5 empty] | 4.1652 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038044 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.6191 | [19 empty] | -1.2268 | [5 empty] | 4.1652 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038044 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.688 | [19 empty] | -1.2269 | [5 empty] | 4.1654 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038045 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.688 | [19 empty] | -1.2268 | [5 empty] | 4.1654 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038045 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.688 | [19 empty] | -1.2268 | [5 empty] | 4.1654 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038045 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.688 | [19 empty] | -1.2268 | [5 empty] | 4.1654 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038045 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.757 | [19 empty] | -1.2269 | [5 empty] | 4.1655 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038046 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.757 | [19 empty] | -1.2268 | [5 empty] | 4.1655 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038046 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:17 | 20 | 20 |  | 2082.757 | [19 empty] | -1.2269 | [5 empty] | 4.1655 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038046 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.757 | [19 empty] | -1.2269 | [5 empty] | 4.1655 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038046 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.826 | [19 empty] | -1.2268 | [5 empty] | 4.1657 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038048 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.826 | [19 empty] | -1.2268 | [5 empty] | 4.1657 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038048 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.826 | [19 empty] | -1.2269 | [5 empty] | 4.1657 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038048 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2268 | [5 empty] | 4.1658 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2268 | [5 empty] | 4.1658 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2268 | [5 empty] | 4.1658 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2268 | [5 empty] | 4.1658 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.9639 | [19 empty] | -1.2269 | [5 empty] | 4.1659 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003805 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:18 | 20 | 20 |  | 2082.9639 | [19 empty] | -1.2269 | [5 empty] | 4.1659 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003805 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2082.9639 | [19 empty] | -1.2269 | [5 empty] | 4.1659 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003805 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2082.9639 | [19 empty] | -1.2268 | [5 empty] | 4.1659 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003805 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2268 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2269 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2268 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2269 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2269 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.1019 | [19 empty] | -1.2269 | [5 empty] | 4.1662 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038053 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.1019 | [19 empty] | -1.2269 | [5 empty] | 4.1662 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038053 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.1019 | [19 empty] | -1.2269 | [5 empty] | 4.1662 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038053 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:19 | 20 | 20 |  | 2083.1019 | [19 empty] | -1.2268 | [5 empty] | 4.1662 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038053 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.1709 | [19 empty] | -1.2268 | [5 empty] | 4.1663 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038054 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.1709 | [19 empty] | -1.2268 | [5 empty] | 4.1663 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038054 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.1709 | [19 empty] | -1.2268 | [5 empty] | 4.1663 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038054 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2269 | [5 empty] | 4.1665 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2269 | [5 empty] | 4.1665 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2268 | [5 empty] | 4.1665 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2269 | [5 empty] | 4.1665 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2268 | [5 empty] | 4.1665 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.3088 | [19 empty] | -1.2269 | [5 empty] | 4.1666 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038056 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:20 | 20 | 20 |  | 2083.3088 | [19 empty] | -1.2268 | [5 empty] | 4.1666 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038056 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.3778 | [19 empty] | -1.2268 | [5 empty] | 4.1668 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038058 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.3778 | [19 empty] | -1.2268 | [5 empty] | 4.1668 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038058 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.3778 | [19 empty] | -1.2268 | [5 empty] | 4.1668 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038058 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.3778 | [19 empty] | -1.2268 | [5 empty] | 4.1668 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038058 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2268 | [5 empty] | 4.1669 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2268 | [5 empty] | 4.1669 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2268 | [5 empty] | 4.1669 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2268 | [5 empty] | 4.1669 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.5157 | [19 empty] | -1.2268 | [5 empty] | 4.167 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003806 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:21 | 20 | 20 |  | 2083.5157 | [19 empty] | -1.2268 | [5 empty] | 4.167 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003806 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.5157 | [19 empty] | -1.2269 | [5 empty] | 4.167 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003806 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.5847 | [19 empty] | -1.2268 | [5 empty] | 4.1672 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038061 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.5847 | [19 empty] | -1.2269 | [5 empty] | 4.1672 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038061 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.5847 | [19 empty] | -1.2268 | [5 empty] | 4.1672 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038061 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.5847 | [19 empty] | -1.2268 | [5 empty] | 4.1672 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038061 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.6537 | [19 empty] | -1.2269 | [5 empty] | 4.1673 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038063 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.6537 | [19 empty] | -1.2268 | [5 empty] | 4.1673 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038063 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.6537 | [19 empty] | -1.2269 | [5 empty] | 4.1673 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038063 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.6537 | [19 empty] | -1.2269 | [5 empty] | 4.1673 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038063 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.7227 | [19 empty] | -1.2269 | [5 empty] | 4.1674 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038064 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:22 | 20 | 20 |  | 2083.7227 | [19 empty] | -1.2269 | [5 empty] | 4.1674 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038064 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.7227 | [19 empty] | -1.2269 | [5 empty] | 4.1674 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038064 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.7227 | [19 empty] | -1.227 | [5 empty] | 4.1674 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038064 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.7916 | [19 empty] | -1.2269 | [5 empty] | 4.1676 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038065 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.7916 | [19 empty] | -1.2269 | [5 empty] | 4.1676 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038065 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.7916 | [19 empty] | -1.2268 | [5 empty] | 4.1676 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038065 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.8606 | [19 empty] | -1.2268 | [5 empty] | 4.1677 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038067 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.8606 | [19 empty] | -1.2269 | [5 empty] | 4.1677 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038067 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.8606 | [19 empty] | -1.2268 | [5 empty] | 4.1677 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038067 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2269 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:23 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2269 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2269 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2269 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2269 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2083.9986 | [19 empty] | -1.2269 | [5 empty] | 4.168 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038069 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2083.9986 | [19 empty] | -1.2268 | [5 empty] | 4.168 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038069 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2084.0675 | [19 empty] | -1.2269 | [5 empty] | 4.1681 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003807 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2084.0675 | [19 empty] | -1.2269 | [5 empty] | 4.1681 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003807 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2084.0675 | [19 empty] | -1.2269 | [5 empty] | 4.1681 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003807 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2084.0675 | [19 empty] | -1.2269 | [5 empty] | 4.1681 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003807 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2084.1365 | [19 empty] | -1.2269 | [5 empty] | 4.1683 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038072 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:24 | 20 | 20 |  | 2084.1365 | [19 empty] | -1.2269 | [5 empty] | 4.1683 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038072 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.1365 | [19 empty] | -1.2269 | [5 empty] | 4.1683 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038072 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.1365 | [19 empty] | -1.2269 | [5 empty] | 4.1683 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038072 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2055 | [19 empty] | -1.2268 | [5 empty] | 4.1684 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038073 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2055 | [19 empty] | -1.2269 | [5 empty] | 4.1684 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038073 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2055 | [19 empty] | -1.2269 | [5 empty] | 4.1684 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038073 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2745 | [19 empty] | -1.2269 | [5 empty] | 4.1685 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038074 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2745 | [19 empty] | -1.2269 | [5 empty] | 4.1685 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038074 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2745 | [19 empty] | -1.2269 | [5 empty] | 4.1685 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038074 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.2745 | [19 empty] | -1.2269 | [5 empty] | 4.1685 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038074 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:25 | 20 | 20 |  | 2084.3434 | [19 empty] | -1.2269 | [5 empty] | 4.1687 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038075 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.3434 | [19 empty] | -1.2269 | [5 empty] | 4.1687 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038075 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.3434 | [19 empty] | -1.2269 | [5 empty] | 4.1687 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038075 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.3434 | [19 empty] | -1.2269 | [5 empty] | 4.1687 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038075 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.3434 | [19 empty] | -1.2269 | [5 empty] | 4.1687 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038075 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.4124 | [19 empty] | -1.2268 | [5 empty] | 4.1688 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038077 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.4124 | [19 empty] | -1.2269 | [5 empty] | 4.1688 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038077 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.4814 | [19 empty] | -1.2269 | [5 empty] | 4.169 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038078 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.4814 | [19 empty] | -1.2269 | [5 empty] | 4.169 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038078 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.4814 | [19 empty] | -1.2269 | [5 empty] | 4.169 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038078 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:26 | 20 | 20 |  | 2084.4814 | [19 empty] | -1.2269 | [5 empty] | 4.169 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038078 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.5504 | [19 empty] | -1.2269 | [5 empty] | 4.1691 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038079 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.5504 | [19 empty] | -1.2269 | [5 empty] | 4.1691 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038079 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.5504 | [19 empty] | -1.2269 | [5 empty] | 4.1691 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038079 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.5504 | [19 empty] | -1.2269 | [5 empty] | 4.1691 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038079 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.6193 | [19 empty] | -1.2269 | [5 empty] | 4.1692 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003808 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.6193 | [19 empty] | -1.2269 | [5 empty] | 4.1692 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003808 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.6193 | [19 empty] | -1.2269 | [5 empty] | 4.1692 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003808 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.6883 | [19 empty] | -1.2269 | [5 empty] | 4.1694 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038082 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.6883 | [19 empty] | -1.2269 | [5 empty] | 4.1694 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038082 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:27 | 20 | 20 |  | 2084.6883 | [19 empty] | -1.2269 | [5 empty] | 4.1694 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038082 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.6883 | [19 empty] | -1.2269 | [5 empty] | 4.1694 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038082 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.7573 | [19 empty] | -1.2269 | [5 empty] | 4.1695 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038083 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.7573 | [19 empty] | -1.2269 | [5 empty] | 4.1695 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038083 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.7573 | [19 empty] | -1.2269 | [5 empty] | 4.1695 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038083 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8263 | [19 empty] | -1.2269 | [5 empty] | 4.1697 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038084 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8263 | [19 empty] | -1.2269 | [5 empty] | 4.1697 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038084 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8263 | [19 empty] | -1.2269 | [5 empty] | 4.1697 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038084 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8263 | [19 empty] | -1.2269 | [5 empty] | 4.1697 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038084 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2268 | [5 empty] | 4.1698 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2269 | [5 empty] | 4.1698 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:28 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2269 | [5 empty] | 4.1698 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2269 | [5 empty] | 4.1698 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2084.9642 | [19 empty] | -1.2269 | [5 empty] | 4.1699 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038087 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2084.9642 | [19 empty] | -1.2269 | [5 empty] | 4.1699 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038087 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.0332 | [19 empty] | -1.2269 | [5 empty] | 4.1701 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038088 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.0332 | [19 empty] | -1.2269 | [5 empty] | 4.1701 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038088 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.0332 | [19 empty] | -1.2269 | [5 empty] | 4.1701 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038088 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.0332 | [19 empty] | -1.2269 | [5 empty] | 4.1701 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038088 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.0332 | [19 empty] | -1.2269 | [5 empty] | 4.1701 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038088 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.1022 | [19 empty] | -1.2269 | [5 empty] | 4.1702 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038089 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:29 | 20 | 20 |  | 2085.1022 | [19 empty] | -1.2269 | [5 empty] | 4.1702 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038089 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.1022 | [19 empty] | -1.2269 | [5 empty] | 4.1702 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038089 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.1711 | [19 empty] | -1.2269 | [5 empty] | 4.1703 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038091 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.1711 | [19 empty] | -1.2269 | [5 empty] | 4.1703 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038091 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.1711 | [19 empty] | -1.2269 | [5 empty] | 4.1703 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038091 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.1711 | [19 empty] | -1.2269 | [5 empty] | 4.1703 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038091 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2269 | [5 empty] | 4.1705 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2269 | [5 empty] | 4.1705 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2269 | [5 empty] | 4.1705 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2269 | [5 empty] | 4.1705 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.3091 | [19 empty] | -1.2269 | [5 empty] | 4.1706 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038093 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:30 | 20 | 20 |  | 2085.3091 | [19 empty] | -1.2269 | [5 empty] | 4.1706 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038093 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.3781 | [19 empty] | -1.2269 | [5 empty] | 4.1708 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038094 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.3781 | [19 empty] | -1.2269 | [5 empty] | 4.1708 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038094 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.3781 | [19 empty] | -1.2269 | [5 empty] | 4.1708 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038094 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.3781 | [19 empty] | -1.2269 | [5 empty] | 4.1708 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038094 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.447 | [19 empty] | -1.2269 | [5 empty] | 4.1709 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038096 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.447 | [19 empty] | -1.2269 | [5 empty] | 4.1709 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038096 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.447 | [19 empty] | -1.2269 | [5 empty] | 4.1709 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038096 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.447 | [19 empty] | -1.2268 | [5 empty] | 4.1709 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038096 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.516 | [19 empty] | -1.2269 | [5 empty] | 4.171 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038097 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:31 | 20 | 20 |  | 2085.516 | [19 empty] | -1.2269 | [5 empty] | 4.171 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038097 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.516 | [19 empty] | -1.2269 | [5 empty] | 4.171 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038097 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.585 | [19 empty] | -1.2269 | [5 empty] | 4.1712 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038098 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.585 | [19 empty] | -1.2269 | [5 empty] | 4.1712 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038098 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.585 | [19 empty] | -1.2269 | [5 empty] | 4.1712 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038098 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.585 | [19 empty] | -1.2269 | [5 empty] | 4.1712 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038098 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.585 | [19 empty] | -1.2269 | [5 empty] | 4.1712 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038098 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.654 | [19 empty] | -1.2269 | [5 empty] | 4.1713 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038099 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.654 | [19 empty] | -1.2269 | [5 empty] | 4.1713 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038099 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.654 | [19 empty] | -1.2269 | [5 empty] | 4.1713 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038099 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:32 | 20 | 20 |  | 2085.654 | [19 empty] | -1.2269 | [5 empty] | 4.1713 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038099 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7229 | [19 empty] | -1.2269 | [5 empty] | 4.1714 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038101 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7229 | [19 empty] | -1.2269 | [5 empty] | 4.1714 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038101 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7229 | [19 empty] | -1.2269 | [5 empty] | 4.1714 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038101 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7919 | [19 empty] | -1.2269 | [5 empty] | 4.1716 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038102 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7919 | [19 empty] | -1.2269 | [5 empty] | 4.1716 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038102 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7919 | [19 empty] | -1.2269 | [5 empty] | 4.1716 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038102 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.7919 | [19 empty] | -1.2269 | [5 empty] | 4.1716 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038102 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.8609 | [19 empty] | -1.2269 | [5 empty] | 4.1717 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038103 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.8609 | [19 empty] | -1.2269 | [5 empty] | 4.1717 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038103 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.8609 | [19 empty] | -1.2269 | [5 empty] | 4.1717 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038103 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:33 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2269 | [5 empty] | 4.1719 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2268 | [5 empty] | 4.1719 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2269 | [5 empty] | 4.1719 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2269 | [5 empty] | 4.1719 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2085.9988 | [19 empty] | -1.2269 | [5 empty] | 4.172 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038106 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2085.9988 | [19 empty] | -1.2268 | [5 empty] | 4.172 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0038106 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2085.9988 | [19 empty] | -1.2269 | [5 empty] | 4.172 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038106 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2086.0678 | [19 empty] | -1.2268 | [5 empty] | 4.1721 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038107 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2086.0678 | [19 empty] | -1.2269 | [5 empty] | 4.1721 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038107 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2086.0678 | [19 empty] | -1.2269 | [5 empty] | 4.1721 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038107 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:34 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2269 | [5 empty] | 4.1723 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2269 | [5 empty] | 4.1723 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2269 | [5 empty] | 4.1723 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2269 | [5 empty] | 4.1723 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2269 | [5 empty] | 4.1723 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.2058 | [19 empty] | -1.2269 | [5 empty] | 4.1724 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003811 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.2058 | [19 empty] | -1.2269 | [5 empty] | 4.1724 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003811 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.2747 | [19 empty] | -1.2269 | [5 empty] | 4.1725 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038111 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.2747 | [19 empty] | -1.2269 | [5 empty] | 4.1725 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038111 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.2747 | [19 empty] | -1.2269 | [5 empty] | 4.1725 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038111 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:35 | 20 | 20 |  | 2086.2747 | [19 empty] | -1.2269 | [5 empty] | 4.1725 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038111 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.3437 | [19 empty] | -1.2269 | [5 empty] | 4.1727 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038112 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.3437 | [19 empty] | -1.2269 | [5 empty] | 4.1727 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038112 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.3437 | [19 empty] | -1.2269 | [5 empty] | 4.1727 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038112 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4127 | [19 empty] | -1.2269 | [5 empty] | 4.1728 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038113 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4127 | [19 empty] | -1.2268 | [5 empty] | 4.1728 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038113 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4127 | [19 empty] | -1.2269 | [5 empty] | 4.1728 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038113 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4127 | [19 empty] | -1.2269 | [5 empty] | 4.1728 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038113 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4816 | [19 empty] | -1.2269 | [5 empty] | 4.173 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038115 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4816 | [19 empty] | -1.2269 | [5 empty] | 4.173 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038115 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.4816 | [19 empty] | -1.2269 | [5 empty] | 4.173 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038115 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:36 | 20 | 20 |  | 2086.5506 | [19 empty] | -1.2269 | [5 empty] | 4.1731 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038116 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.5506 | [19 empty] | -1.2269 | [5 empty] | 4.1731 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038116 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.5506 | [19 empty] | -1.2269 | [5 empty] | 4.1731 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038116 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6196 | [19 empty] | -1.2269 | [5 empty] | 4.1732 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038117 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6196 | [19 empty] | -1.2269 | [5 empty] | 4.1732 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038117 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6196 | [19 empty] | -1.2269 | [5 empty] | 4.1732 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038117 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6196 | [19 empty] | -1.2269 | [5 empty] | 4.1732 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038117 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6196 | [19 empty] | -1.2269 | [5 empty] | 4.1732 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038117 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6886 | [19 empty] | -1.2269 | [5 empty] | 4.1734 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038118 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6886 | [19 empty] | -1.2269 | [5 empty] | 4.1734 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038118 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:37 | 20 | 20 |  | 2086.6886 | [19 empty] | -1.2269 | [5 empty] | 4.1734 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038118 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.7575 | [19 empty] | -1.2269 | [5 empty] | 4.1735 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003812 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.7575 | [19 empty] | -1.2269 | [5 empty] | 4.1735 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003812 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.7575 | [19 empty] | -1.2269 | [5 empty] | 4.1735 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003812 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8265 | [19 empty] | -1.2269 | [5 empty] | 4.1737 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038121 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8265 | [19 empty] | -1.2269 | [5 empty] | 4.1737 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038121 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8265 | [19 empty] | -1.2269 | [5 empty] | 4.1737 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038121 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8265 | [19 empty] | -1.2269 | [5 empty] | 4.1737 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038121 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8955 | [19 empty] | -1.2269 | [5 empty] | 4.1738 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038122 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8955 | [19 empty] | -1.2269 | [5 empty] | 4.1738 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038122 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:38 | 20 | 20 |  | 2086.8955 | [19 empty] | -1.2269 | [5 empty] | 4.1738 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038122 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2086.9645 | [19 empty] | -1.2269 | [5 empty] | 4.1739 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038124 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2086.9645 | [19 empty] | -1.2269 | [5 empty] | 4.1739 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038124 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2086.9645 | [19 empty] | -1.2269 | [5 empty] | 4.1739 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038124 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.0334 | [19 empty] | -1.2269 | [5 empty] | 4.1741 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038125 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.0334 | [19 empty] | -1.227 | [5 empty] | 4.1741 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038125 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.0334 | [19 empty] | -1.227 | [5 empty] | 4.1741 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038125 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.0334 | [19 empty] | -1.2269 | [5 empty] | 4.1741 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038125 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.0334 | [19 empty] | -1.227 | [5 empty] | 4.1741 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038125 | -1.4606 | -1.4414 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.1024 | [19 empty] | -1.2269 | [5 empty] | 4.1742 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038126 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.1024 | [19 empty] | -1.227 | [5 empty] | 4.1742 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038126 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:39 | 20 | 20 |  | 2087.1024 | [19 empty] | -1.227 | [5 empty] | 4.1742 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038126 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.1024 | [19 empty] | -1.2269 | [5 empty] | 4.1742 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038126 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.1714 | [19 empty] | -1.2269 | [5 empty] | 4.1743 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038127 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.1714 | [19 empty] | -1.2269 | [5 empty] | 4.1743 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038127 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.1714 | [19 empty] | -1.2269 | [5 empty] | 4.1743 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038127 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.2404 | [19 empty] | -1.2269 | [5 empty] | 4.1745 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038129 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.2404 | [19 empty] | -1.227 | [5 empty] | 4.1745 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038129 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.2404 | [19 empty] | -1.227 | [5 empty] | 4.1745 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038129 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.2404 | [19 empty] | -1.2269 | [5 empty] | 4.1745 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038129 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.3093 | [19 empty] | -1.227 | [5 empty] | 4.1746 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003813 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:40 | 20 | 20 |  | 2087.3093 | [19 empty] | -1.2269 | [5 empty] | 4.1746 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003813 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.3783 | [19 empty] | -1.2269 | [5 empty] | 4.1748 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038131 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.3783 | [19 empty] | -1.2269 | [5 empty] | 4.1748 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038131 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.3783 | [19 empty] | -1.2269 | [5 empty] | 4.1748 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038131 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.3783 | [19 empty] | -1.227 | [5 empty] | 4.1748 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038131 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.4473 | [19 empty] | -1.227 | [5 empty] | 4.1749 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038132 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.4473 | [19 empty] | -1.227 | [5 empty] | 4.1749 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038132 | -1.4607 | -1.4414 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.4473 | [19 empty] | -1.227 | [5 empty] | 4.1749 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038132 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.4473 | [19 empty] | -1.227 | [5 empty] | 4.1749 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038132 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.5163 | [19 empty] | -1.227 | [5 empty] | 4.175 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038134 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.5163 | [19 empty] | -1.2269 | [5 empty] | 4.175 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038134 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:41 | 20 | 20 |  | 2087.5163 | [19 empty] | -1.2269 | [5 empty] | 4.175 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038134 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.5852 | [19 empty] | -1.2269 | [5 empty] | 4.1752 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038135 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.5852 | [19 empty] | -1.2269 | [5 empty] | 4.1752 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038135 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.5852 | [19 empty] | -1.227 | [5 empty] | 4.1752 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038135 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.5852 | [19 empty] | -1.2269 | [5 empty] | 4.1752 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038135 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.5852 | [19 empty] | -1.2269 | [5 empty] | 4.1752 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038135 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.6542 | [19 empty] | -1.2269 | [5 empty] | 4.1753 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038136 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.6542 | [19 empty] | -1.2269 | [5 empty] | 4.1753 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038136 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.6542 | [19 empty] | -1.2269 | [5 empty] | 4.1753 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038136 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.7232 | [19 empty] | -1.2269 | [5 empty] | 4.1754 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038137 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:42 | 20 | 20 |  | 2087.7232 | [19 empty] | -1.2269 | [5 empty] | 4.1754 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038137 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.7232 | [19 empty] | -1.2269 | [5 empty] | 4.1754 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038137 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.7922 | [19 empty] | -1.2269 | [5 empty] | 4.1756 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038139 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.7922 | [19 empty] | -1.2269 | [5 empty] | 4.1756 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038139 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.7922 | [19 empty] | -1.2269 | [5 empty] | 4.1756 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038139 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.7922 | [19 empty] | -1.2269 | [5 empty] | 4.1756 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038139 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.8611 | [19 empty] | -1.2269 | [5 empty] | 4.1757 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003814 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.8611 | [19 empty] | -1.2269 | [5 empty] | 4.1757 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003814 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.9301 | [19 empty] | -1.2269 | [5 empty] | 4.1759 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038141 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.9301 | [19 empty] | -1.2269 | [5 empty] | 4.1759 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038141 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:43 | 20 | 20 |  | 2087.9301 | [19 empty] | -1.2269 | [5 empty] | 4.1759 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038141 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2087.9301 | [19 empty] | -1.2269 | [5 empty] | 4.1759 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038141 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2087.9991 | [19 empty] | -1.2269 | [5 empty] | 4.176 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038143 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2087.9991 | [19 empty] | -1.227 | [5 empty] | 4.176 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038143 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2087.9991 | [19 empty] | -1.2269 | [5 empty] | 4.176 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038143 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2269 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2269 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2269 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2269 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2269 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.137 | [19 empty] | -1.2269 | [5 empty] | 4.1763 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038145 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:44 | 20 | 20 |  | 2088.137 | [19 empty] | -1.227 | [5 empty] | 4.1763 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038145 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.137 | [19 empty] | -1.2269 | [5 empty] | 4.1763 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038145 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.206 | [19 empty] | -1.2269 | [5 empty] | 4.1764 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038146 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.206 | [19 empty] | -1.2269 | [5 empty] | 4.1764 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038146 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.275 | [19 empty] | -1.227 | [5 empty] | 4.1765 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038148 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.275 | [19 empty] | -1.2269 | [5 empty] | 4.1765 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038148 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.275 | [19 empty] | -1.2269 | [5 empty] | 4.1765 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038148 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.275 | [19 empty] | -1.2269 | [5 empty] | 4.1765 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038148 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.344 | [19 empty] | -1.2269 | [5 empty] | 4.1767 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038149 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.344 | [19 empty] | -1.2269 | [5 empty] | 4.1767 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038149 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:45 | 20 | 20 |  | 2088.344 | [19 empty] | -1.2269 | [5 empty] | 4.1767 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038149 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.344 | [19 empty] | -1.2269 | [5 empty] | 4.1767 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038149 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.4129 | [19 empty] | -1.227 | [5 empty] | 4.1768 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003815 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.4129 | [19 empty] | -1.2269 | [5 empty] | 4.1768 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003815 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.4129 | [19 empty] | -1.2269 | [5 empty] | 4.1768 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003815 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.4819 | [19 empty] | -1.2269 | [5 empty] | 4.177 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038151 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.4819 | [19 empty] | -1.2269 | [5 empty] | 4.177 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038151 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.4819 | [19 empty] | -1.227 | [5 empty] | 4.177 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038151 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.5509 | [19 empty] | -1.2269 | [5 empty] | 4.1771 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038153 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.5509 | [19 empty] | -1.2269 | [5 empty] | 4.1771 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038153 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:46 | 20 | 20 |  | 2088.5509 | [19 empty] | -1.2269 | [5 empty] | 4.1771 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038153 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.5509 | [19 empty] | -1.227 | [5 empty] | 4.1771 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038153 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6199 | [19 empty] | -1.2269 | [5 empty] | 4.1772 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038154 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6199 | [19 empty] | -1.227 | [5 empty] | 4.1772 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038154 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6199 | [19 empty] | -1.227 | [5 empty] | 4.1772 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038154 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6888 | [19 empty] | -1.2269 | [5 empty] | 4.1774 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038155 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6888 | [19 empty] | -1.2269 | [5 empty] | 4.1774 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038155 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6888 | [19 empty] | -1.2269 | [5 empty] | 4.1774 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038155 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.6888 | [19 empty] | -1.2269 | [5 empty] | 4.1774 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038155 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.7578 | [19 empty] | -1.2269 | [5 empty] | 4.1775 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038156 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:47 | 20 | 20 |  | 2088.7578 | [19 empty] | -1.2269 | [5 empty] | 4.1775 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038156 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.7578 | [19 empty] | -1.2269 | [5 empty] | 4.1775 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038156 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8268 | [19 empty] | -1.2269 | [5 empty] | 4.1777 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038158 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8268 | [19 empty] | -1.2269 | [5 empty] | 4.1777 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038158 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8268 | [19 empty] | -1.2269 | [5 empty] | 4.1777 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038158 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8958 | [19 empty] | -1.2269 | [5 empty] | 4.1778 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038159 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8958 | [19 empty] | -1.2269 | [5 empty] | 4.1778 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038159 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8958 | [19 empty] | -1.2269 | [5 empty] | 4.1778 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038159 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.8958 | [19 empty] | -1.227 | [5 empty] | 4.1778 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038159 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.9647 | [19 empty] | -1.227 | [5 empty] | 4.1779 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003816 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.9647 | [19 empty] | -1.227 | [5 empty] | 4.1779 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003816 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:48 | 20 | 20 |  | 2088.9647 | [19 empty] | -1.227 | [5 empty] | 4.1779 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003816 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.2269 | [5 empty] | 4.1781 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.227 | [5 empty] | 4.1781 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.2269 | [5 empty] | 4.1781 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.227 | [5 empty] | 4.1781 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.2269 | [5 empty] | 4.1781 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.1027 | [19 empty] | -1.2269 | [5 empty] | 4.1782 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038163 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.1027 | [19 empty] | -1.2269 | [5 empty] | 4.1782 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038163 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.1717 | [19 empty] | -1.2269 | [5 empty] | 4.1783 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038164 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.1717 | [19 empty] | -1.2269 | [5 empty] | 4.1783 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038164 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:49 | 20 | 20 |  | 2089.1717 | [19 empty] | -1.2269 | [5 empty] | 4.1783 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038164 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.1717 | [19 empty] | -1.2269 | [5 empty] | 4.1783 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038164 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.2406 | [19 empty] | -1.2269 | [5 empty] | 4.1785 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038165 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.2406 | [19 empty] | -1.2269 | [5 empty] | 4.1785 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038165 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.2406 | [19 empty] | -1.2269 | [5 empty] | 4.1785 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038165 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.2406 | [19 empty] | -1.227 | [5 empty] | 4.1785 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038165 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.3096 | [19 empty] | -1.227 | [5 empty] | 4.1786 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038167 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.3096 | [19 empty] | -1.227 | [5 empty] | 4.1786 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038167 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.3096 | [19 empty] | -1.227 | [5 empty] | 4.1786 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038167 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.3786 | [19 empty] | -1.227 | [5 empty] | 4.1788 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038168 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.3786 | [19 empty] | -1.2269 | [5 empty] | 4.1788 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038168 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:50 | 20 | 20 |  | 2089.3786 | [19 empty] | -1.227 | [5 empty] | 4.1788 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038168 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.4476 | [19 empty] | -1.227 | [5 empty] | 4.1789 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038169 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.4476 | [19 empty] | -1.227 | [5 empty] | 4.1789 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038169 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.4476 | [19 empty] | -1.2269 | [5 empty] | 4.1789 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038169 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.4476 | [19 empty] | -1.227 | [5 empty] | 4.1789 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038169 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.5165 | [19 empty] | -1.2269 | [5 empty] | 4.179 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003817 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.5165 | [19 empty] | -1.2269 | [5 empty] | 4.179 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003817 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.5165 | [19 empty] | -1.227 | [5 empty] | 4.179 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003817 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.5855 | [19 empty] | -1.2269 | [5 empty] | 4.1792 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038172 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.5855 | [19 empty] | -1.2269 | [5 empty] | 4.1792 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038172 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:51 | 20 | 20 |  | 2089.5855 | [19 empty] | -1.227 | [5 empty] | 4.1792 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038172 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.5855 | [19 empty] | -1.2269 | [5 empty] | 4.1792 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038172 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.6545 | [19 empty] | -1.227 | [5 empty] | 4.1793 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038173 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.6545 | [19 empty] | -1.2269 | [5 empty] | 4.1793 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038173 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.6545 | [19 empty] | -1.2269 | [5 empty] | 4.1793 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038173 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.7235 | [19 empty] | -1.2269 | [5 empty] | 4.1794 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038174 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.7235 | [19 empty] | -1.2269 | [5 empty] | 4.1794 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038174 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.7235 | [19 empty] | -1.2269 | [5 empty] | 4.1794 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038174 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.7924 | [19 empty] | -1.2269 | [5 empty] | 4.1796 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038175 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.7924 | [19 empty] | -1.2269 | [5 empty] | 4.1796 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038175 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:52 | 20 | 20 |  | 2089.7924 | [19 empty] | -1.2269 | [5 empty] | 4.1796 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038175 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-26 19:53 | 20 | 20 |  | 2089.7924 | [19 empty] | -1.227 | [5 empty] | 4.1796 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038175 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:43 | 20 | 20 |  | 2089.7235 | [19 empty] | -1.2274 | [5 empty] | 4.1794 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038174 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:43 | 20 | 20 |  | 2089.6545 | [19 empty] | -1.2275 | [5 empty] | 4.1793 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038173 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:43 | 20 | 20 |  | 2089.6545 | [19 empty] | -1.2274 | [5 empty] | 4.1793 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038173 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:43 | 20 | 20 |  | 2089.5855 | [19 empty] | -1.2274 | [5 empty] | 4.1792 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038172 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.5165 | [19 empty] | -1.2275 | [5 empty] | 4.179 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003817 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.4476 | [19 empty] | -1.2274 | [5 empty] | 4.1789 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038169 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.3786 | [19 empty] | -1.2274 | [5 empty] | 4.1788 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038168 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.3786 | [19 empty] | -1.2274 | [5 empty] | 4.1788 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038168 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.3096 | [19 empty] | -1.2274 | [5 empty] | 4.1786 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038167 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.2406 | [19 empty] | -1.2274 | [5 empty] | 4.1785 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038165 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.2406 | [19 empty] | -1.2274 | [5 empty] | 4.1785 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038165 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.1717 | [19 empty] | -1.2274 | [5 empty] | 4.1783 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038164 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.1027 | [19 empty] | -1.2274 | [5 empty] | 4.1782 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038163 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:44 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.2274 | [5 empty] | 4.1781 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2089.0337 | [19 empty] | -1.2274 | [5 empty] | 4.1781 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038162 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.9647 | [19 empty] | -1.2274 | [5 empty] | 4.1779 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003816 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.8958 | [19 empty] | -1.2274 | [5 empty] | 4.1778 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038159 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.8268 | [19 empty] | -1.2274 | [5 empty] | 4.1777 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038158 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.7578 | [19 empty] | -1.2274 | [5 empty] | 4.1775 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038156 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.6888 | [19 empty] | -1.2275 | [5 empty] | 4.1774 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038155 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.6888 | [19 empty] | -1.2274 | [5 empty] | 4.1774 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038155 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.6199 | [19 empty] | -1.2275 | [5 empty] | 4.1772 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038154 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.5509 | [19 empty] | -1.2274 | [5 empty] | 4.1771 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038153 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.4819 | [19 empty] | -1.2274 | [5 empty] | 4.177 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038151 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:45 | 20 | 20 |  | 2088.4129 | [19 empty] | -1.2274 | [5 empty] | 4.1768 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003815 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.4129 | [19 empty] | -1.2274 | [5 empty] | 4.1768 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003815 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.344 | [19 empty] | -1.2274 | [5 empty] | 4.1767 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038149 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.275 | [19 empty] | -1.2274 | [5 empty] | 4.1765 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038148 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.206 | [19 empty] | -1.2274 | [5 empty] | 4.1764 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038146 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.137 | [19 empty] | -1.2275 | [5 empty] | 4.1763 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038145 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2274 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2088.0681 | [19 empty] | -1.2274 | [5 empty] | 4.1761 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038144 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2087.9301 | [19 empty] | -1.2275 | [5 empty] | 4.1759 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038141 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2087.9301 | [19 empty] | -1.2275 | [5 empty] | 4.1759 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038141 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:46 | 20 | 20 |  | 2087.8611 | [19 empty] | -1.2275 | [5 empty] | 4.1757 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003814 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.7922 | [19 empty] | -1.2274 | [5 empty] | 4.1756 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038139 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.7232 | [19 empty] | -1.2274 | [5 empty] | 4.1754 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038137 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.7232 | [19 empty] | -1.2275 | [5 empty] | 4.1754 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038137 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.6542 | [19 empty] | -1.2274 | [5 empty] | 4.1753 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038136 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.5852 | [19 empty] | -1.2275 | [5 empty] | 4.1752 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038135 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.5163 | [19 empty] | -1.2274 | [5 empty] | 4.175 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038134 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.4473 | [19 empty] | -1.2275 | [5 empty] | 4.1749 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038132 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.4473 | [19 empty] | -1.2274 | [5 empty] | 4.1749 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038132 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.3783 | [19 empty] | -1.2274 | [5 empty] | 4.1748 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038131 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:47 | 20 | 20 |  | 2087.3093 | [19 empty] | -1.2274 | [5 empty] | 4.1746 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003813 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2087.2404 | [19 empty] | -1.2275 | [5 empty] | 4.1745 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038129 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2087.2404 | [19 empty] | -1.2274 | [5 empty] | 4.1745 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038129 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2087.1714 | [19 empty] | -1.2274 | [5 empty] | 4.1743 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038127 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2087.1024 | [19 empty] | -1.2274 | [5 empty] | 4.1742 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038126 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2087.0334 | [19 empty] | -1.2274 | [5 empty] | 4.1741 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038125 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2086.9645 | [19 empty] | -1.2274 | [5 empty] | 4.1739 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038124 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2086.9645 | [19 empty] | -1.2274 | [5 empty] | 4.1739 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038124 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2086.8955 | [19 empty] | -1.2274 | [5 empty] | 4.1738 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038122 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2086.8265 | [19 empty] | -1.2274 | [5 empty] | 4.1737 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038121 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2086.8265 | [19 empty] | -1.2274 | [5 empty] | 4.1737 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038121 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:48 | 20 | 20 |  | 2086.7575 | [19 empty] | -1.2274 | [5 empty] | 4.1735 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003812 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.6886 | [19 empty] | -1.2274 | [5 empty] | 4.1734 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038118 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.6196 | [19 empty] | -1.2274 | [5 empty] | 4.1732 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038117 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.5506 | [19 empty] | -1.2275 | [5 empty] | 4.1731 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038116 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.4816 | [19 empty] | -1.2274 | [5 empty] | 4.173 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038115 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.4816 | [19 empty] | -1.2274 | [5 empty] | 4.173 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038115 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.4127 | [19 empty] | -1.2274 | [5 empty] | 4.1728 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038113 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.3437 | [19 empty] | -1.2274 | [5 empty] | 4.1727 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038112 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.2747 | [19 empty] | -1.2274 | [5 empty] | 4.1725 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038111 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.2058 | [19 empty] | -1.2274 | [5 empty] | 4.1724 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003811 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:49 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2274 | [5 empty] | 4.1723 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2086.1368 | [19 empty] | -1.2274 | [5 empty] | 4.1723 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038108 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2086.0678 | [19 empty] | -1.2274 | [5 empty] | 4.1721 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038107 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2086.0678 | [19 empty] | -1.2274 | [5 empty] | 4.1721 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038107 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2274 | [5 empty] | 4.1719 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.9299 | [19 empty] | -1.2274 | [5 empty] | 4.1719 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038105 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.8609 | [19 empty] | -1.2274 | [5 empty] | 4.1717 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038103 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.7919 | [19 empty] | -1.2274 | [5 empty] | 4.1716 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038102 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.7229 | [19 empty] | -1.2275 | [5 empty] | 4.1714 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038101 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.654 | [19 empty] | -1.2274 | [5 empty] | 4.1713 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038099 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:50 | 20 | 20 |  | 2085.654 | [19 empty] | -1.2274 | [5 empty] | 4.1713 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038099 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.585 | [19 empty] | -1.2274 | [5 empty] | 4.1712 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038098 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.516 | [19 empty] | -1.2274 | [5 empty] | 4.171 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038097 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.447 | [19 empty] | -1.2274 | [5 empty] | 4.1709 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038096 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.3781 | [19 empty] | -1.2274 | [5 empty] | 4.1708 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038094 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.3781 | [19 empty] | -1.2274 | [5 empty] | 4.1708 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038094 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.3091 | [19 empty] | -1.2274 | [5 empty] | 4.1706 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038093 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2274 | [5 empty] | 4.1705 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.2401 | [19 empty] | -1.2274 | [5 empty] | 4.1705 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038092 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.1711 | [19 empty] | -1.2274 | [5 empty] | 4.1703 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038091 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.1022 | [19 empty] | -1.2274 | [5 empty] | 4.1702 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038089 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:51 | 20 | 20 |  | 2085.0332 | [19 empty] | -1.2274 | [5 empty] | 4.1701 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038088 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.9642 | [19 empty] | -1.2274 | [5 empty] | 4.1699 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038087 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2274 | [5 empty] | 4.1698 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.8952 | [19 empty] | -1.2274 | [5 empty] | 4.1698 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038086 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.8263 | [19 empty] | -1.2274 | [5 empty] | 4.1697 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038084 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.7573 | [19 empty] | -1.2274 | [5 empty] | 4.1695 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038083 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.6883 | [19 empty] | -1.2274 | [5 empty] | 4.1694 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038082 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.6193 | [19 empty] | -1.2274 | [5 empty] | 4.1692 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003808 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.5504 | [19 empty] | -1.2274 | [5 empty] | 4.1691 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038079 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.5504 | [19 empty] | -1.2274 | [5 empty] | 4.1691 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038079 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:52 | 20 | 20 |  | 2084.4814 | [19 empty] | -1.2274 | [5 empty] | 4.169 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038078 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.4124 | [19 empty] | -1.2274 | [5 empty] | 4.1688 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038077 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.3434 | [19 empty] | -1.2274 | [5 empty] | 4.1687 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038075 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.2745 | [19 empty] | -1.2274 | [5 empty] | 4.1685 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038074 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.2055 | [19 empty] | -1.2274 | [5 empty] | 4.1684 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038073 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.1365 | [19 empty] | -1.2274 | [5 empty] | 4.1683 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038072 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.1365 | [19 empty] | -1.2274 | [5 empty] | 4.1683 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038072 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2084.0675 | [19 empty] | -1.2274 | [5 empty] | 4.1681 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003807 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2083.9986 | [19 empty] | -1.2274 | [5 empty] | 4.168 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038069 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2274 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:53 | 20 | 20 |  | 2083.9296 | [19 empty] | -1.2274 | [5 empty] | 4.1679 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038068 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.8606 | [19 empty] | -1.2274 | [5 empty] | 4.1677 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038067 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.7916 | [19 empty] | -1.2275 | [5 empty] | 4.1676 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038065 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.7227 | [19 empty] | -1.2274 | [5 empty] | 4.1674 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038064 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.6537 | [19 empty] | -1.2274 | [5 empty] | 4.1673 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038063 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.5847 | [19 empty] | -1.2274 | [5 empty] | 4.1672 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038061 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.5157 | [19 empty] | -1.2274 | [5 empty] | 4.167 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003806 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2274 | [5 empty] | 4.1669 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.4468 | [19 empty] | -1.2274 | [5 empty] | 4.1669 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038059 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.3778 | [19 empty] | -1.2274 | [5 empty] | 4.1668 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038058 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.3088 | [19 empty] | -1.2275 | [5 empty] | 4.1666 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038056 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:54 | 20 | 20 |  | 2083.2398 | [19 empty] | -1.2274 | [5 empty] | 4.1665 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038055 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2083.1709 | [19 empty] | -1.2274 | [5 empty] | 4.1663 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038054 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2083.1709 | [19 empty] | -1.2274 | [5 empty] | 4.1663 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038054 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2274 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2083.0329 | [19 empty] | -1.2274 | [5 empty] | 4.1661 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038051 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2082.9639 | [19 empty] | -1.2274 | [5 empty] | 4.1659 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003805 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2274 | [5 empty] | 4.1658 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2082.895 | [19 empty] | -1.2274 | [5 empty] | 4.1658 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038049 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2082.826 | [19 empty] | -1.2274 | [5 empty] | 4.1657 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038048 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2082.826 | [19 empty] | -1.2274 | [5 empty] | 4.1657 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038048 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:55 | 20 | 20 |  | 2082.757 | [19 empty] | -1.2274 | [5 empty] | 4.1655 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038046 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.688 | [19 empty] | -1.2274 | [5 empty] | 4.1654 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038045 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.6191 | [19 empty] | -1.2274 | [5 empty] | 4.1652 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038044 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.5501 | [19 empty] | -1.2274 | [5 empty] | 4.1651 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038042 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.4811 | [19 empty] | -1.2273 | [5 empty] | 4.165 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038041 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2274 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.4121 | [19 empty] | -1.2274 | [5 empty] | 4.1648 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003804 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.3432 | [19 empty] | -1.2274 | [5 empty] | 4.1647 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038039 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.2742 | [19 empty] | -1.2274 | [5 empty] | 4.1645 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038037 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.2052 | [19 empty] | -1.2274 | [5 empty] | 4.1644 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038036 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.1363 | [19 empty] | -1.2274 | [5 empty] | 4.1643 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038035 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:56 | 20 | 20 |  | 2082.0673 | [19 empty] | -1.2274 | [5 empty] | 4.1641 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038034 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.9983 | [19 empty] | -1.2274 | [5 empty] | 4.164 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038032 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.9983 | [19 empty] | -1.2274 | [5 empty] | 4.164 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038032 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.9293 | [19 empty] | -1.2274 | [5 empty] | 4.1639 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038031 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.8604 | [19 empty] | -1.2274 | [5 empty] | 4.1637 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003803 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2274 | [5 empty] | 4.1636 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.7914 | [19 empty] | -1.2274 | [5 empty] | 4.1636 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038029 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2275 | [5 empty] | 4.1633 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.6534 | [19 empty] | -1.2274 | [5 empty] | 4.1633 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038026 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2274 | [5 empty] | 4.1632 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:57 | 20 | 20 |  | 2081.5845 | [19 empty] | -1.2274 | [5 empty] | 4.1632 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038025 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2274 | [5 empty] | 4.163 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.5155 | [19 empty] | -1.2274 | [5 empty] | 4.163 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038024 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.4465 | [19 empty] | -1.2274 | [5 empty] | 4.1629 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038022 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.3775 | [19 empty] | -1.2275 | [5 empty] | 4.1628 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0038021 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.3775 | [19 empty] | -1.2274 | [5 empty] | 4.1628 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038021 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.3086 | [19 empty] | -1.2275 | [5 empty] | 4.1626 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003802 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.2396 | [19 empty] | -1.2274 | [5 empty] | 4.1625 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038018 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.1706 | [19 empty] | -1.2274 | [5 empty] | 4.1623 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038017 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.1016 | [19 empty] | -1.2274 | [5 empty] | 4.1622 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038016 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:58 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2274 | [5 empty] | 4.1621 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2081.0327 | [19 empty] | -1.2274 | [5 empty] | 4.1621 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038015 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.9637 | [19 empty] | -1.2274 | [5 empty] | 4.1619 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038013 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.8947 | [19 empty] | -1.2274 | [5 empty] | 4.1618 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038012 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.8257 | [19 empty] | -1.2274 | [5 empty] | 4.1617 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038011 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.7568 | [19 empty] | -1.2274 | [5 empty] | 4.1615 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003801 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.7568 | [19 empty] | -1.2274 | [5 empty] | 4.1615 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003801 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.6878 | [19 empty] | -1.2274 | [5 empty] | 4.1614 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038008 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.6188 | [19 empty] | -1.2274 | [5 empty] | 4.1612 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038007 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.5498 | [19 empty] | -1.2274 | [5 empty] | 4.1611 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038006 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.4809 | [19 empty] | -1.2274 | [5 empty] | 4.161 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038005 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 07:59 | 20 | 20 |  | 2080.4119 | [19 empty] | -1.2274 | [5 empty] | 4.1608 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038003 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.4119 | [19 empty] | -1.2274 | [5 empty] | 4.1608 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038003 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.3429 | [19 empty] | -1.2274 | [5 empty] | 4.1607 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038002 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.2739 | [19 empty] | -1.2274 | [5 empty] | 4.1605 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0038001 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.205 | [19 empty] | -1.2274 | [5 empty] | 4.1604 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037999 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2274 | [5 empty] | 4.1603 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.136 | [19 empty] | -1.2274 | [5 empty] | 4.1603 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037998 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2080.067 | [19 empty] | -1.2274 | [5 empty] | 4.1601 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037997 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2079.998 | [19 empty] | -1.2274 | [5 empty] | 4.16 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037996 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2079.9291 | [19 empty] | -1.2274 | [5 empty] | 4.1599 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037994 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:00 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2274 | [5 empty] | 4.1597 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.8601 | [19 empty] | -1.2274 | [5 empty] | 4.1597 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037993 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.7911 | [19 empty] | -1.2274 | [5 empty] | 4.1596 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037992 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.7221 | [19 empty] | -1.2274 | [5 empty] | 4.1594 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037991 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.6532 | [19 empty] | -1.2274 | [5 empty] | 4.1593 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037989 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.5842 | [19 empty] | -1.2274 | [5 empty] | 4.1592 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037988 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.5152 | [19 empty] | -1.2274 | [5 empty] | 4.159 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037987 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.4462 | [19 empty] | -1.2274 | [5 empty] | 4.1589 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037986 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2274 | [5 empty] | 4.1588 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.3773 | [19 empty] | -1.2274 | [5 empty] | 4.1588 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037984 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:01 | 20 | 20 |  | 2079.3083 | [19 empty] | -1.2274 | [5 empty] | 4.1586 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037983 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2079.2393 | [19 empty] | -1.2274 | [5 empty] | 4.1585 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037982 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2079.1703 | [19 empty] | -1.2274 | [5 empty] | 4.1583 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003798 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2079.1014 | [19 empty] | -1.2274 | [5 empty] | 4.1582 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037979 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2079.0324 | [19 empty] | -1.2274 | [5 empty] | 4.1581 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037978 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.9634 | [19 empty] | -1.2274 | [5 empty] | 4.1579 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037977 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.9634 | [19 empty] | -1.2274 | [5 empty] | 4.1579 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037977 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.8944 | [19 empty] | -1.2274 | [5 empty] | 4.1578 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037975 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.8255 | [19 empty] | -1.2274 | [5 empty] | 4.1577 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037974 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.7565 | [19 empty] | -1.2274 | [5 empty] | 4.1575 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037973 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.6875 | [19 empty] | -1.2274 | [5 empty] | 4.1574 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037972 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:02 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2274 | [5 empty] | 4.1572 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.6185 | [19 empty] | -1.2274 | [5 empty] | 4.1572 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003797 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.5496 | [19 empty] | -1.2274 | [5 empty] | 4.1571 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037969 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.4806 | [19 empty] | -1.2274 | [5 empty] | 4.157 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037968 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.4116 | [19 empty] | -1.2274 | [5 empty] | 4.1568 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037967 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.3426 | [19 empty] | -1.2274 | [5 empty] | 4.1567 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037965 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2273 | [5 empty] | 4.1565 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.2737 | [19 empty] | -1.2273 | [5 empty] | 4.1565 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037964 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2274 | [5 empty] | 4.1563 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.1357 | [19 empty] | -1.2274 | [5 empty] | 4.1563 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037961 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:03 | 20 | 20 |  | 2078.0667 | [19 empty] | -1.2274 | [5 empty] | 4.1561 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003796 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.9978 | [19 empty] | -1.2274 | [5 empty] | 4.156 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037959 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2274 | [5 empty] | 4.1559 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.9288 | [19 empty] | -1.2274 | [5 empty] | 4.1559 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037958 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.8598 | [19 empty] | -1.2274 | [5 empty] | 4.1557 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037956 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.7909 | [19 empty] | -1.2274 | [5 empty] | 4.1556 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037955 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.7219 | [19 empty] | -1.2274 | [5 empty] | 4.1554 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037954 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.6529 | [19 empty] | -1.2274 | [5 empty] | 4.1553 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037953 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.6529 | [19 empty] | -1.2274 | [5 empty] | 4.1553 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037953 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.5839 | [19 empty] | -1.2274 | [5 empty] | 4.1552 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037951 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:04 | 20 | 20 |  | 2077.515 | [19 empty] | -1.2274 | [5 empty] | 4.155 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003795 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.446 | [19 empty] | -1.2274 | [5 empty] | 4.1549 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037949 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.377 | [19 empty] | -1.2274 | [5 empty] | 4.1548 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037948 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.308 | [19 empty] | -1.2274 | [5 empty] | 4.1546 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037946 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.308 | [19 empty] | -1.2275 | [5 empty] | 4.1546 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037946 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.2391 | [19 empty] | -1.2275 | [5 empty] | 4.1545 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037945 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.1701 | [19 empty] | -1.2275 | [5 empty] | 4.1543 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037944 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.1011 | [19 empty] | -1.2274 | [5 empty] | 4.1542 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037942 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.1011 | [19 empty] | -1.2275 | [5 empty] | 4.1542 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037942 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2077.0321 | [19 empty] | -1.2274 | [5 empty] | 4.1541 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037941 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2076.9632 | [19 empty] | -1.2275 | [5 empty] | 4.1539 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003794 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:05 | 20 | 20 |  | 2076.8942 | [19 empty] | -1.2274 | [5 empty] | 4.1538 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037939 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.8252 | [19 empty] | -1.2274 | [5 empty] | 4.1537 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037937 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.8252 | [19 empty] | -1.2274 | [5 empty] | 4.1537 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037937 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.7562 | [19 empty] | -1.2275 | [5 empty] | 4.1535 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037936 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.6873 | [19 empty] | -1.2274 | [5 empty] | 4.1534 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037935 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.6183 | [19 empty] | -1.2274 | [5 empty] | 4.1532 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037934 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.6183 | [19 empty] | -1.2274 | [5 empty] | 4.1532 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037934 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.5493 | [19 empty] | -1.2274 | [5 empty] | 4.1531 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037932 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.4803 | [19 empty] | -1.2275 | [5 empty] | 4.153 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037931 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.4114 | [19 empty] | -1.2275 | [5 empty] | 4.1528 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003793 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:06 | 20 | 20 |  | 2076.3424 | [19 empty] | -1.2274 | [5 empty] | 4.1527 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037929 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2076.2734 | [19 empty] | -1.2275 | [5 empty] | 4.1525 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037927 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2076.2734 | [19 empty] | -1.2275 | [5 empty] | 4.1525 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037927 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2076.2044 | [19 empty] | -1.2274 | [5 empty] | 4.1524 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037926 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2076.1355 | [19 empty] | -1.2275 | [5 empty] | 4.1523 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037925 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2274 | [5 empty] | 4.1521 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2076.0665 | [19 empty] | -1.2274 | [5 empty] | 4.1521 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037923 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2075.9975 | [19 empty] | -1.2274 | [5 empty] | 4.152 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037922 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2075.9285 | [19 empty] | -1.2274 | [5 empty] | 4.1519 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037921 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2075.8596 | [19 empty] | -1.2274 | [5 empty] | 4.1517 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003792 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:07 | 20 | 20 |  | 2075.7906 | [19 empty] | -1.2274 | [5 empty] | 4.1516 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037918 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.7216 | [19 empty] | -1.2274 | [5 empty] | 4.1514 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037917 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2274 | [5 empty] | 4.1513 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.6526 | [19 empty] | -1.2274 | [5 empty] | 4.1513 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037916 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.5837 | [19 empty] | -1.2275 | [5 empty] | 4.1512 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037915 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.5147 | [19 empty] | -1.2275 | [5 empty] | 4.151 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037913 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.4457 | [19 empty] | -1.2274 | [5 empty] | 4.1509 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037912 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.3767 | [19 empty] | -1.2274 | [5 empty] | 4.1508 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037911 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.3078 | [19 empty] | -1.2274 | [5 empty] | 4.1506 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003791 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.2388 | [19 empty] | -1.2274 | [5 empty] | 4.1505 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037908 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2274 | [5 empty] | 4.1503 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:08 | 20 | 20 |  | 2075.1698 | [19 empty] | -1.2274 | [5 empty] | 4.1503 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037907 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2075.1008 | [19 empty] | -1.2274 | [5 empty] | 4.1502 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037906 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2274 | [5 empty] | 4.1499 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.9629 | [19 empty] | -1.2274 | [5 empty] | 4.1499 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037903 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.8939 | [19 empty] | -1.2274 | [5 empty] | 4.1498 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037902 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.8249 | [19 empty] | -1.2274 | [5 empty] | 4.1496 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037901 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.756 | [19 empty] | -1.2275 | [5 empty] | 4.1495 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037899 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2274 | [5 empty] | 4.1494 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.687 | [19 empty] | -1.2274 | [5 empty] | 4.1494 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037898 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.618 | [19 empty] | -1.2274 | [5 empty] | 4.1492 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037897 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:09 | 20 | 20 |  | 2074.549 | [19 empty] | -1.2274 | [5 empty] | 4.1491 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037896 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2275 | [5 empty] | 4.149 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.4801 | [19 empty] | -1.2275 | [5 empty] | 4.149 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037894 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.4111 | [19 empty] | -1.2274 | [5 empty] | 4.1488 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037893 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.3421 | [19 empty] | -1.2275 | [5 empty] | 4.1487 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037892 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.2731 | [19 empty] | -1.2274 | [5 empty] | 4.1485 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037891 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.2042 | [19 empty] | -1.2274 | [5 empty] | 4.1484 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037889 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2274 | [5 empty] | 4.1483 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.1352 | [19 empty] | -1.2274 | [5 empty] | 4.1483 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037888 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2074.0662 | [19 empty] | -1.2275 | [5 empty] | 4.1481 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037887 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2073.9972 | [19 empty] | -1.2274 | [5 empty] | 4.148 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037886 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:10 | 20 | 20 |  | 2073.9283 | [19 empty] | -1.2274 | [5 empty] | 4.1479 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037884 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.8593 | [19 empty] | -1.2274 | [5 empty] | 4.1477 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037883 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2274 | [5 empty] | 4.1476 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.7903 | [19 empty] | -1.2275 | [5 empty] | 4.1476 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037882 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.7213 | [19 empty] | -1.2275 | [5 empty] | 4.1474 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003788 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2275 | [5 empty] | 4.1473 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.6524 | [19 empty] | -1.2274 | [5 empty] | 4.1473 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037879 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.5834 | [19 empty] | -1.2274 | [5 empty] | 4.1472 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037878 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.5144 | [19 empty] | -1.2274 | [5 empty] | 4.147 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037877 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.4455 | [19 empty] | -1.2274 | [5 empty] | 4.1469 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037875 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:11 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2274 | [5 empty] | 4.1468 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.3765 | [19 empty] | -1.2274 | [5 empty] | 4.1468 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037874 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.3075 | [19 empty] | -1.2275 | [5 empty] | 4.1466 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037873 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.2385 | [19 empty] | -1.2274 | [5 empty] | 4.1465 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037872 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.1696 | [19 empty] | -1.2274 | [5 empty] | 4.1463 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003787 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.1006 | [19 empty] | -1.2274 | [5 empty] | 4.1462 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037869 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2275 | [5 empty] | 4.1461 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2073.0316 | [19 empty] | -1.2274 | [5 empty] | 4.1461 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037868 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2072.9626 | [19 empty] | -1.2274 | [5 empty] | 4.1459 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037867 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2072.8937 | [19 empty] | -1.2274 | [5 empty] | 4.1458 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037865 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:12 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2274 | [5 empty] | 4.1456 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.8247 | [19 empty] | -1.2274 | [5 empty] | 4.1456 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037864 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.7557 | [19 empty] | -1.2274 | [5 empty] | 4.1455 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037863 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.6867 | [19 empty] | -1.2275 | [5 empty] | 4.1454 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037861 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.6178 | [19 empty] | -1.2275 | [5 empty] | 4.1452 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003786 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.5488 | [19 empty] | -1.2275 | [5 empty] | 4.1451 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037859 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2274 | [5 empty] | 4.145 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.4798 | [19 empty] | -1.2274 | [5 empty] | 4.145 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037858 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.4108 | [19 empty] | -1.2274 | [5 empty] | 4.1448 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037856 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.3419 | [19 empty] | -1.2275 | [5 empty] | 4.1447 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.0037855 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2274 | [5 empty] | 4.1445 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:13 | 20 | 20 |  | 2072.2729 | [19 empty] | -1.2274 | [5 empty] | 4.1445 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037854 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2072.2039 | [19 empty] | -1.2275 | [5 empty] | 4.1444 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037853 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2072.1349 | [19 empty] | -1.2274 | [5 empty] | 4.1443 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037851 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2072.066 | [19 empty] | -1.2274 | [5 empty] | 4.1441 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003785 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.997 | [19 empty] | -1.2274 | [5 empty] | 4.144 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037849 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2274 | [5 empty] | 4.1439 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.928 | [19 empty] | -1.2274 | [5 empty] | 4.1439 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037848 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.859 | [19 empty] | -1.2274 | [5 empty] | 4.1437 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037846 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.7901 | [19 empty] | -1.2274 | [5 empty] | 4.1436 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037845 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.7211 | [19 empty] | -1.2274 | [5 empty] | 4.1434 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037844 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:14 | 20 | 20 |  | 2071.6521 | [19 empty] | -1.2274 | [5 empty] | 4.1433 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037842 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2274 | [5 empty] | 4.1432 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.5831 | [19 empty] | -1.2274 | [5 empty] | 4.1432 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037841 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.5142 | [19 empty] | -1.2275 | [5 empty] | 4.143 | -1.2267 | 0.90444 | 17.6664 | 0.23773 | -0.003784 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.4452 | [19 empty] | -1.2275 | [5 empty] | 4.1429 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037839 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.3762 | [19 empty] | -1.2274 | [5 empty] | 4.1428 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037837 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.3072 | [19 empty] | -1.2274 | [5 empty] | 4.1426 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037836 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2274 | [5 empty] | 4.1425 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.2383 | [19 empty] | -1.2274 | [5 empty] | 4.1425 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037835 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.1693 | [19 empty] | -1.2274 | [5 empty] | 4.1423 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037834 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:15 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2274 | [5 empty] | 4.1422 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2071.1003 | [19 empty] | -1.2274 | [5 empty] | 4.1422 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037832 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2274 | [5 empty] | 4.1419 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.9624 | [19 empty] | -1.2274 | [5 empty] | 4.1419 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003783 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.8934 | [19 empty] | -1.2274 | [5 empty] | 4.1418 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037829 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.8244 | [19 empty] | -1.2274 | [5 empty] | 4.1416 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037827 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.7554 | [19 empty] | -1.2274 | [5 empty] | 4.1415 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037826 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.7554 | [19 empty] | -1.2274 | [5 empty] | 4.1415 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037826 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.6175 | [19 empty] | -1.2274 | [5 empty] | 4.1412 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037823 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.6175 | [19 empty] | -1.2274 | [5 empty] | 4.1412 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037823 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.5485 | [19 empty] | -1.2274 | [5 empty] | 4.1411 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037822 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:16 | 20 | 20 |  | 2070.4795 | [19 empty] | -1.2274 | [5 empty] | 4.141 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037821 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.4795 | [19 empty] | -1.2274 | [5 empty] | 4.141 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037821 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.4106 | [19 empty] | -1.2274 | [5 empty] | 4.1408 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003782 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.3416 | [19 empty] | -1.2274 | [5 empty] | 4.1407 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037818 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.2726 | [19 empty] | -1.2274 | [5 empty] | 4.1405 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037817 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.2036 | [19 empty] | -1.2274 | [5 empty] | 4.1404 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037816 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.1347 | [19 empty] | -1.2274 | [5 empty] | 4.1403 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037815 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.1347 | [19 empty] | -1.2274 | [5 empty] | 4.1403 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037815 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2070.0657 | [19 empty] | -1.2274 | [5 empty] | 4.1401 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037813 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2069.9967 | [19 empty] | -1.2274 | [5 empty] | 4.14 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037812 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:17 | 20 | 20 |  | 2069.9277 | [19 empty] | -1.2274 | [5 empty] | 4.1399 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037811 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.9277 | [19 empty] | -1.2274 | [5 empty] | 4.1399 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037811 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.9277 | [19 empty] | -1.2274 | [5 empty] | 4.1399 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037811 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.8588 | [19 empty] | -1.2274 | [5 empty] | 4.1397 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003781 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.7898 | [19 empty] | -1.2274 | [5 empty] | 4.1396 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037808 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.7208 | [19 empty] | -1.2274 | [5 empty] | 4.1394 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037807 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.7208 | [19 empty] | -1.2274 | [5 empty] | 4.1394 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037807 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.6518 | [19 empty] | -1.2274 | [5 empty] | 4.1393 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037806 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.5829 | [19 empty] | -1.2274 | [5 empty] | 4.1392 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037804 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.5139 | [19 empty] | -1.2274 | [5 empty] | 4.139 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037803 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:18 | 20 | 20 |  | 2069.5139 | [19 empty] | -1.2274 | [5 empty] | 4.139 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037803 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.4449 | [19 empty] | -1.2274 | [5 empty] | 4.1389 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037802 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.3759 | [19 empty] | -1.2274 | [5 empty] | 4.1388 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037801 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.307 | [19 empty] | -1.2274 | [5 empty] | 4.1386 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037799 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.238 | [19 empty] | -1.2274 | [5 empty] | 4.1385 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037798 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.238 | [19 empty] | -1.2274 | [5 empty] | 4.1385 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037798 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.169 | [19 empty] | -1.2274 | [5 empty] | 4.1383 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037797 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.1001 | [19 empty] | -1.2274 | [5 empty] | 4.1382 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037796 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.0311 | [19 empty] | -1.2274 | [5 empty] | 4.1381 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037794 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2069.0311 | [19 empty] | -1.2274 | [5 empty] | 4.1381 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037794 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2068.9621 | [19 empty] | -1.2274 | [5 empty] | 4.1379 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037793 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:19 | 20 | 20 |  | 2068.8931 | [19 empty] | -1.2274 | [5 empty] | 4.1378 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037792 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.8242 | [19 empty] | -1.2274 | [5 empty] | 4.1376 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037791 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.7552 | [19 empty] | -1.2274 | [5 empty] | 4.1375 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037789 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.7552 | [19 empty] | -1.2274 | [5 empty] | 4.1375 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037789 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.6862 | [19 empty] | -1.2274 | [5 empty] | 4.1374 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037788 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.6172 | [19 empty] | -1.2275 | [5 empty] | 4.1372 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037787 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.6172 | [19 empty] | -1.2274 | [5 empty] | 4.1372 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037787 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.5483 | [19 empty] | -1.2274 | [5 empty] | 4.1371 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037785 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.5483 | [19 empty] | -1.2274 | [5 empty] | 4.1371 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037785 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.4793 | [19 empty] | -1.2274 | [5 empty] | 4.137 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037784 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:20 | 20 | 20 |  | 2068.4103 | [19 empty] | -1.2274 | [5 empty] | 4.1368 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037783 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2068.3413 | [19 empty] | -1.2274 | [5 empty] | 4.1367 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037782 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2068.2724 | [19 empty] | -1.2274 | [5 empty] | 4.1365 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003778 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2068.2034 | [19 empty] | -1.2274 | [5 empty] | 4.1364 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037779 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2068.2034 | [19 empty] | -1.2274 | [5 empty] | 4.1364 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037779 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2068.1344 | [19 empty] | -1.2274 | [5 empty] | 4.1363 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037778 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2068.0654 | [19 empty] | -1.2274 | [5 empty] | 4.1361 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037777 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2067.9965 | [19 empty] | -1.2274 | [5 empty] | 4.136 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037775 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2067.9275 | [19 empty] | -1.2274 | [5 empty] | 4.1359 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037774 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2067.8585 | [19 empty] | -1.2274 | [5 empty] | 4.1357 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037773 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:21 | 20 | 20 |  | 2067.8585 | [19 empty] | -1.2274 | [5 empty] | 4.1357 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037773 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.7895 | [19 empty] | -1.2274 | [5 empty] | 4.1356 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037772 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.7206 | [19 empty] | -1.2274 | [5 empty] | 4.1354 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003777 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.7206 | [19 empty] | -1.2274 | [5 empty] | 4.1354 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003777 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.6516 | [19 empty] | -1.2274 | [5 empty] | 4.1353 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037769 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.6516 | [19 empty] | -1.2274 | [5 empty] | 4.1353 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037769 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.5826 | [19 empty] | -1.2274 | [5 empty] | 4.1352 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037768 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.5136 | [19 empty] | -1.2274 | [5 empty] | 4.135 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037766 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.5136 | [19 empty] | -1.2274 | [5 empty] | 4.135 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037766 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.4447 | [19 empty] | -1.2273 | [5 empty] | 4.1349 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037765 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.3757 | [19 empty] | -1.2274 | [5 empty] | 4.1348 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037764 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:22 | 20 | 20 |  | 2067.3757 | [19 empty] | -1.2274 | [5 empty] | 4.1348 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037764 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.3067 | [19 empty] | -1.2274 | [5 empty] | 4.1346 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037763 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.3067 | [19 empty] | -1.2274 | [5 empty] | 4.1346 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037763 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.2377 | [19 empty] | -1.2274 | [5 empty] | 4.1345 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037761 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.1688 | [19 empty] | -1.2274 | [5 empty] | 4.1343 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003776 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.1688 | [19 empty] | -1.2274 | [5 empty] | 4.1343 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003776 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.0998 | [19 empty] | -1.2274 | [5 empty] | 4.1342 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037759 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.0998 | [19 empty] | -1.2273 | [5 empty] | 4.1342 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037759 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2067.0308 | [19 empty] | -1.2274 | [5 empty] | 4.1341 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037758 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2066.9618 | [19 empty] | -1.2274 | [5 empty] | 4.1339 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037756 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:23 | 20 | 20 |  | 2066.8929 | [19 empty] | -1.2274 | [5 empty] | 4.1338 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037755 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.8239 | [19 empty] | -1.2274 | [5 empty] | 4.1336 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037754 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.7549 | [19 empty] | -1.2274 | [5 empty] | 4.1335 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037753 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.7549 | [19 empty] | -1.2274 | [5 empty] | 4.1335 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037753 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.6859 | [19 empty] | -1.2274 | [5 empty] | 4.1334 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037751 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.617 | [19 empty] | -1.2274 | [5 empty] | 4.1332 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003775 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.617 | [19 empty] | -1.2274 | [5 empty] | 4.1332 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003775 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.548 | [19 empty] | -1.2274 | [5 empty] | 4.1331 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037749 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.479 | [19 empty] | -1.2274 | [5 empty] | 4.133 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037748 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.41 | [19 empty] | -1.2274 | [5 empty] | 4.1328 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037746 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.41 | [19 empty] | -1.2274 | [5 empty] | 4.1328 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037746 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:24 | 20 | 20 |  | 2066.3411 | [19 empty] | -1.2274 | [5 empty] | 4.1327 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037745 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2066.2721 | [19 empty] | -1.2274 | [5 empty] | 4.1325 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037744 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2066.2031 | [19 empty] | -1.2274 | [5 empty] | 4.1324 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037742 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2066.2031 | [19 empty] | -1.2274 | [5 empty] | 4.1324 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037742 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2066.1341 | [19 empty] | -1.2274 | [5 empty] | 4.1323 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037741 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2066.0652 | [19 empty] | -1.2274 | [5 empty] | 4.1321 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003774 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2066.0652 | [19 empty] | -1.2274 | [5 empty] | 4.1321 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003774 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2065.9962 | [19 empty] | -1.2274 | [5 empty] | 4.132 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037739 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2065.9272 | [19 empty] | -1.2274 | [5 empty] | 4.1319 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037737 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2065.8582 | [19 empty] | -1.2274 | [5 empty] | 4.1317 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037736 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:25 | 20 | 20 |  | 2065.8582 | [19 empty] | -1.2274 | [5 empty] | 4.1317 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037736 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.7893 | [19 empty] | -1.2274 | [5 empty] | 4.1316 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037735 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.7203 | [19 empty] | -1.2274 | [5 empty] | 4.1314 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037734 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.6513 | [19 empty] | -1.2274 | [5 empty] | 4.1313 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037732 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.6513 | [19 empty] | -1.2274 | [5 empty] | 4.1313 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037732 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.5823 | [19 empty] | -1.2274 | [5 empty] | 4.1312 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037731 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.5134 | [19 empty] | -1.2274 | [5 empty] | 4.131 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003773 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.5134 | [19 empty] | -1.2274 | [5 empty] | 4.131 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003773 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.4444 | [19 empty] | -1.2274 | [5 empty] | 4.1309 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037729 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.3754 | [19 empty] | -1.2274 | [5 empty] | 4.1308 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037727 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:26 | 20 | 20 |  | 2065.3064 | [19 empty] | -1.2274 | [5 empty] | 4.1306 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037726 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.3064 | [19 empty] | -1.2273 | [5 empty] | 4.1306 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037726 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.2375 | [19 empty] | -1.2274 | [5 empty] | 4.1305 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037725 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.2375 | [19 empty] | -1.2274 | [5 empty] | 4.1305 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037725 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.1685 | [19 empty] | -1.2273 | [5 empty] | 4.1303 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037723 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.1685 | [19 empty] | -1.2274 | [5 empty] | 4.1303 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037723 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.0995 | [19 empty] | -1.2274 | [5 empty] | 4.1302 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037722 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.0305 | [19 empty] | -1.2274 | [5 empty] | 4.1301 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037721 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2065.0305 | [19 empty] | -1.2274 | [5 empty] | 4.1301 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037721 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2064.9616 | [19 empty] | -1.2274 | [5 empty] | 4.1299 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003772 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:27 | 20 | 20 |  | 2064.8926 | [19 empty] | -1.2274 | [5 empty] | 4.1298 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037718 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.8926 | [19 empty] | -1.2274 | [5 empty] | 4.1298 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037718 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.8236 | [19 empty] | -1.2274 | [5 empty] | 4.1296 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037717 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.7547 | [19 empty] | -1.2274 | [5 empty] | 4.1295 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037716 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.7547 | [19 empty] | -1.2274 | [5 empty] | 4.1295 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037716 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.6857 | [19 empty] | -1.2274 | [5 empty] | 4.1294 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037715 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.6167 | [19 empty] | -1.2274 | [5 empty] | 4.1292 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037713 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.6167 | [19 empty] | -1.2274 | [5 empty] | 4.1292 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037713 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.5477 | [19 empty] | -1.2274 | [5 empty] | 4.1291 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037712 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.5477 | [19 empty] | -1.2274 | [5 empty] | 4.1291 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037712 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.4788 | [19 empty] | -1.2273 | [5 empty] | 4.129 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037711 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:28 | 20 | 20 |  | 2064.4788 | [19 empty] | -1.2274 | [5 empty] | 4.129 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037711 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.4098 | [19 empty] | -1.2274 | [5 empty] | 4.1288 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003771 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.3408 | [19 empty] | -1.2274 | [5 empty] | 4.1287 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037708 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.3408 | [19 empty] | -1.2274 | [5 empty] | 4.1287 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037708 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.2718 | [19 empty] | -1.2274 | [5 empty] | 4.1285 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037707 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.2029 | [19 empty] | -1.2274 | [5 empty] | 4.1284 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037706 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.2029 | [19 empty] | -1.2274 | [5 empty] | 4.1284 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037706 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.1339 | [19 empty] | -1.2274 | [5 empty] | 4.1283 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037704 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.0649 | [19 empty] | -1.2274 | [5 empty] | 4.1281 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037703 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2064.0649 | [19 empty] | -1.2273 | [5 empty] | 4.1281 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037703 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:29 | 20 | 20 |  | 2063.9959 | [19 empty] | -1.2274 | [5 empty] | 4.128 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037702 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.9959 | [19 empty] | -1.2274 | [5 empty] | 4.128 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037702 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.927 | [19 empty] | -1.2274 | [5 empty] | 4.1279 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037701 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.858 | [19 empty] | -1.2274 | [5 empty] | 4.1277 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037699 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.789 | [19 empty] | -1.2274 | [5 empty] | 4.1276 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037698 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.789 | [19 empty] | -1.2274 | [5 empty] | 4.1276 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037698 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.789 | [19 empty] | -1.2274 | [5 empty] | 4.1276 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037698 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.72 | [19 empty] | -1.2274 | [5 empty] | 4.1274 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037697 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.6511 | [19 empty] | -1.2274 | [5 empty] | 4.1273 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037696 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.6511 | [19 empty] | -1.2274 | [5 empty] | 4.1273 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037696 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.5821 | [19 empty] | -1.2274 | [5 empty] | 4.1272 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037694 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:30 | 20 | 20 |  | 2063.5821 | [19 empty] | -1.2274 | [5 empty] | 4.1272 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037694 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.5131 | [19 empty] | -1.2274 | [5 empty] | 4.127 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037693 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.4441 | [19 empty] | -1.2274 | [5 empty] | 4.1269 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037692 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.4441 | [19 empty] | -1.2273 | [5 empty] | 4.1269 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037692 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.3752 | [19 empty] | -1.2274 | [5 empty] | 4.1268 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037691 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.3752 | [19 empty] | -1.2274 | [5 empty] | 4.1268 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037691 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.3062 | [19 empty] | -1.2274 | [5 empty] | 4.1266 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037689 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.3062 | [19 empty] | -1.2274 | [5 empty] | 4.1266 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037689 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.2372 | [19 empty] | -1.2274 | [5 empty] | 4.1265 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037688 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.1682 | [19 empty] | -1.2274 | [5 empty] | 4.1263 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037687 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:31 | 20 | 20 |  | 2063.1682 | [19 empty] | -1.2274 | [5 empty] | 4.1263 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037687 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2063.0993 | [19 empty] | -1.2274 | [5 empty] | 4.1262 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037685 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2063.0303 | [19 empty] | -1.2274 | [5 empty] | 4.1261 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037684 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2063.0303 | [19 empty] | -1.2274 | [5 empty] | 4.1261 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037684 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.9613 | [19 empty] | -1.2274 | [5 empty] | 4.1259 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037683 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.9613 | [19 empty] | -1.2274 | [5 empty] | 4.1259 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037683 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.8923 | [19 empty] | -1.2274 | [5 empty] | 4.1258 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037682 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.8923 | [19 empty] | -1.2274 | [5 empty] | 4.1258 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037682 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.8234 | [19 empty] | -1.2274 | [5 empty] | 4.1256 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003768 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.8234 | [19 empty] | -1.2274 | [5 empty] | 4.1256 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003768 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:32 | 20 | 20 |  | 2062.7544 | [19 empty] | -1.2274 | [5 empty] | 4.1255 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037679 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.7544 | [19 empty] | -1.2274 | [5 empty] | 4.1255 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037679 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.6854 | [19 empty] | -1.2274 | [5 empty] | 4.1254 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037678 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.6854 | [19 empty] | -1.2274 | [5 empty] | 4.1254 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037678 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.6164 | [19 empty] | -1.2274 | [5 empty] | 4.1252 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037677 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.6164 | [19 empty] | -1.2274 | [5 empty] | 4.1252 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037677 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.5475 | [19 empty] | -1.2274 | [5 empty] | 4.1251 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037675 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.4785 | [19 empty] | -1.2274 | [5 empty] | 4.125 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037674 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.4785 | [19 empty] | -1.2274 | [5 empty] | 4.125 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037674 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.4095 | [19 empty] | -1.2274 | [5 empty] | 4.1248 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037673 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.3405 | [19 empty] | -1.2274 | [5 empty] | 4.1247 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037672 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:33 | 20 | 20 |  | 2062.3405 | [19 empty] | -1.2274 | [5 empty] | 4.1247 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037672 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.2716 | [19 empty] | -1.2273 | [5 empty] | 4.1245 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003767 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.2716 | [19 empty] | -1.2274 | [5 empty] | 4.1245 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003767 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.2026 | [19 empty] | -1.2274 | [5 empty] | 4.1244 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037669 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.1336 | [19 empty] | -1.2274 | [5 empty] | 4.1243 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037668 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.0646 | [19 empty] | -1.2274 | [5 empty] | 4.1241 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037666 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.0646 | [19 empty] | -1.2274 | [5 empty] | 4.1241 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037666 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2062.0646 | [19 empty] | -1.2274 | [5 empty] | 4.1241 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037666 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2061.9957 | [19 empty] | -1.2274 | [5 empty] | 4.124 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037665 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2061.9267 | [19 empty] | -1.2273 | [5 empty] | 4.1239 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037664 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:34 | 20 | 20 |  | 2061.9267 | [19 empty] | -1.2274 | [5 empty] | 4.1239 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037664 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.8577 | [19 empty] | -1.2274 | [5 empty] | 4.1237 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037663 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.8577 | [19 empty] | -1.2274 | [5 empty] | 4.1237 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037663 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.7887 | [19 empty] | -1.2274 | [5 empty] | 4.1236 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037661 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.7887 | [19 empty] | -1.2273 | [5 empty] | 4.1236 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037661 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.7198 | [19 empty] | -1.2274 | [5 empty] | 4.1234 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003766 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.6508 | [19 empty] | -1.2274 | [5 empty] | 4.1233 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037659 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.6508 | [19 empty] | -1.2273 | [5 empty] | 4.1233 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037659 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.5818 | [19 empty] | -1.2273 | [5 empty] | 4.1232 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037658 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.5128 | [19 empty] | -1.2273 | [5 empty] | 4.123 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037656 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:35 | 20 | 20 |  | 2061.5128 | [19 empty] | -1.2274 | [5 empty] | 4.123 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037656 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.4439 | [19 empty] | -1.2273 | [5 empty] | 4.1229 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037655 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.4439 | [19 empty] | -1.2274 | [5 empty] | 4.1229 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037655 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.3749 | [19 empty] | -1.2274 | [5 empty] | 4.1227 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037654 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.3749 | [19 empty] | -1.2274 | [5 empty] | 4.1227 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037654 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.3059 | [19 empty] | -1.2274 | [5 empty] | 4.1226 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037653 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.2369 | [19 empty] | -1.2273 | [5 empty] | 4.1225 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037651 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.2369 | [19 empty] | -1.2273 | [5 empty] | 4.1225 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037651 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.168 | [19 empty] | -1.2274 | [5 empty] | 4.1223 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003765 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.168 | [19 empty] | -1.2274 | [5 empty] | 4.1223 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003765 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.099 | [19 empty] | -1.2274 | [5 empty] | 4.1222 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037649 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:36 | 20 | 20 |  | 2061.099 | [19 empty] | -1.2274 | [5 empty] | 4.1222 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037649 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2061.03 | [19 empty] | -1.2274 | [5 empty] | 4.1221 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037647 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.961 | [19 empty] | -1.2274 | [5 empty] | 4.1219 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037646 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.961 | [19 empty] | -1.2274 | [5 empty] | 4.1219 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037646 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.961 | [19 empty] | -1.2274 | [5 empty] | 4.1219 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037646 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.8921 | [19 empty] | -1.2274 | [5 empty] | 4.1218 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037645 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.8231 | [19 empty] | -1.2273 | [5 empty] | 4.1216 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037644 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.8231 | [19 empty] | -1.2274 | [5 empty] | 4.1216 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037644 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.7541 | [19 empty] | -1.2274 | [5 empty] | 4.1215 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037642 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.7541 | [19 empty] | -1.2274 | [5 empty] | 4.1215 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037642 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:37 | 20 | 20 |  | 2060.6851 | [19 empty] | -1.2274 | [5 empty] | 4.1214 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037641 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.6162 | [19 empty] | -1.2274 | [5 empty] | 4.1212 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003764 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.6162 | [19 empty] | -1.2274 | [5 empty] | 4.1212 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003764 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.5472 | [19 empty] | -1.2274 | [5 empty] | 4.1211 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037639 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.5472 | [19 empty] | -1.2274 | [5 empty] | 4.1211 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037639 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.4782 | [19 empty] | -1.2274 | [5 empty] | 4.121 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037637 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.4093 | [19 empty] | -1.2274 | [5 empty] | 4.1208 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037636 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.4093 | [19 empty] | -1.2274 | [5 empty] | 4.1208 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037636 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.3403 | [19 empty] | -1.2274 | [5 empty] | 4.1207 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037635 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.3403 | [19 empty] | -1.2274 | [5 empty] | 4.1207 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037635 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.2713 | [19 empty] | -1.2274 | [5 empty] | 4.1205 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037634 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:38 | 20 | 20 |  | 2060.2713 | [19 empty] | -1.2274 | [5 empty] | 4.1205 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037634 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2060.2023 | [19 empty] | -1.2274 | [5 empty] | 4.1204 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037632 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2060.2023 | [19 empty] | -1.2274 | [5 empty] | 4.1204 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037632 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2060.1334 | [19 empty] | -1.2273 | [5 empty] | 4.1203 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037631 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2060.1334 | [19 empty] | -1.2273 | [5 empty] | 4.1203 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037631 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2060.0644 | [19 empty] | -1.2273 | [5 empty] | 4.1201 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003763 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2060.0644 | [19 empty] | -1.2273 | [5 empty] | 4.1201 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003763 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2059.9954 | [19 empty] | -1.2274 | [5 empty] | 4.12 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037628 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2059.9954 | [19 empty] | -1.2274 | [5 empty] | 4.12 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037628 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2059.9264 | [19 empty] | -1.2274 | [5 empty] | 4.1199 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037627 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:39 | 20 | 20 |  | 2059.8575 | [19 empty] | -1.2274 | [5 empty] | 4.1197 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037626 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.8575 | [19 empty] | -1.2274 | [5 empty] | 4.1197 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037626 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.7885 | [19 empty] | -1.2274 | [5 empty] | 4.1196 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037625 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.7195 | [19 empty] | -1.2274 | [5 empty] | 4.1194 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037623 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.7195 | [19 empty] | -1.2274 | [5 empty] | 4.1194 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037623 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.6505 | [19 empty] | -1.2274 | [5 empty] | 4.1193 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037622 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.6505 | [19 empty] | -1.2274 | [5 empty] | 4.1193 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037622 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.5816 | [19 empty] | -1.2274 | [5 empty] | 4.1192 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037621 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.5816 | [19 empty] | -1.2274 | [5 empty] | 4.1192 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037621 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.5126 | [19 empty] | -1.2274 | [5 empty] | 4.119 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003762 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:40 | 20 | 20 |  | 2059.5126 | [19 empty] | -1.2274 | [5 empty] | 4.119 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003762 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.4436 | [19 empty] | -1.2274 | [5 empty] | 4.1189 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037618 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.4436 | [19 empty] | -1.2274 | [5 empty] | 4.1189 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037618 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.3746 | [19 empty] | -1.2274 | [5 empty] | 4.1187 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037617 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.3746 | [19 empty] | -1.2274 | [5 empty] | 4.1187 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037617 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.3057 | [19 empty] | -1.2274 | [5 empty] | 4.1186 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037616 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.3057 | [19 empty] | -1.2274 | [5 empty] | 4.1186 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037616 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.2367 | [19 empty] | -1.2274 | [5 empty] | 4.1185 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037615 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.2367 | [19 empty] | -1.2273 | [5 empty] | 4.1185 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037615 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.1677 | [19 empty] | -1.2273 | [5 empty] | 4.1183 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037613 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:41 | 20 | 20 |  | 2059.1677 | [19 empty] | -1.2273 | [5 empty] | 4.1183 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037613 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2059.0987 | [19 empty] | -1.2273 | [5 empty] | 4.1182 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037612 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2059.0987 | [19 empty] | -1.2274 | [5 empty] | 4.1182 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037612 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2059.0987 | [19 empty] | -1.2274 | [5 empty] | 4.1182 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037612 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2059.0298 | [19 empty] | -1.2274 | [5 empty] | 4.1181 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037611 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2059.0298 | [19 empty] | -1.2274 | [5 empty] | 4.1181 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037611 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2058.9608 | [19 empty] | -1.2273 | [5 empty] | 4.1179 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003761 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2058.9608 | [19 empty] | -1.2274 | [5 empty] | 4.1179 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003761 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2058.8918 | [19 empty] | -1.2274 | [5 empty] | 4.1178 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037608 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2058.8228 | [19 empty] | -1.2274 | [5 empty] | 4.1176 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037607 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2058.8228 | [19 empty] | -1.2274 | [5 empty] | 4.1176 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037607 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:42 | 20 | 20 |  | 2058.7539 | [19 empty] | -1.2274 | [5 empty] | 4.1175 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037606 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.7539 | [19 empty] | -1.2274 | [5 empty] | 4.1175 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037606 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.6849 | [19 empty] | -1.2274 | [5 empty] | 4.1174 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037604 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.6159 | [19 empty] | -1.2274 | [5 empty] | 4.1172 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037603 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.6159 | [19 empty] | -1.2274 | [5 empty] | 4.1172 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037603 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.5469 | [19 empty] | -1.2274 | [5 empty] | 4.1171 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037602 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.5469 | [19 empty] | -1.2274 | [5 empty] | 4.1171 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037602 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.5469 | [19 empty] | -1.2273 | [5 empty] | 4.1171 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037602 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.478 | [19 empty] | -1.2273 | [5 empty] | 4.117 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037601 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.478 | [19 empty] | -1.2273 | [5 empty] | 4.117 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037601 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:43 | 20 | 20 |  | 2058.409 | [19 empty] | -1.2274 | [5 empty] | 4.1168 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037599 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.409 | [19 empty] | -1.2274 | [5 empty] | 4.1168 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037599 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.34 | [19 empty] | -1.2274 | [5 empty] | 4.1167 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037598 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.34 | [19 empty] | -1.2273 | [5 empty] | 4.1167 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037598 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.34 | [19 empty] | -1.2274 | [5 empty] | 4.1167 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037598 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.271 | [19 empty] | -1.2273 | [5 empty] | 4.1165 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037597 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.271 | [19 empty] | -1.2273 | [5 empty] | 4.1165 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037597 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.2021 | [19 empty] | -1.2273 | [5 empty] | 4.1164 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037596 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.2021 | [19 empty] | -1.2273 | [5 empty] | 4.1164 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037596 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.2021 | [19 empty] | -1.2273 | [5 empty] | 4.1164 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037596 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.1331 | [19 empty] | -1.2273 | [5 empty] | 4.1163 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037594 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:44 | 20 | 20 |  | 2058.0641 | [19 empty] | -1.2274 | [5 empty] | 4.1161 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037593 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2058.0641 | [19 empty] | -1.2274 | [5 empty] | 4.1161 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037593 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.9951 | [19 empty] | -1.2273 | [5 empty] | 4.116 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037592 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.9951 | [19 empty] | -1.2273 | [5 empty] | 4.116 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037592 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.9262 | [19 empty] | -1.2273 | [5 empty] | 4.1159 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037591 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.8572 | [19 empty] | -1.2274 | [5 empty] | 4.1157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037589 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.8572 | [19 empty] | -1.2273 | [5 empty] | 4.1157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037589 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.8572 | [19 empty] | -1.2273 | [5 empty] | 4.1157 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037589 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.7882 | [19 empty] | -1.2273 | [5 empty] | 4.1156 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037588 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.7882 | [19 empty] | -1.2274 | [5 empty] | 4.1156 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037588 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:45 | 20 | 20 |  | 2057.7192 | [19 empty] | -1.2273 | [5 empty] | 4.1154 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037587 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.7192 | [19 empty] | -1.2273 | [5 empty] | 4.1154 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037587 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.6503 | [19 empty] | -1.2273 | [5 empty] | 4.1153 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037585 | -1.4604 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.6503 | [19 empty] | -1.2273 | [5 empty] | 4.1153 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037585 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.5813 | [19 empty] | -1.2273 | [5 empty] | 4.1152 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037584 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.5813 | [19 empty] | -1.2273 | [5 empty] | 4.1152 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037584 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.5123 | [19 empty] | -1.2273 | [5 empty] | 4.115 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037583 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.5123 | [19 empty] | -1.2273 | [5 empty] | 4.115 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037583 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.5123 | [19 empty] | -1.2274 | [5 empty] | 4.115 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037583 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.4433 | [19 empty] | -1.2274 | [5 empty] | 4.1149 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037582 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:46 | 20 | 20 |  | 2057.3744 | [19 empty] | -1.2274 | [5 empty] | 4.1147 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003758 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.3744 | [19 empty] | -1.2273 | [5 empty] | 4.1147 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003758 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.3744 | [19 empty] | -1.2273 | [5 empty] | 4.1147 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003758 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.3054 | [19 empty] | -1.2273 | [5 empty] | 4.1146 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037579 | -1.4604 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.3054 | [19 empty] | -1.2274 | [5 empty] | 4.1146 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037579 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.3054 | [19 empty] | -1.2273 | [5 empty] | 4.1146 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037579 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.2364 | [19 empty] | -1.2273 | [5 empty] | 4.1145 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037578 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.2364 | [19 empty] | -1.2274 | [5 empty] | 4.1145 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037578 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.1674 | [19 empty] | -1.2273 | [5 empty] | 4.1143 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037577 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.1674 | [19 empty] | -1.2274 | [5 empty] | 4.1143 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037577 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.0985 | [19 empty] | -1.2274 | [5 empty] | 4.1142 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037575 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:47 | 20 | 20 |  | 2057.0985 | [19 empty] | -1.2274 | [5 empty] | 4.1142 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037575 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2057.0985 | [19 empty] | -1.2273 | [5 empty] | 4.1142 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037575 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2057.0295 | [19 empty] | -1.2273 | [5 empty] | 4.1141 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037574 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2057.0295 | [19 empty] | -1.2273 | [5 empty] | 4.1141 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037574 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.9605 | [19 empty] | -1.2273 | [5 empty] | 4.1139 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037573 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.9605 | [19 empty] | -1.2274 | [5 empty] | 4.1139 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037573 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.9605 | [19 empty] | -1.2273 | [5 empty] | 4.1139 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037573 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.8915 | [19 empty] | -1.2273 | [5 empty] | 4.1138 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037572 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.8915 | [19 empty] | -1.2274 | [5 empty] | 4.1138 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037572 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.8915 | [19 empty] | -1.2273 | [5 empty] | 4.1138 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037572 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:48 | 20 | 20 |  | 2056.8226 | [19 empty] | -1.2274 | [5 empty] | 4.1136 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003757 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.8226 | [19 empty] | -1.2274 | [5 empty] | 4.1136 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003757 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.8226 | [19 empty] | -1.2273 | [5 empty] | 4.1136 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003757 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.7536 | [19 empty] | -1.2274 | [5 empty] | 4.1135 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037569 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.7536 | [19 empty] | -1.2274 | [5 empty] | 4.1135 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037569 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.6846 | [19 empty] | -1.2274 | [5 empty] | 4.1134 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037568 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.6846 | [19 empty] | -1.2273 | [5 empty] | 4.1134 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037568 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.6846 | [19 empty] | -1.2273 | [5 empty] | 4.1134 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037568 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.6156 | [19 empty] | -1.2274 | [5 empty] | 4.1132 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037566 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.6156 | [19 empty] | -1.2274 | [5 empty] | 4.1132 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037566 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:49 | 20 | 20 |  | 2056.5467 | [19 empty] | -1.2274 | [5 empty] | 4.1131 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037565 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.5467 | [19 empty] | -1.2273 | [5 empty] | 4.1131 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037565 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.5467 | [19 empty] | -1.2273 | [5 empty] | 4.1131 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037565 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.4777 | [19 empty] | -1.2274 | [5 empty] | 4.113 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037564 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.4777 | [19 empty] | -1.2273 | [5 empty] | 4.113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037564 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.4087 | [19 empty] | -1.2273 | [5 empty] | 4.1128 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037563 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.4087 | [19 empty] | -1.2273 | [5 empty] | 4.1128 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037563 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.3397 | [19 empty] | -1.2273 | [5 empty] | 4.1127 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037561 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.3397 | [19 empty] | -1.2273 | [5 empty] | 4.1127 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037561 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.3397 | [19 empty] | -1.2274 | [5 empty] | 4.1127 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037561 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.2708 | [19 empty] | -1.2273 | [5 empty] | 4.1125 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003756 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:50 | 20 | 20 |  | 2056.2708 | [19 empty] | -1.2273 | [5 empty] | 4.1125 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003756 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.2708 | [19 empty] | -1.2273 | [5 empty] | 4.1125 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003756 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.2018 | [19 empty] | -1.2273 | [5 empty] | 4.1124 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037559 | -1.4604 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.2018 | [19 empty] | -1.2274 | [5 empty] | 4.1124 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037559 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.1328 | [19 empty] | -1.2273 | [5 empty] | 4.1123 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037558 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.1328 | [19 empty] | -1.2273 | [5 empty] | 4.1123 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037558 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.1328 | [19 empty] | -1.2273 | [5 empty] | 4.1123 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037558 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.0639 | [19 empty] | -1.2274 | [5 empty] | 4.1121 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037556 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2056.0639 | [19 empty] | -1.2273 | [5 empty] | 4.1121 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037556 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2055.9949 | [19 empty] | -1.2274 | [5 empty] | 4.112 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037555 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:51 | 20 | 20 |  | 2055.9949 | [19 empty] | -1.2274 | [5 empty] | 4.112 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037555 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.9949 | [19 empty] | -1.2273 | [5 empty] | 4.112 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037555 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.9949 | [19 empty] | -1.2273 | [5 empty] | 4.112 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037555 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.9259 | [19 empty] | -1.2273 | [5 empty] | 4.1119 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037554 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.9259 | [19 empty] | -1.2274 | [5 empty] | 4.1119 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037554 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.8569 | [19 empty] | -1.2273 | [5 empty] | 4.1117 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037553 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.8569 | [19 empty] | -1.2273 | [5 empty] | 4.1117 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037553 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.8569 | [19 empty] | -1.2273 | [5 empty] | 4.1117 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037553 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.788 | [19 empty] | -1.2273 | [5 empty] | 4.1116 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037551 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.788 | [19 empty] | -1.2273 | [5 empty] | 4.1116 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037551 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.719 | [19 empty] | -1.2273 | [5 empty] | 4.1114 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003755 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:52 | 20 | 20 |  | 2055.719 | [19 empty] | -1.2274 | [5 empty] | 4.1114 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003755 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2274 | [5 empty] | 4.1113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2273 | [5 empty] | 4.1113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2274 | [5 empty] | 4.1113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2274 | [5 empty] | 4.1113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2274 | [5 empty] | 4.1113 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2274 | [5 empty] | 4.1113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.65 | [19 empty] | -1.2274 | [5 empty] | 4.1113 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037549 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.581 | [19 empty] | -1.2274 | [5 empty] | 4.1112 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037547 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.581 | [19 empty] | -1.2274 | [5 empty] | 4.1112 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037547 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:53 | 20 | 20 |  | 2055.5121 | [19 empty] | -1.2274 | [5 empty] | 4.111 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037546 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.5121 | [19 empty] | -1.2274 | [5 empty] | 4.111 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037546 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.5121 | [19 empty] | -1.2274 | [5 empty] | 4.111 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037546 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.5121 | [19 empty] | -1.2274 | [5 empty] | 4.111 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037546 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.4431 | [19 empty] | -1.2274 | [5 empty] | 4.1109 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037545 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.4431 | [19 empty] | -1.2274 | [5 empty] | 4.1109 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037545 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.3741 | [19 empty] | -1.2274 | [5 empty] | 4.1107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037544 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.3741 | [19 empty] | -1.2274 | [5 empty] | 4.1107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037544 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.3051 | [19 empty] | -1.2274 | [5 empty] | 4.1106 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037542 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.3051 | [19 empty] | -1.2273 | [5 empty] | 4.1106 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037542 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:54 | 20 | 20 |  | 2055.3051 | [19 empty] | -1.2274 | [5 empty] | 4.1106 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037542 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.2362 | [19 empty] | -1.2274 | [5 empty] | 4.1105 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037541 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.2362 | [19 empty] | -1.2274 | [5 empty] | 4.1105 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037541 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.1672 | [19 empty] | -1.2274 | [5 empty] | 4.1103 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003754 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.1672 | [19 empty] | -1.2274 | [5 empty] | 4.1103 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003754 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.1672 | [19 empty] | -1.2274 | [5 empty] | 4.1103 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003754 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.0982 | [19 empty] | -1.2274 | [5 empty] | 4.1102 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037539 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.0982 | [19 empty] | -1.2274 | [5 empty] | 4.1102 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037539 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.0982 | [19 empty] | -1.2274 | [5 empty] | 4.1102 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037539 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.0292 | [19 empty] | -1.2274 | [5 empty] | 4.1101 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037537 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2055.0292 | [19 empty] | -1.2274 | [5 empty] | 4.1101 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037537 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:55 | 20 | 20 |  | 2054.9603 | [19 empty] | -1.2274 | [5 empty] | 4.1099 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037536 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.9603 | [19 empty] | -1.2273 | [5 empty] | 4.1099 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037536 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.9603 | [19 empty] | -1.2273 | [5 empty] | 4.1099 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037536 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.8913 | [19 empty] | -1.2274 | [5 empty] | 4.1098 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037535 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.8913 | [19 empty] | -1.2273 | [5 empty] | 4.1098 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037535 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.8913 | [19 empty] | -1.2273 | [5 empty] | 4.1098 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037535 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.8223 | [19 empty] | -1.2274 | [5 empty] | 4.1096 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037534 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.8223 | [19 empty] | -1.2274 | [5 empty] | 4.1096 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037534 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.8223 | [19 empty] | -1.2273 | [5 empty] | 4.1096 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037534 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.7533 | [19 empty] | -1.2274 | [5 empty] | 4.1095 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037532 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:56 | 20 | 20 |  | 2054.7533 | [19 empty] | -1.2274 | [5 empty] | 4.1095 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037532 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.7533 | [19 empty] | -1.2274 | [5 empty] | 4.1095 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037532 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.6844 | [19 empty] | -1.2274 | [5 empty] | 4.1094 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037531 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.6844 | [19 empty] | -1.2274 | [5 empty] | 4.1094 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037531 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.6844 | [19 empty] | -1.2274 | [5 empty] | 4.1094 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037531 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.6154 | [19 empty] | -1.2273 | [5 empty] | 4.1092 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003753 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.6154 | [19 empty] | -1.2273 | [5 empty] | 4.1092 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003753 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.6154 | [19 empty] | -1.2273 | [5 empty] | 4.1092 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003753 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.5464 | [19 empty] | -1.2274 | [5 empty] | 4.1091 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037528 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.5464 | [19 empty] | -1.2273 | [5 empty] | 4.1091 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037528 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:57 | 20 | 20 |  | 2054.5464 | [19 empty] | -1.2272 | [5 empty] | 4.1091 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037528 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4774 | [19 empty] | -1.2274 | [5 empty] | 4.109 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037527 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4774 | [19 empty] | -1.2274 | [5 empty] | 4.109 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037527 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4774 | [19 empty] | -1.2274 | [5 empty] | 4.109 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037527 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4085 | [19 empty] | -1.2274 | [5 empty] | 4.1088 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037526 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4085 | [19 empty] | -1.2274 | [5 empty] | 4.1088 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037526 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4085 | [19 empty] | -1.2274 | [5 empty] | 4.1088 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037526 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.4085 | [19 empty] | -1.2274 | [5 empty] | 4.1088 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037526 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.3395 | [19 empty] | -1.2273 | [5 empty] | 4.1087 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037525 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.3395 | [19 empty] | -1.2274 | [5 empty] | 4.1087 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037525 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.3395 | [19 empty] | -1.2274 | [5 empty] | 4.1087 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037525 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:58 | 20 | 20 |  | 2054.2705 | [19 empty] | -1.2273 | [5 empty] | 4.1085 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037523 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.2705 | [19 empty] | -1.2274 | [5 empty] | 4.1085 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037523 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.2015 | [19 empty] | -1.2274 | [5 empty] | 4.1084 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037522 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.2015 | [19 empty] | -1.2273 | [5 empty] | 4.1084 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037522 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.2015 | [19 empty] | -1.2274 | [5 empty] | 4.1084 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037522 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.2015 | [19 empty] | -1.2274 | [5 empty] | 4.1084 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037522 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.1326 | [19 empty] | -1.2274 | [5 empty] | 4.1083 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037521 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.1326 | [19 empty] | -1.2274 | [5 empty] | 4.1083 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037521 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.1326 | [19 empty] | -1.2274 | [5 empty] | 4.1083 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037521 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.1326 | [19 empty] | -1.2274 | [5 empty] | 4.1083 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037521 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 08:59 | 20 | 20 |  | 2054.0636 | [19 empty] | -1.2274 | [5 empty] | 4.1081 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003752 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2054.0636 | [19 empty] | -1.2274 | [5 empty] | 4.1081 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003752 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2054.0636 | [19 empty] | -1.2273 | [5 empty] | 4.1081 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003752 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2054.0636 | [19 empty] | -1.2273 | [5 empty] | 4.1081 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003752 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.9946 | [19 empty] | -1.2273 | [5 empty] | 4.108 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037518 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.9946 | [19 empty] | -1.2273 | [5 empty] | 4.108 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037518 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.9946 | [19 empty] | -1.2274 | [5 empty] | 4.108 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037518 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.9256 | [19 empty] | -1.2274 | [5 empty] | 4.1079 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037517 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.9256 | [19 empty] | -1.2274 | [5 empty] | 4.1079 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037517 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.8567 | [19 empty] | -1.2273 | [5 empty] | 4.1077 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037516 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:00 | 20 | 20 |  | 2053.8567 | [19 empty] | -1.2274 | [5 empty] | 4.1077 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037516 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.8567 | [19 empty] | -1.2274 | [5 empty] | 4.1077 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037516 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.8567 | [19 empty] | -1.2274 | [5 empty] | 4.1077 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037516 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.7877 | [19 empty] | -1.2274 | [5 empty] | 4.1076 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037515 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.7877 | [19 empty] | -1.2273 | [5 empty] | 4.1076 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037515 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.7877 | [19 empty] | -1.2274 | [5 empty] | 4.1076 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037515 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.7187 | [19 empty] | -1.2274 | [5 empty] | 4.1074 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037513 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.7187 | [19 empty] | -1.2274 | [5 empty] | 4.1074 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037513 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.7187 | [19 empty] | -1.2274 | [5 empty] | 4.1074 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037513 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.6497 | [19 empty] | -1.2274 | [5 empty] | 4.1073 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037512 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.6497 | [19 empty] | -1.2274 | [5 empty] | 4.1073 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037512 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:01 | 20 | 20 |  | 2053.6497 | [19 empty] | -1.2274 | [5 empty] | 4.1073 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037512 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5808 | [19 empty] | -1.2274 | [5 empty] | 4.1072 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037511 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5808 | [19 empty] | -1.2274 | [5 empty] | 4.1072 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037511 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5808 | [19 empty] | -1.2274 | [5 empty] | 4.1072 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037511 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5118 | [19 empty] | -1.2274 | [5 empty] | 4.107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037509 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5118 | [19 empty] | -1.2273 | [5 empty] | 4.107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037509 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5118 | [19 empty] | -1.2274 | [5 empty] | 4.107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037509 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5118 | [19 empty] | -1.2273 | [5 empty] | 4.107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037509 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5118 | [19 empty] | -1.2274 | [5 empty] | 4.107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037509 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.5118 | [19 empty] | -1.2273 | [5 empty] | 4.107 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037509 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:02 | 20 | 20 |  | 2053.4428 | [19 empty] | -1.2274 | [5 empty] | 4.1069 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037508 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.4428 | [19 empty] | -1.2274 | [5 empty] | 4.1069 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037508 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.4428 | [19 empty] | -1.2274 | [5 empty] | 4.1069 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037508 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.4428 | [19 empty] | -1.2274 | [5 empty] | 4.1069 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037508 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.4428 | [19 empty] | -1.2274 | [5 empty] | 4.1069 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037508 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.3738 | [19 empty] | -1.2274 | [5 empty] | 4.1067 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037507 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.3738 | [19 empty] | -1.2273 | [5 empty] | 4.1067 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037507 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.3738 | [19 empty] | -1.2274 | [5 empty] | 4.1067 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037507 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.3738 | [19 empty] | -1.2273 | [5 empty] | 4.1067 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037507 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.3738 | [19 empty] | -1.2273 | [5 empty] | 4.1067 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037507 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:03 | 20 | 20 |  | 2053.3738 | [19 empty] | -1.2273 | [5 empty] | 4.1067 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037507 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.3049 | [19 empty] | -1.2274 | [5 empty] | 4.1066 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037506 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.3049 | [19 empty] | -1.2273 | [5 empty] | 4.1066 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037506 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.3049 | [19 empty] | -1.2274 | [5 empty] | 4.1066 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037506 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.3049 | [19 empty] | -1.2274 | [5 empty] | 4.1066 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037506 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.2359 | [19 empty] | -1.2274 | [5 empty] | 4.1065 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037504 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.2359 | [19 empty] | -1.2274 | [5 empty] | 4.1065 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037504 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.2359 | [19 empty] | -1.2273 | [5 empty] | 4.1065 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037504 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.1669 | [19 empty] | -1.2273 | [5 empty] | 4.1063 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037503 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.1669 | [19 empty] | -1.2274 | [5 empty] | 4.1063 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037503 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.1669 | [19 empty] | -1.2274 | [5 empty] | 4.1063 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037503 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:04 | 20 | 20 |  | 2053.0979 | [19 empty] | -1.2274 | [5 empty] | 4.1062 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037502 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2053.0979 | [19 empty] | -1.2274 | [5 empty] | 4.1062 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037502 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2053.0979 | [19 empty] | -1.2274 | [5 empty] | 4.1062 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037502 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2053.029 | [19 empty] | -1.2274 | [5 empty] | 4.1061 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037501 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2053.029 | [19 empty] | -1.2274 | [5 empty] | 4.1061 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037501 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2053.029 | [19 empty] | -1.2274 | [5 empty] | 4.1061 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037501 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2053.029 | [19 empty] | -1.2274 | [5 empty] | 4.1061 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037501 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2274 | [5 empty] | 4.1059 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2274 | [5 empty] | 4.1059 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2274 | [5 empty] | 4.1059 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:05 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2274 | [5 empty] | 4.1059 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2274 | [5 empty] | 4.1059 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2273 | [5 empty] | 4.1059 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.96 | [19 empty] | -1.2273 | [5 empty] | 4.1059 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037499 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.891 | [19 empty] | -1.2273 | [5 empty] | 4.1058 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037498 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.891 | [19 empty] | -1.2274 | [5 empty] | 4.1058 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037498 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.891 | [19 empty] | -1.2273 | [5 empty] | 4.1058 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037498 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.822 | [19 empty] | -1.2274 | [5 empty] | 4.1056 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037497 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.822 | [19 empty] | -1.2274 | [5 empty] | 4.1056 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037497 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.822 | [19 empty] | -1.2274 | [5 empty] | 4.1056 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037497 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:06 | 20 | 20 |  | 2052.822 | [19 empty] | -1.2273 | [5 empty] | 4.1056 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037497 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.822 | [19 empty] | -1.2274 | [5 empty] | 4.1056 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037497 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.7531 | [19 empty] | -1.2273 | [5 empty] | 4.1055 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037496 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.7531 | [19 empty] | -1.2274 | [5 empty] | 4.1055 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037496 | -1.4606 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.7531 | [19 empty] | -1.2273 | [5 empty] | 4.1055 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037496 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.6841 | [19 empty] | -1.2273 | [5 empty] | 4.1054 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037494 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.6841 | [19 empty] | -1.2274 | [5 empty] | 4.1054 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037494 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.6841 | [19 empty] | -1.2273 | [5 empty] | 4.1054 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037494 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.6151 | [19 empty] | -1.2273 | [5 empty] | 4.1052 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037493 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.6151 | [19 empty] | -1.2274 | [5 empty] | 4.1052 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037493 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.6151 | [19 empty] | -1.2274 | [5 empty] | 4.1052 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037493 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:07 | 20 | 20 |  | 2052.5461 | [19 empty] | -1.2274 | [5 empty] | 4.1051 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037492 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.5461 | [19 empty] | -1.2274 | [5 empty] | 4.1051 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037492 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4772 | [19 empty] | -1.2273 | [5 empty] | 4.105 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003749 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4772 | [19 empty] | -1.2273 | [5 empty] | 4.105 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003749 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4772 | [19 empty] | -1.2273 | [5 empty] | 4.105 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003749 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4082 | [19 empty] | -1.2273 | [5 empty] | 4.1048 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037489 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4082 | [19 empty] | -1.2273 | [5 empty] | 4.1048 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037489 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4082 | [19 empty] | -1.2273 | [5 empty] | 4.1048 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037489 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.4082 | [19 empty] | -1.2273 | [5 empty] | 4.1048 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037489 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.3392 | [19 empty] | -1.2273 | [5 empty] | 4.1047 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037488 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:08 | 20 | 20 |  | 2052.3392 | [19 empty] | -1.2273 | [5 empty] | 4.1047 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037488 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.3392 | [19 empty] | -1.2274 | [5 empty] | 4.1047 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037488 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.3392 | [19 empty] | -1.2273 | [5 empty] | 4.1047 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037488 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.3392 | [19 empty] | -1.2273 | [5 empty] | 4.1047 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037488 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2702 | [19 empty] | -1.2273 | [5 empty] | 4.1045 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037487 | -1.4604 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2702 | [19 empty] | -1.2274 | [5 empty] | 4.1045 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037487 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2702 | [19 empty] | -1.2274 | [5 empty] | 4.1045 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037487 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2702 | [19 empty] | -1.2273 | [5 empty] | 4.1045 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037487 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2013 | [19 empty] | -1.2273 | [5 empty] | 4.1044 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037485 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2013 | [19 empty] | -1.2273 | [5 empty] | 4.1044 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037485 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:09 | 20 | 20 |  | 2052.2013 | [19 empty] | -1.2273 | [5 empty] | 4.1044 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037485 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2052.1323 | [19 empty] | -1.2273 | [5 empty] | 4.1043 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037484 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2052.1323 | [19 empty] | -1.2273 | [5 empty] | 4.1043 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037484 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2052.0633 | [19 empty] | -1.2273 | [5 empty] | 4.1041 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037483 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2052.0633 | [19 empty] | -1.2273 | [5 empty] | 4.1041 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037483 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2052.0633 | [19 empty] | -1.2273 | [5 empty] | 4.1041 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037483 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2052.0633 | [19 empty] | -1.2273 | [5 empty] | 4.1041 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037483 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2051.9943 | [19 empty] | -1.2273 | [5 empty] | 4.104 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037482 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2051.9943 | [19 empty] | -1.2273 | [5 empty] | 4.104 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037482 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2051.9254 | [19 empty] | -1.2273 | [5 empty] | 4.1039 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003748 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2051.9254 | [19 empty] | -1.2273 | [5 empty] | 4.1039 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003748 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:10 | 20 | 20 |  | 2051.9254 | [19 empty] | -1.2273 | [5 empty] | 4.1039 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003748 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.9254 | [19 empty] | -1.2273 | [5 empty] | 4.1039 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003748 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.8564 | [19 empty] | -1.2273 | [5 empty] | 4.1037 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037479 | -1.4604 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.8564 | [19 empty] | -1.2274 | [5 empty] | 4.1037 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037479 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.8564 | [19 empty] | -1.2274 | [5 empty] | 4.1037 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037479 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.7874 | [19 empty] | -1.2273 | [5 empty] | 4.1036 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037478 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.7874 | [19 empty] | -1.2274 | [5 empty] | 4.1036 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037478 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.7874 | [19 empty] | -1.2273 | [5 empty] | 4.1036 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037478 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.7874 | [19 empty] | -1.2273 | [5 empty] | 4.1036 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037478 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.7874 | [19 empty] | -1.2273 | [5 empty] | 4.1036 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037478 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:11 | 20 | 20 |  | 2051.7185 | [19 empty] | -1.2274 | [5 empty] | 4.1034 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037477 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.7185 | [19 empty] | -1.2273 | [5 empty] | 4.1034 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037477 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.7185 | [19 empty] | -1.2274 | [5 empty] | 4.1034 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037477 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.7185 | [19 empty] | -1.2274 | [5 empty] | 4.1034 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037477 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.7185 | [19 empty] | -1.2274 | [5 empty] | 4.1034 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037477 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.6495 | [19 empty] | -1.2274 | [5 empty] | 4.1033 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037475 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.6495 | [19 empty] | -1.2274 | [5 empty] | 4.1033 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037475 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.6495 | [19 empty] | -1.2274 | [5 empty] | 4.1033 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037475 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.5805 | [19 empty] | -1.2274 | [5 empty] | 4.1032 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037474 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.5805 | [19 empty] | -1.2274 | [5 empty] | 4.1032 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037474 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.5805 | [19 empty] | -1.2274 | [5 empty] | 4.1032 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037474 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:12 | 20 | 20 |  | 2051.5115 | [19 empty] | -1.2274 | [5 empty] | 4.103 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037473 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.5115 | [19 empty] | -1.2274 | [5 empty] | 4.103 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037473 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.5115 | [19 empty] | -1.2274 | [5 empty] | 4.103 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037473 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.5115 | [19 empty] | -1.2274 | [5 empty] | 4.103 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037473 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.4426 | [19 empty] | -1.2274 | [5 empty] | 4.1029 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037472 | -1.4605 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.4426 | [19 empty] | -1.2274 | [5 empty] | 4.1029 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.0037472 | -1.4605 | -1.4413 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.4426 | [19 empty] | -1.2274 | [5 empty] | 4.1029 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037472 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.3736 | [19 empty] | -1.2274 | [5 empty] | 4.1027 | -1.2266 | 0.90444 | 17.6664 | 0.23773 | -0.003747 | -1.4606 | -1.4413 | -1.4413 | -1.4413 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.3736 | [19 empty] | -1.2273 | [5 empty] | 4.1027 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003747 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.3736 | [19 empty] | -1.2274 | [5 empty] | 4.1027 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003747 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:13 | 20 | 20 |  | 2051.3736 | [19 empty] | -1.2273 | [5 empty] | 4.1027 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003747 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.3046 | [19 empty] | -1.2274 | [5 empty] | 4.1026 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037469 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.3046 | [19 empty] | -1.2274 | [5 empty] | 4.1026 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037469 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.3046 | [19 empty] | -1.2273 | [5 empty] | 4.1026 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037469 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.3046 | [19 empty] | -1.2274 | [5 empty] | 4.1026 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037469 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.2356 | [19 empty] | -1.2273 | [5 empty] | 4.1025 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037468 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.2356 | [19 empty] | -1.2273 | [5 empty] | 4.1025 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037468 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.2356 | [19 empty] | -1.2273 | [5 empty] | 4.1025 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037468 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.1667 | [19 empty] | -1.2274 | [5 empty] | 4.1023 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037466 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.1667 | [19 empty] | -1.2274 | [5 empty] | 4.1023 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037466 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:14 | 20 | 20 |  | 2051.1667 | [19 empty] | -1.2274 | [5 empty] | 4.1023 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037466 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.1667 | [19 empty] | -1.2274 | [5 empty] | 4.1023 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037466 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.1667 | [19 empty] | -1.2273 | [5 empty] | 4.1023 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037466 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.0977 | [19 empty] | -1.2274 | [5 empty] | 4.1022 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037465 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.0977 | [19 empty] | -1.2273 | [5 empty] | 4.1022 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037465 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.0977 | [19 empty] | -1.2273 | [5 empty] | 4.1022 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037465 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.0287 | [19 empty] | -1.2273 | [5 empty] | 4.1021 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037464 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.0287 | [19 empty] | -1.2273 | [5 empty] | 4.1021 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037464 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2051.0287 | [19 empty] | -1.2273 | [5 empty] | 4.1021 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037464 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2050.9597 | [19 empty] | -1.2273 | [5 empty] | 4.1019 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037463 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2050.9597 | [19 empty] | -1.2274 | [5 empty] | 4.1019 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037463 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:15 | 20 | 20 |  | 2050.9597 | [19 empty] | -1.2273 | [5 empty] | 4.1019 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037463 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.9597 | [19 empty] | -1.2274 | [5 empty] | 4.1019 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037463 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.9597 | [19 empty] | -1.2273 | [5 empty] | 4.1019 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037463 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.9597 | [19 empty] | -1.2273 | [5 empty] | 4.1019 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037463 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8908 | [19 empty] | -1.2273 | [5 empty] | 4.1018 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037461 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8908 | [19 empty] | -1.2273 | [5 empty] | 4.1018 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037461 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8908 | [19 empty] | -1.2274 | [5 empty] | 4.1018 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037461 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8908 | [19 empty] | -1.2273 | [5 empty] | 4.1018 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037461 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8218 | [19 empty] | -1.2273 | [5 empty] | 4.1016 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003746 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8218 | [19 empty] | -1.2273 | [5 empty] | 4.1016 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003746 | -1.4604 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:16 | 20 | 20 |  | 2050.8218 | [19 empty] | -1.2273 | [5 empty] | 4.1016 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003746 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.8218 | [19 empty] | -1.2273 | [5 empty] | 4.1016 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.003746 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.7528 | [19 empty] | -1.2273 | [5 empty] | 4.1015 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037459 | -1.4605 | -1.4412 | -1.4412 | -1.4412 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.7528 | [19 empty] | -1.2273 | [5 empty] | 4.1015 | -1.2265 | 0.90444 | 17.6664 | 0.23773 | -0.0037459 | -1.4605 | -1.4412 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.7528 | [19 empty] | -1.2273 | [5 empty] | 4.1015 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037459 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.7528 | [19 empty] | -1.2273 | [5 empty] | 4.1015 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037459 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.6838 | [19 empty] | -1.2272 | [5 empty] | 4.1014 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037458 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.6838 | [19 empty] | -1.2273 | [5 empty] | 4.1014 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037458 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.6838 | [19 empty] | -1.2273 | [5 empty] | 4.1014 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037458 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.6149 | [19 empty] | -1.2272 | [5 empty] | 4.1012 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037456 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:17 | 20 | 20 |  | 2050.6149 | [19 empty] | -1.2272 | [5 empty] | 4.1012 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037456 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.6149 | [19 empty] | -1.2273 | [5 empty] | 4.1012 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037456 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.6149 | [19 empty] | -1.2273 | [5 empty] | 4.1012 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037456 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2273 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2273 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2272 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2273 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2272 | [5 empty] | 4.1011 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2272 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2272 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.5459 | [19 empty] | -1.2272 | [5 empty] | 4.1011 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037455 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:18 | 20 | 20 |  | 2050.4769 | [19 empty] | -1.2272 | [5 empty] | 4.101 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037454 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.4769 | [19 empty] | -1.2272 | [5 empty] | 4.101 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037454 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.4079 | [19 empty] | -1.2272 | [5 empty] | 4.1008 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037453 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.4079 | [19 empty] | -1.2272 | [5 empty] | 4.1008 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037453 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.4079 | [19 empty] | -1.2272 | [5 empty] | 4.1008 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037453 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.4079 | [19 empty] | -1.2272 | [5 empty] | 4.1008 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037453 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.4079 | [19 empty] | -1.2272 | [5 empty] | 4.1008 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037453 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.339 | [19 empty] | -1.2272 | [5 empty] | 4.1007 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037451 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.339 | [19 empty] | -1.2272 | [5 empty] | 4.1007 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037451 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.339 | [19 empty] | -1.2272 | [5 empty] | 4.1007 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037451 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:19 | 20 | 20 |  | 2050.339 | [19 empty] | -1.2273 | [5 empty] | 4.1007 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037451 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.27 | [19 empty] | -1.2272 | [5 empty] | 4.1005 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003745 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.27 | [19 empty] | -1.2272 | [5 empty] | 4.1005 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003745 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.27 | [19 empty] | -1.2272 | [5 empty] | 4.1005 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003745 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.27 | [19 empty] | -1.2273 | [5 empty] | 4.1005 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003745 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.201 | [19 empty] | -1.2272 | [5 empty] | 4.1004 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037449 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.201 | [19 empty] | -1.2272 | [5 empty] | 4.1004 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037449 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.201 | [19 empty] | -1.2272 | [5 empty] | 4.1004 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037449 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.201 | [19 empty] | -1.2272 | [5 empty] | 4.1004 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037449 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.201 | [19 empty] | -1.2272 | [5 empty] | 4.1004 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037449 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:20 | 20 | 20 |  | 2050.201 | [19 empty] | -1.2272 | [5 empty] | 4.1004 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037449 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.132 | [19 empty] | -1.2272 | [5 empty] | 4.1003 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037447 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.132 | [19 empty] | -1.2272 | [5 empty] | 4.1003 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037447 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.132 | [19 empty] | -1.2272 | [5 empty] | 4.1003 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037447 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.132 | [19 empty] | -1.2273 | [5 empty] | 4.1003 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037447 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.132 | [19 empty] | -1.2273 | [5 empty] | 4.1003 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037447 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2273 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2272 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2273 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2272 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2273 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:21 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2273 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2050.0631 | [19 empty] | -1.2273 | [5 empty] | 4.1001 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037446 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2273 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2272 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2272 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2273 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2273 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2272 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9941 | [19 empty] | -1.2273 | [5 empty] | 4.1 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037445 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9251 | [19 empty] | -1.2272 | [5 empty] | 4.0999 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037444 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:22 | 20 | 20 |  | 2049.9251 | [19 empty] | -1.2272 | [5 empty] | 4.0999 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037444 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.9251 | [19 empty] | -1.2272 | [5 empty] | 4.0999 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037444 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.9251 | [19 empty] | -1.2272 | [5 empty] | 4.0999 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037444 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.9251 | [19 empty] | -1.2273 | [5 empty] | 4.0999 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037444 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.8561 | [19 empty] | -1.2273 | [5 empty] | 4.0997 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037442 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.8561 | [19 empty] | -1.2273 | [5 empty] | 4.0997 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037442 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.8561 | [19 empty] | -1.2272 | [5 empty] | 4.0997 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037442 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.8561 | [19 empty] | -1.2272 | [5 empty] | 4.0997 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037442 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.7872 | [19 empty] | -1.2272 | [5 empty] | 4.0996 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037441 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.7872 | [19 empty] | -1.2272 | [5 empty] | 4.0996 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037441 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:23 | 20 | 20 |  | 2049.7872 | [19 empty] | -1.2272 | [5 empty] | 4.0996 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037441 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7872 | [19 empty] | -1.2272 | [5 empty] | 4.0996 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037441 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7872 | [19 empty] | -1.2272 | [5 empty] | 4.0996 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037441 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7872 | [19 empty] | -1.2272 | [5 empty] | 4.0996 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037441 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:24 | 20 | 20 |  | 2049.7182 | [19 empty] | -1.2272 | [5 empty] | 4.0994 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003744 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.6492 | [19 empty] | -1.2272 | [5 empty] | 4.0993 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037439 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.5802 | [19 empty] | -1.2272 | [5 empty] | 4.0992 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037437 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.5802 | [19 empty] | -1.2272 | [5 empty] | 4.0992 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037437 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:25 | 20 | 20 |  | 2049.5802 | [19 empty] | -1.2272 | [5 empty] | 4.0992 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037437 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5802 | [19 empty] | -1.2271 | [5 empty] | 4.0992 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037437 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5802 | [19 empty] | -1.2272 | [5 empty] | 4.0992 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037437 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.5113 | [19 empty] | -1.2272 | [5 empty] | 4.099 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037436 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:26 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.4423 | [19 empty] | -1.2272 | [5 empty] | 4.0989 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037435 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:27 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3733 | [19 empty] | -1.2272 | [5 empty] | 4.0987 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037434 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2273 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2273 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2273 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2272 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2273 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:28 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2272 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2273 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2272 | [5 empty] | 4.0985 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2273 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2272 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:29 | 20 | 20 |  | 2049.3043 | [19 empty] | -1.2273 | [5 empty] | 4.0986 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037432 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2272 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2272 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2273 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.2354 | [19 empty] | -1.2272 | [5 empty] | 4.0985 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037431 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.1664 | [19 empty] | -1.2272 | [5 empty] | 4.0983 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003743 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.1664 | [19 empty] | -1.2272 | [5 empty] | 4.0983 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003743 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.1664 | [19 empty] | -1.2272 | [5 empty] | 4.0983 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003743 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.1664 | [19 empty] | -1.2272 | [5 empty] | 4.0983 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003743 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.1664 | [19 empty] | -1.2272 | [5 empty] | 4.0983 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003743 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2273 | [5 empty] | 4.0982 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:30 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2272 | [5 empty] | 4.0982 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2272 | [5 empty] | 4.0982 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2272 | [5 empty] | 4.0982 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2272 | [5 empty] | 4.0982 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2273 | [5 empty] | 4.0982 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0974 | [19 empty] | -1.2272 | [5 empty] | 4.0982 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037428 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2272 | [5 empty] | 4.0981 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2272 | [5 empty] | 4.0981 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2273 | [5 empty] | 4.0981 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2273 | [5 empty] | 4.0981 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:31 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2272 | [5 empty] | 4.0981 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2272 | [5 empty] | 4.0981 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2272 | [5 empty] | 4.0981 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2049.0284 | [19 empty] | -1.2272 | [5 empty] | 4.0981 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037427 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:32 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2271 | [5 empty] | 4.0978 | -1.2262 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4602 | -1.4409 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2271 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:33 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:34 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2273 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:35 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:36 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:37 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:38 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:39 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:40 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:41 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:42 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:43 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2272 | [5 empty] | 4.0971 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.5456 | [19 empty] | -1.2273 | [5 empty] | 4.0971 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037418 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:44 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:45 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:46 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:47 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2273 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:48 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4604 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6146 | [19 empty] | -1.2272 | [5 empty] | 4.0972 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.003742 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:49 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:50 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:51 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.4411 | -1.4411 | -1.4411 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2273 | [5 empty] | 4.0974 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.6836 | [19 empty] | -1.2272 | [5 empty] | 4.0974 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037421 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2273 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2273 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:52 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2273 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2273 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2273 | [5 empty] | 4.0975 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.7525 | [19 empty] | -1.2272 | [5 empty] | 4.0975 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037422 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2273 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:53 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2273 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2273 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:54 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2273 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2273 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2273 | [5 empty] | 4.0976 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:55 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2273 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:56 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2273 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8215 | [19 empty] | -1.2272 | [5 empty] | 4.0976 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037423 | -1.4603 | -1.441 | -1.441 | -1.441 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:57 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2273 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.8905 | [19 empty] | -1.2272 | [5 empty] | 4.0978 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037425 | -1.4603 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:58 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2264 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4604 | -1.4411 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 2048.9595 | [19 empty] | -1.2272 | [5 empty] | 4.0979 | -1.2263 | 0.90444 | 17.6664 | 0.23773 | -0.0037426 | -1.4603 | -1.441 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 2011.2993 | [19 empty] | -1.2267 | [5 empty] | 4.0226 | -1.2258 | 0.90444 | 17.6664 | 0.23773 | -0.0036735 | -1.4598 | -1.4406 | -1.4406 | -1.4406 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 1811.6179 | [19 empty] | -1.2262 | [5 empty] | 3.6232 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0033069 | -1.4597 | -1.4405 | -1.4405 | -1.4405 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 1725.2617 | [19 empty] | -1.2262 | [5 empty] | 3.4505 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031484 | -1.4599 | -1.4406 | -1.4406 | -1.4406 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 1722.4337 | [19 empty] | -1.2262 | [5 empty] | 3.4449 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031432 | -1.4599 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 1723.0545 | [19 empty] | -1.2262 | [5 empty] | 3.4461 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031444 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 09:59 | 20 | 20 |  | 1723.4683 | [19 empty] | -1.2262 | [5 empty] | 3.4469 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031451 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1723.8132 | [19 empty] | -1.2263 | [5 empty] | 3.4476 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031458 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1724.1581 | [19 empty] | -1.2263 | [5 empty] | 3.4483 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031464 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1724.434 | [19 empty] | -1.2263 | [5 empty] | 3.4489 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031469 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1724.7099 | [19 empty] | -1.2263 | [5 empty] | 3.4494 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031474 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1724.9168 | [19 empty] | -1.2262 | [5 empty] | 3.4498 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031478 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1725.1237 | [19 empty] | -1.2263 | [5 empty] | 3.4502 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031482 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1725.3306 | [19 empty] | -1.2263 | [5 empty] | 3.4507 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031486 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1725.4686 | [19 empty] | -1.2263 | [5 empty] | 3.4509 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031488 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1725.6755 | [19 empty] | -1.2263 | [5 empty] | 3.4514 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031492 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1725.7445 | [19 empty] | -1.2263 | [5 empty] | 3.4515 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031493 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:00 | 20 | 20 |  | 1725.9514 | [19 empty] | -1.2263 | [5 empty] | 3.4519 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031497 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.0894 | [19 empty] | -1.2262 | [5 empty] | 3.4522 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00315 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.2273 | [19 empty] | -1.2263 | [5 empty] | 3.4525 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031502 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.3653 | [19 empty] | -1.2263 | [5 empty] | 3.4527 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031505 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.5032 | [19 empty] | -1.2262 | [5 empty] | 3.453 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031507 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.6412 | [19 empty] | -1.2262 | [5 empty] | 3.4533 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003151 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.7101 | [19 empty] | -1.2262 | [5 empty] | 3.4534 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031511 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.8481 | [19 empty] | -1.2262 | [5 empty] | 3.4537 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031513 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1726.9171 | [19 empty] | -1.2262 | [5 empty] | 3.4538 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031515 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1727.055 | [19 empty] | -1.2262 | [5 empty] | 3.4541 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031517 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:01 | 20 | 20 |  | 1727.193 | [19 empty] | -1.2262 | [5 empty] | 3.4544 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003152 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.2619 | [19 empty] | -1.2262 | [5 empty] | 3.4545 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031521 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.3999 | [19 empty] | -1.2262 | [5 empty] | 3.4548 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031524 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.4689 | [19 empty] | -1.2262 | [5 empty] | 3.4549 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031525 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.6068 | [19 empty] | -1.2262 | [5 empty] | 3.4552 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031527 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.6758 | [19 empty] | -1.2263 | [5 empty] | 3.4554 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031529 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.7448 | [19 empty] | -1.2263 | [5 empty] | 3.4555 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003153 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.8137 | [19 empty] | -1.2262 | [5 empty] | 3.4556 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031531 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1727.9517 | [19 empty] | -1.2262 | [5 empty] | 3.4559 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031534 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1728.0207 | [19 empty] | -1.2263 | [5 empty] | 3.456 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031535 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:02 | 20 | 20 |  | 1728.0896 | [19 empty] | -1.2263 | [5 empty] | 3.4562 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031536 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.1586 | [19 empty] | -1.2262 | [5 empty] | 3.4563 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031537 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.2966 | [19 empty] | -1.2262 | [5 empty] | 3.4566 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003154 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.3655 | [19 empty] | -1.2263 | [5 empty] | 3.4567 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031541 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.4345 | [19 empty] | -1.2263 | [5 empty] | 3.4569 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031543 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.5035 | [19 empty] | -1.2263 | [5 empty] | 3.457 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031544 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.5725 | [19 empty] | -1.2262 | [5 empty] | 3.4571 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031545 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.6414 | [19 empty] | -1.2263 | [5 empty] | 3.4573 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031546 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.7104 | [19 empty] | -1.2263 | [5 empty] | 3.4574 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031548 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.8483 | [19 empty] | -1.2262 | [5 empty] | 3.4577 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003155 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.8483 | [19 empty] | -1.2262 | [5 empty] | 3.4577 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003155 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:03 | 20 | 20 |  | 1728.9863 | [19 empty] | -1.2262 | [5 empty] | 3.458 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031553 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.0553 | [19 empty] | -1.2263 | [5 empty] | 3.4581 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031554 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.1242 | [19 empty] | -1.2262 | [5 empty] | 3.4582 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031555 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.1932 | [19 empty] | -1.2263 | [5 empty] | 3.4584 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031556 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.2622 | [19 empty] | -1.2263 | [5 empty] | 3.4585 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031558 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.3312 | [19 empty] | -1.2262 | [5 empty] | 3.4587 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031559 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.3312 | [19 empty] | -1.2262 | [5 empty] | 3.4587 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031559 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.4691 | [19 empty] | -1.2262 | [5 empty] | 3.4589 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031562 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.5381 | [19 empty] | -1.2263 | [5 empty] | 3.4591 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031563 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.5381 | [19 empty] | -1.2263 | [5 empty] | 3.4591 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031563 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:04 | 20 | 20 |  | 1729.6071 | [19 empty] | -1.2263 | [5 empty] | 3.4592 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031564 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1729.676 | [19 empty] | -1.2263 | [5 empty] | 3.4594 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031565 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1729.745 | [19 empty] | -1.2262 | [5 empty] | 3.4595 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031567 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1729.814 | [19 empty] | -1.2262 | [5 empty] | 3.4596 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031568 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1729.883 | [19 empty] | -1.2262 | [5 empty] | 3.4598 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031569 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2263 | [5 empty] | 3.4599 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1729.9519 | [19 empty] | -1.2263 | [5 empty] | 3.4599 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003157 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1730.0209 | [19 empty] | -1.2263 | [5 empty] | 3.46 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031572 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1730.0899 | [19 empty] | -1.2263 | [5 empty] | 3.4602 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031573 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1730.1589 | [19 empty] | -1.2263 | [5 empty] | 3.4603 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031574 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:05 | 20 | 20 |  | 1730.2278 | [19 empty] | -1.2263 | [5 empty] | 3.4605 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031575 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2263 | [5 empty] | 3.4606 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.2968 | [19 empty] | -1.2263 | [5 empty] | 3.4606 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031577 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.3658 | [19 empty] | -1.2262 | [5 empty] | 3.4607 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031578 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.4348 | [19 empty] | -1.2262 | [5 empty] | 3.4609 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031579 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.5037 | [19 empty] | -1.2263 | [5 empty] | 3.461 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031581 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.5727 | [19 empty] | -1.2262 | [5 empty] | 3.4611 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031582 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.5727 | [19 empty] | -1.2263 | [5 empty] | 3.4611 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031582 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.6417 | [19 empty] | -1.2263 | [5 empty] | 3.4613 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031583 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.7107 | [19 empty] | -1.2262 | [5 empty] | 3.4614 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031584 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.7796 | [19 empty] | -1.2263 | [5 empty] | 3.4616 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031586 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:06 | 20 | 20 |  | 1730.7796 | [19 empty] | -1.2263 | [5 empty] | 3.4616 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031586 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1730.8486 | [19 empty] | -1.2263 | [5 empty] | 3.4617 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031587 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1730.9176 | [19 empty] | -1.2263 | [5 empty] | 3.4618 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031588 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1730.9866 | [19 empty] | -1.2262 | [5 empty] | 3.462 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031589 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1730.9866 | [19 empty] | -1.2262 | [5 empty] | 3.462 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031589 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1731.0555 | [19 empty] | -1.2263 | [5 empty] | 3.4621 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031591 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2263 | [5 empty] | 3.4622 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1731.1245 | [19 empty] | -1.2263 | [5 empty] | 3.4622 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031592 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1731.1935 | [19 empty] | -1.2262 | [5 empty] | 3.4624 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031593 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1731.2625 | [19 empty] | -1.2262 | [5 empty] | 3.4625 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031594 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:07 | 20 | 20 |  | 1731.2625 | [19 empty] | -1.2262 | [5 empty] | 3.4625 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031594 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.3314 | [19 empty] | -1.2262 | [5 empty] | 3.4627 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031596 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.4004 | [19 empty] | -1.2262 | [5 empty] | 3.4628 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031597 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.4694 | [19 empty] | -1.2263 | [5 empty] | 3.4629 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031598 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.4694 | [19 empty] | -1.2262 | [5 empty] | 3.4629 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031598 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.5384 | [19 empty] | -1.2262 | [5 empty] | 3.4631 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00316 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.6073 | [19 empty] | -1.2262 | [5 empty] | 3.4632 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031601 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.6073 | [19 empty] | -1.2263 | [5 empty] | 3.4632 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031601 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.6763 | [19 empty] | -1.2263 | [5 empty] | 3.4634 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031602 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.7453 | [19 empty] | -1.2263 | [5 empty] | 3.4635 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031603 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:08 | 20 | 20 |  | 1731.7453 | [19 empty] | -1.2262 | [5 empty] | 3.4635 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031603 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1731.8143 | [19 empty] | -1.2262 | [5 empty] | 3.4636 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031605 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1731.8143 | [19 empty] | -1.2262 | [5 empty] | 3.4636 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031605 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1731.8832 | [19 empty] | -1.2263 | [5 empty] | 3.4638 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031606 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1731.9522 | [19 empty] | -1.2262 | [5 empty] | 3.4639 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031607 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.0212 | [19 empty] | -1.2262 | [5 empty] | 3.464 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031608 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.0212 | [19 empty] | -1.2262 | [5 empty] | 3.464 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031608 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.0902 | [19 empty] | -1.2262 | [5 empty] | 3.4642 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003161 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.1591 | [19 empty] | -1.2262 | [5 empty] | 3.4643 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031611 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.1591 | [19 empty] | -1.2263 | [5 empty] | 3.4643 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031611 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.1591 | [19 empty] | -1.2263 | [5 empty] | 3.4643 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031611 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:09 | 20 | 20 |  | 1732.2281 | [19 empty] | -1.2262 | [5 empty] | 3.4645 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031612 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.2971 | [19 empty] | -1.2262 | [5 empty] | 3.4646 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031613 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.2971 | [19 empty] | -1.2262 | [5 empty] | 3.4646 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031613 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.3661 | [19 empty] | -1.2262 | [5 empty] | 3.4647 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031615 | -1.4598 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.3661 | [19 empty] | -1.2262 | [5 empty] | 3.4647 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031615 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.435 | [19 empty] | -1.2263 | [5 empty] | 3.4649 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031616 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.435 | [19 empty] | -1.2262 | [5 empty] | 3.4649 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031616 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.504 | [19 empty] | -1.2263 | [5 empty] | 3.465 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.504 | [19 empty] | -1.2263 | [5 empty] | 3.465 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.504 | [19 empty] | -1.2263 | [5 empty] | 3.465 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031617 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:10 | 20 | 20 |  | 1732.573 | [19 empty] | -1.2263 | [5 empty] | 3.4651 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031619 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.573 | [19 empty] | -1.2262 | [5 empty] | 3.4651 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031619 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.642 | [19 empty] | -1.2262 | [5 empty] | 3.4653 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003162 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.7109 | [19 empty] | -1.2262 | [5 empty] | 3.4654 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031621 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.7109 | [19 empty] | -1.2263 | [5 empty] | 3.4654 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031621 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.7109 | [19 empty] | -1.2262 | [5 empty] | 3.4654 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031621 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.7799 | [19 empty] | -1.2263 | [5 empty] | 3.4656 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031622 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.8489 | [19 empty] | -1.2263 | [5 empty] | 3.4657 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031624 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.8489 | [19 empty] | -1.2263 | [5 empty] | 3.4657 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031624 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.9179 | [19 empty] | -1.2263 | [5 empty] | 3.4658 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031625 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:11 | 20 | 20 |  | 1732.9179 | [19 empty] | -1.2263 | [5 empty] | 3.4658 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031625 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1732.9179 | [19 empty] | -1.2262 | [5 empty] | 3.4658 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031625 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1732.9868 | [19 empty] | -1.2262 | [5 empty] | 3.466 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031626 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.0558 | [19 empty] | -1.2263 | [5 empty] | 3.4661 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031627 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.1248 | [19 empty] | -1.2262 | [5 empty] | 3.4662 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031629 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.1248 | [19 empty] | -1.2262 | [5 empty] | 3.4662 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031629 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.1937 | [19 empty] | -1.2262 | [5 empty] | 3.4664 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003163 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.1937 | [19 empty] | -1.2262 | [5 empty] | 3.4664 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003163 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.2627 | [19 empty] | -1.2262 | [5 empty] | 3.4665 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031631 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.2627 | [19 empty] | -1.2262 | [5 empty] | 3.4665 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031631 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.2627 | [19 empty] | -1.2262 | [5 empty] | 3.4665 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031631 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:12 | 20 | 20 |  | 1733.3317 | [19 empty] | -1.2263 | [5 empty] | 3.4667 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031632 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.3317 | [19 empty] | -1.2263 | [5 empty] | 3.4667 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031632 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.4007 | [19 empty] | -1.2263 | [5 empty] | 3.4668 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031634 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.4007 | [19 empty] | -1.2262 | [5 empty] | 3.4668 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031634 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.4696 | [19 empty] | -1.2263 | [5 empty] | 3.4669 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031635 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.4696 | [19 empty] | -1.2263 | [5 empty] | 3.4669 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031635 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.4696 | [19 empty] | -1.2263 | [5 empty] | 3.4669 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031635 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.5386 | [19 empty] | -1.2262 | [5 empty] | 3.4671 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031636 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.5386 | [19 empty] | -1.2263 | [5 empty] | 3.4671 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031636 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.6076 | [19 empty] | -1.2263 | [5 empty] | 3.4672 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031638 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:13 | 20 | 20 |  | 1733.6076 | [19 empty] | -1.2263 | [5 empty] | 3.4672 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031638 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.6076 | [19 empty] | -1.2263 | [5 empty] | 3.4672 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031638 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.6076 | [19 empty] | -1.2262 | [5 empty] | 3.4672 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031638 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.6766 | [19 empty] | -1.2263 | [5 empty] | 3.4674 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031639 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.6766 | [19 empty] | -1.2262 | [5 empty] | 3.4674 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031639 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.7455 | [19 empty] | -1.2263 | [5 empty] | 3.4675 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003164 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.7455 | [19 empty] | -1.2262 | [5 empty] | 3.4675 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003164 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.7455 | [19 empty] | -1.2262 | [5 empty] | 3.4675 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003164 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.8145 | [19 empty] | -1.2263 | [5 empty] | 3.4676 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031641 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.8145 | [19 empty] | -1.2263 | [5 empty] | 3.4676 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031641 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.8835 | [19 empty] | -1.2263 | [5 empty] | 3.4678 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031643 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:14 | 20 | 20 |  | 1733.8835 | [19 empty] | -1.2263 | [5 empty] | 3.4678 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031643 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1733.8835 | [19 empty] | -1.2263 | [5 empty] | 3.4678 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031643 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1733.9525 | [19 empty] | -1.2263 | [5 empty] | 3.4679 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031644 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1733.9525 | [19 empty] | -1.2262 | [5 empty] | 3.4679 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031644 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.0214 | [19 empty] | -1.2262 | [5 empty] | 3.468 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031645 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.0214 | [19 empty] | -1.2263 | [5 empty] | 3.468 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031645 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.0214 | [19 empty] | -1.2262 | [5 empty] | 3.468 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031645 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.0904 | [19 empty] | -1.2262 | [5 empty] | 3.4682 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031646 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.0904 | [19 empty] | -1.2262 | [5 empty] | 3.4682 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031646 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.0904 | [19 empty] | -1.2262 | [5 empty] | 3.4682 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031646 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:15 | 20 | 20 |  | 1734.1594 | [19 empty] | -1.2262 | [5 empty] | 3.4683 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031648 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.1594 | [19 empty] | -1.2262 | [5 empty] | 3.4683 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031648 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.2284 | [19 empty] | -1.2262 | [5 empty] | 3.4685 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031649 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.2284 | [19 empty] | -1.2262 | [5 empty] | 3.4685 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031649 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.2284 | [19 empty] | -1.2262 | [5 empty] | 3.4685 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031649 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.2973 | [19 empty] | -1.2262 | [5 empty] | 3.4686 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003165 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.2973 | [19 empty] | -1.2262 | [5 empty] | 3.4686 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003165 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.2973 | [19 empty] | -1.2263 | [5 empty] | 3.4686 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003165 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.3663 | [19 empty] | -1.2263 | [5 empty] | 3.4687 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031651 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.3663 | [19 empty] | -1.2262 | [5 empty] | 3.4687 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031651 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:16 | 20 | 20 |  | 1734.3663 | [19 empty] | -1.2262 | [5 empty] | 3.4687 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031651 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.4353 | [19 empty] | -1.2262 | [5 empty] | 3.4689 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031653 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.4353 | [19 empty] | -1.2263 | [5 empty] | 3.4689 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031653 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.4353 | [19 empty] | -1.2262 | [5 empty] | 3.4689 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031653 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.4353 | [19 empty] | -1.2262 | [5 empty] | 3.4689 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031653 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.5043 | [19 empty] | -1.2262 | [5 empty] | 3.469 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031654 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.5043 | [19 empty] | -1.2262 | [5 empty] | 3.469 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031654 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.5732 | [19 empty] | -1.2263 | [5 empty] | 3.4691 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031655 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.5732 | [19 empty] | -1.2262 | [5 empty] | 3.4691 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031655 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.5732 | [19 empty] | -1.2263 | [5 empty] | 3.4691 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031655 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.6422 | [19 empty] | -1.2263 | [5 empty] | 3.4693 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031657 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:17 | 20 | 20 |  | 1734.6422 | [19 empty] | -1.2262 | [5 empty] | 3.4693 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031657 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.6422 | [19 empty] | -1.2263 | [5 empty] | 3.4693 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031657 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.6422 | [19 empty] | -1.2263 | [5 empty] | 3.4693 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031657 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.7112 | [19 empty] | -1.2263 | [5 empty] | 3.4694 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031658 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.7112 | [19 empty] | -1.2262 | [5 empty] | 3.4694 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031658 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.7802 | [19 empty] | -1.2263 | [5 empty] | 3.4696 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031659 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.7802 | [19 empty] | -1.2263 | [5 empty] | 3.4696 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031659 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.7802 | [19 empty] | -1.2263 | [5 empty] | 3.4696 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031659 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.8491 | [19 empty] | -1.2263 | [5 empty] | 3.4697 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003166 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.8491 | [19 empty] | -1.2263 | [5 empty] | 3.4697 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003166 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:18 | 20 | 20 |  | 1734.8491 | [19 empty] | -1.2263 | [5 empty] | 3.4697 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003166 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.8491 | [19 empty] | -1.2263 | [5 empty] | 3.4697 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.003166 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9181 | [19 empty] | -1.2263 | [5 empty] | 3.4698 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031662 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9181 | [19 empty] | -1.2263 | [5 empty] | 3.4698 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031662 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9181 | [19 empty] | -1.2262 | [5 empty] | 3.4698 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031662 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9871 | [19 empty] | -1.2263 | [5 empty] | 3.47 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031663 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9871 | [19 empty] | -1.2263 | [5 empty] | 3.47 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031663 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9871 | [19 empty] | -1.2263 | [5 empty] | 3.47 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031663 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1734.9871 | [19 empty] | -1.2262 | [5 empty] | 3.47 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031663 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1735.0561 | [19 empty] | -1.2263 | [5 empty] | 3.4701 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031664 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:19 | 20 | 20 |  | 1735.0561 | [19 empty] | -1.2263 | [5 empty] | 3.4701 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031664 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.0561 | [19 empty] | -1.2262 | [5 empty] | 3.4701 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031664 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.0561 | [19 empty] | -1.2262 | [5 empty] | 3.4701 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031664 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.125 | [19 empty] | -1.2262 | [5 empty] | 3.4703 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031665 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.125 | [19 empty] | -1.2263 | [5 empty] | 3.4703 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031665 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.125 | [19 empty] | -1.2262 | [5 empty] | 3.4703 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031665 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2262 | [5 empty] | 3.4704 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2263 | [5 empty] | 3.4704 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2262 | [5 empty] | 3.4704 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2262 | [5 empty] | 3.4704 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2262 | [5 empty] | 3.4704 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:20 | 20 | 20 |  | 1735.194 | [19 empty] | -1.2263 | [5 empty] | 3.4704 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031667 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.263 | [19 empty] | -1.2263 | [5 empty] | 3.4705 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031668 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.263 | [19 empty] | -1.2262 | [5 empty] | 3.4705 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031668 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.263 | [19 empty] | -1.2262 | [5 empty] | 3.4705 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031668 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.332 | [19 empty] | -1.2262 | [5 empty] | 3.4707 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031669 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.332 | [19 empty] | -1.2262 | [5 empty] | 3.4707 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031669 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.332 | [19 empty] | -1.2263 | [5 empty] | 3.4707 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031669 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.332 | [19 empty] | -1.2262 | [5 empty] | 3.4707 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031669 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.4009 | [19 empty] | -1.2262 | [5 empty] | 3.4708 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003167 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.4009 | [19 empty] | -1.2262 | [5 empty] | 3.4708 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003167 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:21 | 20 | 20 |  | 1735.4009 | [19 empty] | -1.2262 | [5 empty] | 3.4708 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003167 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4009 | [19 empty] | -1.2262 | [5 empty] | 3.4708 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003167 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.2263 | [5 empty] | 3.4709 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.2262 | [5 empty] | 3.4709 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.2263 | [5 empty] | 3.4709 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.2263 | [5 empty] | 3.4709 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.2263 | [5 empty] | 3.4709 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.4699 | [19 empty] | -1.2262 | [5 empty] | 3.4709 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031672 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.5389 | [19 empty] | -1.2263 | [5 empty] | 3.4711 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031673 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.5389 | [19 empty] | -1.2262 | [5 empty] | 3.4711 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031673 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:22 | 20 | 20 |  | 1735.5389 | [19 empty] | -1.2262 | [5 empty] | 3.4711 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031673 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6079 | [19 empty] | -1.2262 | [5 empty] | 3.4712 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031674 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6079 | [19 empty] | -1.2262 | [5 empty] | 3.4712 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031674 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6079 | [19 empty] | -1.2263 | [5 empty] | 3.4712 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031674 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6079 | [19 empty] | -1.2262 | [5 empty] | 3.4712 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031674 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.2262 | [5 empty] | 3.4714 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.2262 | [5 empty] | 3.4714 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.2263 | [5 empty] | 3.4714 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.2262 | [5 empty] | 3.4714 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.2263 | [5 empty] | 3.4714 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.6768 | [19 empty] | -1.2262 | [5 empty] | 3.4714 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031675 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:23 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.2262 | [5 empty] | 3.4715 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.2262 | [5 empty] | 3.4715 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.2263 | [5 empty] | 3.4715 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.2263 | [5 empty] | 3.4715 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.2262 | [5 empty] | 3.4715 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.7458 | [19 empty] | -1.2262 | [5 empty] | 3.4715 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031677 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.8148 | [19 empty] | -1.2263 | [5 empty] | 3.4716 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031678 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.8148 | [19 empty] | -1.2263 | [5 empty] | 3.4716 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031678 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.8148 | [19 empty] | -1.2263 | [5 empty] | 3.4716 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031678 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.8148 | [19 empty] | -1.2262 | [5 empty] | 3.4716 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031678 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:24 | 20 | 20 |  | 1735.8148 | [19 empty] | -1.2263 | [5 empty] | 3.4716 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031678 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.8838 | [19 empty] | -1.2263 | [5 empty] | 3.4718 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031679 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.8838 | [19 empty] | -1.2262 | [5 empty] | 3.4718 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031679 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.8838 | [19 empty] | -1.2262 | [5 empty] | 3.4718 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031679 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2263 | [5 empty] | 3.4719 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2262 | [5 empty] | 3.4719 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2262 | [5 empty] | 3.4719 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2262 | [5 empty] | 3.4719 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2262 | [5 empty] | 3.4719 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2263 | [5 empty] | 3.4719 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1735.9527 | [19 empty] | -1.2262 | [5 empty] | 3.4719 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031681 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:25 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4598 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0217 | [19 empty] | -1.2262 | [5 empty] | 3.472 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031682 | -1.4598 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.2262 | [5 empty] | 3.4722 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.2262 | [5 empty] | 3.4722 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:26 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.2262 | [5 empty] | 3.4722 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.2263 | [5 empty] | 3.4722 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.2262 | [5 empty] | 3.4722 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.0907 | [19 empty] | -1.2262 | [5 empty] | 3.4722 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031683 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.2262 | [5 empty] | 3.4723 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.2262 | [5 empty] | 3.4723 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.2262 | [5 empty] | 3.4723 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.2262 | [5 empty] | 3.4723 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.2262 | [5 empty] | 3.4723 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.1597 | [19 empty] | -1.2262 | [5 empty] | 3.4723 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031684 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:27 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2262 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2263 | [5 empty] | 3.4725 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2262 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2263 | [5 empty] | 3.4725 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2262 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2262 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2263 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2262 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2286 | [19 empty] | -1.2262 | [5 empty] | 3.4725 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031686 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.2262 | [5 empty] | 3.4726 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.2263 | [5 empty] | 3.4726 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:28 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.2262 | [5 empty] | 3.4726 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.2263 | [5 empty] | 3.4726 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.2262 | [5 empty] | 3.4726 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.2976 | [19 empty] | -1.2263 | [5 empty] | 3.4726 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031687 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2263 | [5 empty] | 3.4727 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2263 | [5 empty] | 3.4727 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2263 | [5 empty] | 3.4727 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2264 | [5 empty] | 3.4727 | -1.2255 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2263 | [5 empty] | 3.4727 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2264 | [5 empty] | 3.4727 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:29 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.3666 | [19 empty] | -1.2263 | [5 empty] | 3.4727 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031688 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.4356 | [19 empty] | -1.2263 | [5 empty] | 3.4729 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031689 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.2263 | [5 empty] | 3.473 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:30 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.2263 | [5 empty] | 3.473 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.2263 | [5 empty] | 3.473 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.2263 | [5 empty] | 3.473 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.2263 | [5 empty] | 3.473 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5045 | [19 empty] | -1.2263 | [5 empty] | 3.473 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031691 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:31 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2262 | [5 empty] | 3.4731 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.5735 | [19 empty] | -1.2263 | [5 empty] | 3.4731 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031692 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:32 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.6425 | [19 empty] | -1.2263 | [5 empty] | 3.4733 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031693 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:33 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7115 | [19 empty] | -1.2263 | [5 empty] | 3.4734 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031694 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:34 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.7804 | [19 empty] | -1.2263 | [5 empty] | 3.4736 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031696 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2264 | [5 empty] | 3.4737 | -1.2255 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:35 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.8494 | [19 empty] | -1.2263 | [5 empty] | 3.4737 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031697 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2262 | [5 empty] | 3.4738 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:36 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2262 | [5 empty] | 3.4738 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2262 | [5 empty] | 3.4738 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2263 | [5 empty] | 3.4738 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:37 | 20 | 20 |  | 1736.9184 | [19 empty] | -1.2262 | [5 empty] | 3.4738 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031698 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2262 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2262 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2262 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:38 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1736.9874 | [19 empty] | -1.2263 | [5 empty] | 3.474 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.00317 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:39 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.0563 | [19 empty] | -1.2263 | [5 empty] | 3.4741 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031701 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:40 | 20 | 20 |  | 1737.1253 | [19 empty] | -1.2263 | [5 empty] | 3.4743 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031702 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:41 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2262 | [5 empty] | 3.4744 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:42 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.1943 | [19 empty] | -1.2263 | [5 empty] | 3.4744 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031703 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2262 | [5 empty] | 3.4745 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:43 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:44 | 20 | 20 |  | 1737.2633 | [19 empty] | -1.2263 | [5 empty] | 3.4745 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031705 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2262 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:45 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2262 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2262 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:46 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.3322 | [19 empty] | -1.2263 | [5 empty] | 3.4747 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031706 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:47 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.4599 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:48 | 20 | 20 |  | 1737.4012 | [19 empty] | -1.2263 | [5 empty] | 3.4748 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031707 | -1.46 | -1.4407 | -1.4407 | -1.4407 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.46 | -1.4407 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:49 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.4702 | [19 empty] | -1.2263 | [5 empty] | 3.4749 | -1.2254 | 0.90444 | 17.6664 | 0.23773 | -0.0031708 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2262 | [5 empty] | 3.4751 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2262 | [5 empty] | 3.4751 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2263 | [5 empty] | 3.4751 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2262 | [5 empty] | 3.4751 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2263 | [5 empty] | 3.4751 | -1.2253 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4599 | -1.4406 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1737.5391 | [19 empty] | -1.2255 | [5 empty] | 3.4751 | -1.2246 | 0.90444 | 17.6664 | 0.23773 | -0.003171 | -1.4592 | -1.4399 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1539.1682 | [19 empty] | -1.2241 | [5 empty] | 3.0783 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0028068 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:50 | 20 | 20 |  | 1374.3189 | [19 empty] | -1.224 | [5 empty] | 2.7486 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025043 | -1.4584 | -1.4391 | -1.4391 | -1.4391 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1376.9399 | [19 empty] | -1.224 | [5 empty] | 2.7539 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025091 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1377.9746 | [19 empty] | -1.224 | [5 empty] | 2.7559 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002511 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1378.5264 | [19 empty] | -1.2241 | [5 empty] | 2.7571 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.002512 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1378.8712 | [19 empty] | -1.2241 | [5 empty] | 2.7577 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025126 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1379.2161 | [19 empty] | -1.2241 | [5 empty] | 2.7584 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025133 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1379.492 | [19 empty] | -1.2241 | [5 empty] | 2.759 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025138 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1379.7679 | [19 empty] | -1.2241 | [5 empty] | 2.7595 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025143 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1379.9748 | [19 empty] | -1.2241 | [5 empty] | 2.7599 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025146 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1380.1817 | [19 empty] | -1.2241 | [5 empty] | 2.7604 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.002515 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1380.3887 | [19 empty] | -1.2241 | [5 empty] | 2.7608 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025154 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:51 | 20 | 20 |  | 1380.5956 | [19 empty] | -1.2241 | [5 empty] | 2.7612 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025158 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1380.7335 | [19 empty] | -1.2241 | [5 empty] | 2.7615 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.002516 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1380.8715 | [19 empty] | -1.224 | [5 empty] | 2.7617 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025163 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.0784 | [19 empty] | -1.2241 | [5 empty] | 2.7622 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025167 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.2164 | [19 empty] | -1.224 | [5 empty] | 2.7624 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025169 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.3543 | [19 empty] | -1.2241 | [5 empty] | 2.7627 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025172 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.4923 | [19 empty] | -1.2241 | [5 empty] | 2.763 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025174 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.6302 | [19 empty] | -1.2241 | [5 empty] | 2.7633 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025177 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.7682 | [19 empty] | -1.2241 | [5 empty] | 2.7635 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025179 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.8371 | [19 empty] | -1.224 | [5 empty] | 2.7637 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025181 | -1.4583 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:52 | 20 | 20 |  | 1381.9751 | [19 empty] | -1.2241 | [5 empty] | 2.764 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025183 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1382.113 | [19 empty] | -1.2241 | [5 empty] | 2.7642 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025186 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1382.251 | [19 empty] | -1.2241 | [5 empty] | 2.7645 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025188 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2241 | [5 empty] | 2.7646 | -1.2232 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4584 | -1.4391 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.224 | [5 empty] | 2.7649 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1384.941 | [19 empty] | -1.224 | [5 empty] | 2.7699 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025238 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1375.8363 | [19 empty] | -1.2239 | [5 empty] | 2.7517 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002507 | -1.4583 | -1.439 | -1.439 | -1.439 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1375.9743 | [19 empty] | -1.224 | [5 empty] | 2.7519 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025073 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1376.0433 | [19 empty] | -1.2239 | [5 empty] | 2.7521 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025074 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1376.1812 | [19 empty] | -1.2239 | [5 empty] | 2.7524 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025077 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:53 | 20 | 20 |  | 1376.2502 | [19 empty] | -1.2239 | [5 empty] | 2.7525 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025078 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.3881 | [19 empty] | -1.2239 | [5 empty] | 2.7528 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025081 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.4571 | [19 empty] | -1.2239 | [5 empty] | 2.7529 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025082 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.5261 | [19 empty] | -1.224 | [5 empty] | 2.7531 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025083 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.664 | [19 empty] | -1.224 | [5 empty] | 2.7533 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025086 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.733 | [19 empty] | -1.2239 | [5 empty] | 2.7535 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025087 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.802 | [19 empty] | -1.224 | [5 empty] | 2.7536 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025088 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.871 | [19 empty] | -1.2239 | [5 empty] | 2.7537 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025089 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1376.9399 | [19 empty] | -1.224 | [5 empty] | 2.7539 | -1.2231 | 0.90444 | 17.6664 | 0.23773 | -0.0025091 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1377.0779 | [19 empty] | -1.2239 | [5 empty] | 2.7542 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025093 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1377.1469 | [19 empty] | -1.2239 | [5 empty] | 2.7543 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025095 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:54 | 20 | 20 |  | 1377.2158 | [19 empty] | -1.2239 | [5 empty] | 2.7544 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025096 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.2848 | [19 empty] | -1.2239 | [5 empty] | 2.7546 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025097 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.3538 | [19 empty] | -1.2239 | [5 empty] | 2.7547 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025098 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.4228 | [19 empty] | -1.2239 | [5 empty] | 2.7548 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.00251 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.4917 | [19 empty] | -1.2239 | [5 empty] | 2.755 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025101 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.5607 | [19 empty] | -1.2239 | [5 empty] | 2.7551 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025102 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.6297 | [19 empty] | -1.2239 | [5 empty] | 2.7553 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025103 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.6987 | [19 empty] | -1.2239 | [5 empty] | 2.7554 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025105 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.7676 | [19 empty] | -1.2239 | [5 empty] | 2.7555 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025106 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.8366 | [19 empty] | -1.2239 | [5 empty] | 2.7557 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025107 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:55 | 20 | 20 |  | 1377.9056 | [19 empty] | -1.2239 | [5 empty] | 2.7558 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025108 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1377.9746 | [19 empty] | -1.2239 | [5 empty] | 2.7559 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002511 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.0435 | [19 empty] | -1.2239 | [5 empty] | 2.7561 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025111 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.1125 | [19 empty] | -1.2239 | [5 empty] | 2.7562 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025112 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.1815 | [19 empty] | -1.224 | [5 empty] | 2.7564 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025114 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.2505 | [19 empty] | -1.2239 | [5 empty] | 2.7565 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025115 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.3194 | [19 empty] | -1.2239 | [5 empty] | 2.7566 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025116 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.3884 | [19 empty] | -1.2239 | [5 empty] | 2.7568 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025117 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.3884 | [19 empty] | -1.2239 | [5 empty] | 2.7568 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025117 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.4574 | [19 empty] | -1.2239 | [5 empty] | 2.7569 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025119 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.5264 | [19 empty] | -1.224 | [5 empty] | 2.7571 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002512 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:56 | 20 | 20 |  | 1378.5953 | [19 empty] | -1.2239 | [5 empty] | 2.7572 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025121 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1378.6643 | [19 empty] | -1.2239 | [5 empty] | 2.7573 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025122 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1378.7333 | [19 empty] | -1.2239 | [5 empty] | 2.7575 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025124 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1378.7333 | [19 empty] | -1.2239 | [5 empty] | 2.7575 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025124 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1378.8712 | [19 empty] | -1.2239 | [5 empty] | 2.7577 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025126 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1378.8712 | [19 empty] | -1.2239 | [5 empty] | 2.7577 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025126 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1378.9402 | [19 empty] | -1.2239 | [5 empty] | 2.7579 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025127 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1379.0092 | [19 empty] | -1.2239 | [5 empty] | 2.758 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025129 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1379.0782 | [19 empty] | -1.2239 | [5 empty] | 2.7582 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002513 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1379.0782 | [19 empty] | -1.2239 | [5 empty] | 2.7582 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002513 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:57 | 20 | 20 |  | 1379.1471 | [19 empty] | -1.2239 | [5 empty] | 2.7583 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025131 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.2161 | [19 empty] | -1.2239 | [5 empty] | 2.7584 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025133 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.2851 | [19 empty] | -1.2239 | [5 empty] | 2.7586 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025134 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.2851 | [19 empty] | -1.2239 | [5 empty] | 2.7586 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025134 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.3541 | [19 empty] | -1.2239 | [5 empty] | 2.7587 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025135 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.423 | [19 empty] | -1.2239 | [5 empty] | 2.7588 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025136 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.423 | [19 empty] | -1.2239 | [5 empty] | 2.7588 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025136 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.492 | [19 empty] | -1.2239 | [5 empty] | 2.759 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025138 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.561 | [19 empty] | -1.2239 | [5 empty] | 2.7591 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025139 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.6299 | [19 empty] | -1.2239 | [5 empty] | 2.7593 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002514 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:58 | 20 | 20 |  | 1379.6299 | [19 empty] | -1.2239 | [5 empty] | 2.7593 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002514 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1379.6989 | [19 empty] | -1.2239 | [5 empty] | 2.7594 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025141 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.2505 | [19 empty] | -1.2239 | [5 empty] | 2.7565 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025115 | -1.4582 | -1.4389 | -1.4389 | -1.4389 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.3194 | [19 empty] | -1.2239 | [5 empty] | 2.7566 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025116 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.3884 | [19 empty] | -1.2239 | [5 empty] | 2.7568 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025117 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.3884 | [19 empty] | -1.2239 | [5 empty] | 2.7568 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025117 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.4574 | [19 empty] | -1.2239 | [5 empty] | 2.7569 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025119 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.5264 | [19 empty] | -1.2239 | [5 empty] | 2.7571 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002512 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.5953 | [19 empty] | -1.2239 | [5 empty] | 2.7572 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025121 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.5953 | [19 empty] | -1.2239 | [5 empty] | 2.7572 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025121 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 10:59 | 20 | 20 |  | 1378.5953 | [19 empty] | -1.2239 | [5 empty] | 2.7572 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025121 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.6643 | [19 empty] | -1.2239 | [5 empty] | 2.7573 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025122 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.7333 | [19 empty] | -1.2239 | [5 empty] | 2.7575 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025124 | -1.4582 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.7333 | [19 empty] | -1.2239 | [5 empty] | 2.7575 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025124 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.8023 | [19 empty] | -1.2239 | [5 empty] | 2.7576 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025125 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.8712 | [19 empty] | -1.2239 | [5 empty] | 2.7577 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025126 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.8712 | [19 empty] | -1.2239 | [5 empty] | 2.7577 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025126 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.9402 | [19 empty] | -1.2239 | [5 empty] | 2.7579 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025127 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1378.9402 | [19 empty] | -1.2239 | [5 empty] | 2.7579 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025127 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1379.0092 | [19 empty] | -1.2239 | [5 empty] | 2.758 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025129 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1379.0782 | [19 empty] | -1.2239 | [5 empty] | 2.7582 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002513 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:00 | 20 | 20 |  | 1379.0782 | [19 empty] | -1.2239 | [5 empty] | 2.7582 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002513 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.1471 | [19 empty] | -1.2239 | [5 empty] | 2.7583 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025131 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.1471 | [19 empty] | -1.2239 | [5 empty] | 2.7583 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025131 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.2161 | [19 empty] | -1.2239 | [5 empty] | 2.7584 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025133 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.2161 | [19 empty] | -1.2238 | [5 empty] | 2.7584 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025133 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.2851 | [19 empty] | -1.2239 | [5 empty] | 2.7586 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025134 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.2851 | [19 empty] | -1.2239 | [5 empty] | 2.7586 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025134 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.3541 | [19 empty] | -1.2239 | [5 empty] | 2.7587 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025135 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.423 | [19 empty] | -1.2239 | [5 empty] | 2.7588 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025136 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.423 | [19 empty] | -1.2239 | [5 empty] | 2.7588 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025136 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:01 | 20 | 20 |  | 1379.423 | [19 empty] | -1.2239 | [5 empty] | 2.7588 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025136 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.492 | [19 empty] | -1.2239 | [5 empty] | 2.759 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025138 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.561 | [19 empty] | -1.2239 | [5 empty] | 2.7591 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025139 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.561 | [19 empty] | -1.2238 | [5 empty] | 2.7591 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025139 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.6299 | [19 empty] | -1.2239 | [5 empty] | 2.7593 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002514 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.6299 | [19 empty] | -1.2239 | [5 empty] | 2.7593 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002514 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.6989 | [19 empty] | -1.2239 | [5 empty] | 2.7594 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025141 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.6989 | [19 empty] | -1.2239 | [5 empty] | 2.7594 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025141 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.7679 | [19 empty] | -1.2239 | [5 empty] | 2.7595 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025143 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.7679 | [19 empty] | -1.2239 | [5 empty] | 2.7595 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025143 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.7679 | [19 empty] | -1.2239 | [5 empty] | 2.7595 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025143 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:02 | 20 | 20 |  | 1379.8369 | [19 empty] | -1.2239 | [5 empty] | 2.7597 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025144 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1379.9058 | [19 empty] | -1.2239 | [5 empty] | 2.7598 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025145 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1379.9058 | [19 empty] | -1.2239 | [5 empty] | 2.7598 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025145 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1379.9058 | [19 empty] | -1.2239 | [5 empty] | 2.7598 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025145 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1379.9748 | [19 empty] | -1.2239 | [5 empty] | 2.7599 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025146 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1380.0438 | [19 empty] | -1.2239 | [5 empty] | 2.7601 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025148 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1380.0438 | [19 empty] | -1.2239 | [5 empty] | 2.7601 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025148 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1380.0438 | [19 empty] | -1.2239 | [5 empty] | 2.7601 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025148 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1380.1128 | [19 empty] | -1.2239 | [5 empty] | 2.7602 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025149 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1380.1128 | [19 empty] | -1.224 | [5 empty] | 2.7602 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025149 | -1.4583 | -1.439 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:03 | 20 | 20 |  | 1380.1128 | [19 empty] | -1.2239 | [5 empty] | 2.7602 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025149 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.1817 | [19 empty] | -1.2239 | [5 empty] | 2.7604 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002515 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.1817 | [19 empty] | -1.2239 | [5 empty] | 2.7604 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002515 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.2507 | [19 empty] | -1.2239 | [5 empty] | 2.7605 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025152 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.2507 | [19 empty] | -1.2239 | [5 empty] | 2.7605 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025152 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.3197 | [19 empty] | -1.2239 | [5 empty] | 2.7606 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025153 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.3197 | [19 empty] | -1.2239 | [5 empty] | 2.7606 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025153 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.3197 | [19 empty] | -1.2239 | [5 empty] | 2.7606 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025153 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.3887 | [19 empty] | -1.2238 | [5 empty] | 2.7608 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025154 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.3887 | [19 empty] | -1.2239 | [5 empty] | 2.7608 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025154 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:04 | 20 | 20 |  | 1380.4576 | [19 empty] | -1.2239 | [5 empty] | 2.7609 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025155 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.4576 | [19 empty] | -1.2239 | [5 empty] | 2.7609 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025155 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.4576 | [19 empty] | -1.2239 | [5 empty] | 2.7609 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025155 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.5266 | [19 empty] | -1.2239 | [5 empty] | 2.7611 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025157 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.5266 | [19 empty] | -1.2238 | [5 empty] | 2.7611 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025157 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.5956 | [19 empty] | -1.2238 | [5 empty] | 2.7612 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025158 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.5956 | [19 empty] | -1.2239 | [5 empty] | 2.7612 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025158 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.5956 | [19 empty] | -1.2238 | [5 empty] | 2.7612 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025158 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.6646 | [19 empty] | -1.2239 | [5 empty] | 2.7613 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025159 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.6646 | [19 empty] | -1.2239 | [5 empty] | 2.7613 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025159 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.6646 | [19 empty] | -1.2239 | [5 empty] | 2.7613 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025159 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:05 | 20 | 20 |  | 1380.7335 | [19 empty] | -1.2239 | [5 empty] | 2.7615 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002516 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.7335 | [19 empty] | -1.2239 | [5 empty] | 2.7615 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002516 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.7335 | [19 empty] | -1.2239 | [5 empty] | 2.7615 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002516 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.8025 | [19 empty] | -1.2239 | [5 empty] | 2.7616 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025162 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.8025 | [19 empty] | -1.2238 | [5 empty] | 2.7616 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025162 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.8715 | [19 empty] | -1.2238 | [5 empty] | 2.7617 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025163 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.8715 | [19 empty] | -1.2238 | [5 empty] | 2.7617 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025163 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.8715 | [19 empty] | -1.2238 | [5 empty] | 2.7617 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025163 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.9405 | [19 empty] | -1.2239 | [5 empty] | 2.7619 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025164 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.9405 | [19 empty] | -1.2239 | [5 empty] | 2.7619 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025164 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:06 | 20 | 20 |  | 1380.9405 | [19 empty] | -1.2238 | [5 empty] | 2.7619 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025164 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0094 | [19 empty] | -1.2239 | [5 empty] | 2.762 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025165 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0094 | [19 empty] | -1.2239 | [5 empty] | 2.762 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025165 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0094 | [19 empty] | -1.2239 | [5 empty] | 2.762 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025165 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0784 | [19 empty] | -1.2238 | [5 empty] | 2.7622 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025167 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0784 | [19 empty] | -1.2238 | [5 empty] | 2.7622 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025167 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0784 | [19 empty] | -1.2238 | [5 empty] | 2.7622 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025167 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.0784 | [19 empty] | -1.2238 | [5 empty] | 2.7622 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025167 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.1474 | [19 empty] | -1.2238 | [5 empty] | 2.7623 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025168 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.1474 | [19 empty] | -1.2239 | [5 empty] | 2.7623 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025168 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:07 | 20 | 20 |  | 1381.1474 | [19 empty] | -1.2239 | [5 empty] | 2.7623 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025168 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.1474 | [19 empty] | -1.2239 | [5 empty] | 2.7623 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025168 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.2164 | [19 empty] | -1.2239 | [5 empty] | 2.7624 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025169 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.2164 | [19 empty] | -1.2238 | [5 empty] | 2.7624 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025169 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.2853 | [19 empty] | -1.2239 | [5 empty] | 2.7626 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002517 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.2853 | [19 empty] | -1.2239 | [5 empty] | 2.7626 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.002517 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.2853 | [19 empty] | -1.2238 | [5 empty] | 2.7626 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002517 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.2853 | [19 empty] | -1.2238 | [5 empty] | 2.7626 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.002517 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.3543 | [19 empty] | -1.2238 | [5 empty] | 2.7627 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025172 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.3543 | [19 empty] | -1.2239 | [5 empty] | 2.7627 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025172 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.3543 | [19 empty] | -1.2238 | [5 empty] | 2.7627 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025172 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:08 | 20 | 20 |  | 1381.3543 | [19 empty] | -1.2238 | [5 empty] | 2.7627 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025172 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4233 | [19 empty] | -1.2239 | [5 empty] | 2.7628 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025173 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4233 | [19 empty] | -1.2239 | [5 empty] | 2.7628 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025173 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4233 | [19 empty] | -1.2238 | [5 empty] | 2.7628 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025173 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4233 | [19 empty] | -1.2239 | [5 empty] | 2.7628 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025173 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4923 | [19 empty] | -1.2239 | [5 empty] | 2.763 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025174 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4923 | [19 empty] | -1.2239 | [5 empty] | 2.763 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025174 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4923 | [19 empty] | -1.2239 | [5 empty] | 2.763 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025174 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4923 | [19 empty] | -1.2239 | [5 empty] | 2.763 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025174 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.4923 | [19 empty] | -1.2239 | [5 empty] | 2.763 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025174 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:09 | 20 | 20 |  | 1381.5612 | [19 empty] | -1.2239 | [5 empty] | 2.7631 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025176 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.5612 | [19 empty] | -1.2239 | [5 empty] | 2.7631 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025176 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.5612 | [19 empty] | -1.2239 | [5 empty] | 2.7631 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025176 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6302 | [19 empty] | -1.2239 | [5 empty] | 2.7633 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025177 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6302 | [19 empty] | -1.2239 | [5 empty] | 2.7633 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025177 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6302 | [19 empty] | -1.2239 | [5 empty] | 2.7633 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025177 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6302 | [19 empty] | -1.2239 | [5 empty] | 2.7633 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025177 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6992 | [19 empty] | -1.2239 | [5 empty] | 2.7634 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025178 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6992 | [19 empty] | -1.2239 | [5 empty] | 2.7634 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025178 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6992 | [19 empty] | -1.2239 | [5 empty] | 2.7634 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025178 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.6992 | [19 empty] | -1.2238 | [5 empty] | 2.7634 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025178 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:10 | 20 | 20 |  | 1381.7682 | [19 empty] | -1.2239 | [5 empty] | 2.7635 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025179 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.7682 | [19 empty] | -1.2239 | [5 empty] | 2.7635 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025179 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.7682 | [19 empty] | -1.2239 | [5 empty] | 2.7635 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025179 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.7682 | [19 empty] | -1.2239 | [5 empty] | 2.7635 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025179 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.8371 | [19 empty] | -1.2239 | [5 empty] | 2.7637 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025181 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.8371 | [19 empty] | -1.2239 | [5 empty] | 2.7637 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025181 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.8371 | [19 empty] | -1.2239 | [5 empty] | 2.7637 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025181 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.8371 | [19 empty] | -1.2238 | [5 empty] | 2.7637 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025181 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.8371 | [19 empty] | -1.2239 | [5 empty] | 2.7637 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025181 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.9061 | [19 empty] | -1.2239 | [5 empty] | 2.7638 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025182 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:11 | 20 | 20 |  | 1381.9061 | [19 empty] | -1.2238 | [5 empty] | 2.7638 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025182 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9061 | [19 empty] | -1.2238 | [5 empty] | 2.7638 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025182 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9061 | [19 empty] | -1.2238 | [5 empty] | 2.7638 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025182 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9751 | [19 empty] | -1.2238 | [5 empty] | 2.764 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025183 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9751 | [19 empty] | -1.2239 | [5 empty] | 2.764 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025183 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9751 | [19 empty] | -1.2238 | [5 empty] | 2.764 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025183 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9751 | [19 empty] | -1.2239 | [5 empty] | 2.764 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025183 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1381.9751 | [19 empty] | -1.2239 | [5 empty] | 2.764 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025183 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1382.0441 | [19 empty] | -1.2238 | [5 empty] | 2.7641 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025184 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1382.0441 | [19 empty] | -1.2238 | [5 empty] | 2.7641 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025184 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:12 | 20 | 20 |  | 1382.0441 | [19 empty] | -1.2239 | [5 empty] | 2.7641 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025184 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.0441 | [19 empty] | -1.2239 | [5 empty] | 2.7641 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025184 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.0441 | [19 empty] | -1.2238 | [5 empty] | 2.7641 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025184 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.113 | [19 empty] | -1.2238 | [5 empty] | 2.7642 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025186 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.113 | [19 empty] | -1.2238 | [5 empty] | 2.7642 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025186 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.113 | [19 empty] | -1.2238 | [5 empty] | 2.7642 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025186 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.113 | [19 empty] | -1.2239 | [5 empty] | 2.7642 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025186 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.182 | [19 empty] | -1.2239 | [5 empty] | 2.7644 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025187 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.182 | [19 empty] | -1.2239 | [5 empty] | 2.7644 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025187 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.182 | [19 empty] | -1.2239 | [5 empty] | 2.7644 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025187 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.182 | [19 empty] | -1.2238 | [5 empty] | 2.7644 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025187 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:13 | 20 | 20 |  | 1382.182 | [19 empty] | -1.2239 | [5 empty] | 2.7644 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025187 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.251 | [19 empty] | -1.2239 | [5 empty] | 2.7645 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025188 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.251 | [19 empty] | -1.2239 | [5 empty] | 2.7645 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025188 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.251 | [19 empty] | -1.2238 | [5 empty] | 2.7645 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025188 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.251 | [19 empty] | -1.2238 | [5 empty] | 2.7645 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025188 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2239 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2239 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2239 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2239 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2238 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:14 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2238 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.32 | [19 empty] | -1.2238 | [5 empty] | 2.7646 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025189 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2238 | [5 empty] | 2.7648 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2238 | [5 empty] | 2.7648 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2238 | [5 empty] | 2.7648 | -1.2228 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2239 | [5 empty] | 2.7648 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2239 | [5 empty] | 2.7648 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2238 | [5 empty] | 2.7648 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.3889 | [19 empty] | -1.2238 | [5 empty] | 2.7648 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025191 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.2238 | [5 empty] | 2.7649 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:15 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.2238 | [5 empty] | 2.7649 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.2238 | [5 empty] | 2.7649 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.2238 | [5 empty] | 2.7649 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.2238 | [5 empty] | 2.7649 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.4579 | [19 empty] | -1.2238 | [5 empty] | 2.7649 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025192 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2238 | [5 empty] | 2.7651 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2238 | [5 empty] | 2.7651 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2239 | [5 empty] | 2.7651 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2239 | [5 empty] | 2.7651 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2239 | [5 empty] | 2.7651 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:16 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2238 | [5 empty] | 2.7651 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.5269 | [19 empty] | -1.2239 | [5 empty] | 2.7651 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025193 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.5959 | [19 empty] | -1.2239 | [5 empty] | 2.7652 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025195 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.5959 | [19 empty] | -1.2239 | [5 empty] | 2.7652 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025195 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.5959 | [19 empty] | -1.2239 | [5 empty] | 2.7652 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025195 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.5959 | [19 empty] | -1.2239 | [5 empty] | 2.7652 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025195 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.5959 | [19 empty] | -1.2239 | [5 empty] | 2.7652 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025195 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2239 | [5 empty] | 2.7653 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2239 | [5 empty] | 2.7653 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2239 | [5 empty] | 2.7653 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2238 | [5 empty] | 2.7653 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:17 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2239 | [5 empty] | 2.7653 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2239 | [5 empty] | 2.7653 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.6648 | [19 empty] | -1.2239 | [5 empty] | 2.7653 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025196 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.7338 | [19 empty] | -1.2239 | [5 empty] | 2.7655 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025197 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.7338 | [19 empty] | -1.2238 | [5 empty] | 2.7655 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025197 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.7338 | [19 empty] | -1.2238 | [5 empty] | 2.7655 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025197 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.7338 | [19 empty] | -1.2239 | [5 empty] | 2.7655 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025197 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.7338 | [19 empty] | -1.2238 | [5 empty] | 2.7655 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025197 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.7338 | [19 empty] | -1.2239 | [5 empty] | 2.7655 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025197 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.8028 | [19 empty] | -1.2239 | [5 empty] | 2.7656 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025198 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:18 | 20 | 20 |  | 1382.8028 | [19 empty] | -1.2239 | [5 empty] | 2.7656 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025198 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8028 | [19 empty] | -1.2239 | [5 empty] | 2.7656 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025198 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8028 | [19 empty] | -1.2239 | [5 empty] | 2.7656 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025198 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8028 | [19 empty] | -1.2239 | [5 empty] | 2.7656 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025198 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8028 | [19 empty] | -1.2238 | [5 empty] | 2.7656 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025198 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2239 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2239 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2238 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2239 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2239 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2238 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:19 | 20 | 20 |  | 1382.8718 | [19 empty] | -1.2239 | [5 empty] | 2.7657 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.00252 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1382.9407 | [19 empty] | -1.2239 | [5 empty] | 2.7659 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025201 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1382.9407 | [19 empty] | -1.2239 | [5 empty] | 2.7659 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025201 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1382.9407 | [19 empty] | -1.2239 | [5 empty] | 2.7659 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025201 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1382.9407 | [19 empty] | -1.2239 | [5 empty] | 2.7659 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025201 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:20 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0097 | [19 empty] | -1.2239 | [5 empty] | 2.766 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025202 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2239 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2238 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2239 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2239 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2239 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2239 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:21 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2238 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2238 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.1477 | [19 empty] | -1.2239 | [5 empty] | 2.7663 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025205 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.0787 | [19 empty] | -1.2238 | [5 empty] | 2.7662 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025203 | -1.4581 | -1.4388 | -1.4388 | -1.4388 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.1477 | [19 empty] | -1.2238 | [5 empty] | 2.7663 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025205 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.1477 | [19 empty] | -1.2239 | [5 empty] | 2.7663 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025205 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.1477 | [19 empty] | -1.2239 | [5 empty] | 2.7663 | -1.223 | 0.90444 | 17.6664 | 0.23773 | -0.0025205 | -1.4582 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.1477 | [19 empty] | -1.2239 | [5 empty] | 2.7663 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025205 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.2166 | [19 empty] | -1.2238 | [5 empty] | 2.7664 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025206 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.2166 | [19 empty] | -1.2239 | [5 empty] | 2.7664 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025206 | -1.4581 | -1.4389 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.2166 | [19 empty] | -1.2238 | [5 empty] | 2.7664 | -1.2229 | 0.90444 | 17.6664 | 0.23773 | -0.0025206 | -1.4581 | -1.4388 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:22 | 20 | 20 |  | 1383.2166 | [19 empty] | -1.2237 | [5 empty] | 2.7664 | -1.2228 | 0.90444 | 17.6664 | 0.23773 | -0.0025206 | -1.458 | -1.4387 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1330.52 | [19 empty] | -1.2212 | [5 empty] | 2.661 | -1.2203 | 0.90444 | 17.6664 | 0.23773 | -0.0024239 | -1.4556 | -1.4363 | -1.4363 | -1.4363 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1063.8642 | [19 empty] | -1.2207 | [5 empty] | 2.1277 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0019344 | -1.4555 | -1.4362 | -1.4362 | -1.4362 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1027.3767 | [19 empty] | -1.2207 | [5 empty] | 2.0548 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018674 | -1.4557 | -1.4364 | -1.4364 | -1.4364 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1028.7562 | [19 empty] | -1.2208 | [5 empty] | 2.0575 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00187 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1029.308 | [19 empty] | -1.2208 | [5 empty] | 2.0586 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001871 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1029.7908 | [19 empty] | -1.2208 | [5 empty] | 2.0596 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018719 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1030.1357 | [19 empty] | -1.2208 | [5 empty] | 2.0603 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018725 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1030.4805 | [19 empty] | -1.2208 | [5 empty] | 2.061 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018731 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1030.7564 | [19 empty] | -1.2208 | [5 empty] | 2.0615 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018736 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:23 | 20 | 20 |  | 1030.9633 | [19 empty] | -1.2208 | [5 empty] | 2.0619 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001874 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1031.2392 | [19 empty] | -1.2208 | [5 empty] | 2.0625 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018745 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1031.3772 | [19 empty] | -1.2208 | [5 empty] | 2.0628 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018748 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1031.5841 | [19 empty] | -1.2208 | [5 empty] | 2.0632 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018752 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1031.7221 | [19 empty] | -1.2208 | [5 empty] | 2.0634 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018754 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1031.929 | [19 empty] | -1.2208 | [5 empty] | 2.0639 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018758 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1032.0669 | [19 empty] | -1.2208 | [5 empty] | 2.0641 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1032.2739 | [19 empty] | -1.2208 | [5 empty] | 2.0645 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018764 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1032.4118 | [19 empty] | -1.2208 | [5 empty] | 2.0648 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018767 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1032.5498 | [19 empty] | -1.2208 | [5 empty] | 2.0651 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018769 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:24 | 20 | 20 |  | 1032.6187 | [19 empty] | -1.2208 | [5 empty] | 2.0652 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018771 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1032.7567 | [19 empty] | -1.2208 | [5 empty] | 2.0655 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018773 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1032.8946 | [19 empty] | -1.2208 | [5 empty] | 2.0658 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018776 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.0326 | [19 empty] | -1.2209 | [5 empty] | 2.0661 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018778 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.1016 | [19 empty] | -1.2208 | [5 empty] | 2.0662 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018779 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.2395 | [19 empty] | -1.2208 | [5 empty] | 2.0665 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018782 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.3775 | [19 empty] | -1.2209 | [5 empty] | 2.0668 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018785 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.4464 | [19 empty] | -1.2208 | [5 empty] | 2.0669 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018786 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.5844 | [19 empty] | -1.2208 | [5 empty] | 2.0672 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018788 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.6534 | [19 empty] | -1.2208 | [5 empty] | 2.0673 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001879 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.7223 | [19 empty] | -1.2208 | [5 empty] | 2.0674 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018791 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:25 | 20 | 20 |  | 1033.8603 | [19 empty] | -1.2209 | [5 empty] | 2.0677 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018793 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1033.9293 | [19 empty] | -1.2209 | [5 empty] | 2.0679 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018795 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.0672 | [19 empty] | -1.2209 | [5 empty] | 2.0681 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018797 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.1362 | [19 empty] | -1.2208 | [5 empty] | 2.0683 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018798 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.2741 | [19 empty] | -1.2208 | [5 empty] | 2.0685 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018801 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.3431 | [19 empty] | -1.2209 | [5 empty] | 2.0687 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018802 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.4121 | [19 empty] | -1.2209 | [5 empty] | 2.0688 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018803 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.4811 | [19 empty] | -1.2208 | [5 empty] | 2.069 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018805 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.55 | [19 empty] | -1.2208 | [5 empty] | 2.0691 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018806 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.619 | [19 empty] | -1.2208 | [5 empty] | 2.0692 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018807 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:26 | 20 | 20 |  | 1034.619 | [19 empty] | -1.2208 | [5 empty] | 2.0692 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018807 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1034.7569 | [19 empty] | -1.2208 | [5 empty] | 2.0695 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001881 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1034.8259 | [19 empty] | -1.2209 | [5 empty] | 2.0697 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018811 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1034.8949 | [19 empty] | -1.2209 | [5 empty] | 2.0698 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018812 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1034.9639 | [19 empty] | -1.2209 | [5 empty] | 2.0699 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018814 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1035.1018 | [19 empty] | -1.2209 | [5 empty] | 2.0702 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018816 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1035.1708 | [19 empty] | -1.2209 | [5 empty] | 2.0703 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018817 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1035.2398 | [19 empty] | -1.2208 | [5 empty] | 2.0705 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018819 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1035.3087 | [19 empty] | -1.2208 | [5 empty] | 2.0706 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1035.3777 | [19 empty] | -1.2208 | [5 empty] | 2.0708 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018821 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:27 | 20 | 20 |  | 1035.4467 | [19 empty] | -1.2207 | [5 empty] | 2.0709 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018822 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.5157 | [19 empty] | -1.2208 | [5 empty] | 2.071 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018824 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.5846 | [19 empty] | -1.2208 | [5 empty] | 2.0712 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018825 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.6536 | [19 empty] | -1.2208 | [5 empty] | 2.0713 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018826 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.7226 | [19 empty] | -1.2208 | [5 empty] | 2.0714 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018828 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.7916 | [19 empty] | -1.2208 | [5 empty] | 2.0716 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018829 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.8605 | [19 empty] | -1.2208 | [5 empty] | 2.0717 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.9295 | [19 empty] | -1.2208 | [5 empty] | 2.0719 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018831 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1035.9985 | [19 empty] | -1.2208 | [5 empty] | 2.072 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018833 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1036.0675 | [19 empty] | -1.2208 | [5 empty] | 2.0721 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018834 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1036.1364 | [19 empty] | -1.2208 | [5 empty] | 2.0723 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018835 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:28 | 20 | 20 |  | 1036.2054 | [19 empty] | -1.2208 | [5 empty] | 2.0724 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018836 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1036.2054 | [19 empty] | -1.2208 | [5 empty] | 2.0724 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018836 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1036.3434 | [19 empty] | -1.2207 | [5 empty] | 2.0727 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018839 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.0672 | [19 empty] | -1.2208 | [5 empty] | 2.0681 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018797 | -1.4557 | -1.4364 | -1.4364 | -1.4364 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.1362 | [19 empty] | -1.2208 | [5 empty] | 2.0683 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018798 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.2052 | [19 empty] | -1.2208 | [5 empty] | 2.0684 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00188 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.2741 | [19 empty] | -1.2208 | [5 empty] | 2.0685 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018801 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.3431 | [19 empty] | -1.2208 | [5 empty] | 2.0687 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018802 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.4121 | [19 empty] | -1.2208 | [5 empty] | 2.0688 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018803 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.4121 | [19 empty] | -1.2208 | [5 empty] | 2.0688 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018803 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:29 | 20 | 20 |  | 1034.4121 | [19 empty] | -1.2208 | [5 empty] | 2.0688 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018803 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.4811 | [19 empty] | -1.2208 | [5 empty] | 2.069 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018805 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.55 | [19 empty] | -1.2208 | [5 empty] | 2.0691 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018806 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.55 | [19 empty] | -1.2208 | [5 empty] | 2.0691 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018806 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.619 | [19 empty] | -1.2208 | [5 empty] | 2.0692 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018807 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.619 | [19 empty] | -1.2208 | [5 empty] | 2.0692 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018807 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.688 | [19 empty] | -1.2208 | [5 empty] | 2.0694 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018809 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.7569 | [19 empty] | -1.2208 | [5 empty] | 2.0695 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001881 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.8259 | [19 empty] | -1.2208 | [5 empty] | 2.0697 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018811 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.8259 | [19 empty] | -1.2208 | [5 empty] | 2.0697 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018811 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.8949 | [19 empty] | -1.2208 | [5 empty] | 2.0698 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018812 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:30 | 20 | 20 |  | 1034.9639 | [19 empty] | -1.2208 | [5 empty] | 2.0699 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018814 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1034.9639 | [19 empty] | -1.2208 | [5 empty] | 2.0699 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018814 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.0328 | [19 empty] | -1.2209 | [5 empty] | 2.0701 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018815 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.1018 | [19 empty] | -1.2208 | [5 empty] | 2.0702 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018816 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.1018 | [19 empty] | -1.2208 | [5 empty] | 2.0702 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018816 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.1708 | [19 empty] | -1.2208 | [5 empty] | 2.0703 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018817 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.2398 | [19 empty] | -1.2208 | [5 empty] | 2.0705 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018819 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.2398 | [19 empty] | -1.2208 | [5 empty] | 2.0705 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018819 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.3087 | [19 empty] | -1.2208 | [5 empty] | 2.0706 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.3777 | [19 empty] | -1.2208 | [5 empty] | 2.0708 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018821 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:31 | 20 | 20 |  | 1035.3777 | [19 empty] | -1.2208 | [5 empty] | 2.0708 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018821 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.4467 | [19 empty] | -1.2208 | [5 empty] | 2.0709 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018822 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.4467 | [19 empty] | -1.2208 | [5 empty] | 2.0709 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018822 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.5157 | [19 empty] | -1.2208 | [5 empty] | 2.071 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018824 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.5157 | [19 empty] | -1.2208 | [5 empty] | 2.071 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018824 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.5846 | [19 empty] | -1.2208 | [5 empty] | 2.0712 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018825 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.5846 | [19 empty] | -1.2208 | [5 empty] | 2.0712 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018825 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.6536 | [19 empty] | -1.2208 | [5 empty] | 2.0713 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018826 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.7226 | [19 empty] | -1.2208 | [5 empty] | 2.0714 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018828 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.7226 | [19 empty] | -1.2208 | [5 empty] | 2.0714 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018828 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:32 | 20 | 20 |  | 1035.7916 | [19 empty] | -1.2208 | [5 empty] | 2.0716 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018829 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1035.8605 | [19 empty] | -1.2208 | [5 empty] | 2.0717 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1035.8605 | [19 empty] | -1.2208 | [5 empty] | 2.0717 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1035.9295 | [19 empty] | -1.2208 | [5 empty] | 2.0719 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018831 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1035.9295 | [19 empty] | -1.2208 | [5 empty] | 2.0719 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018831 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1035.9985 | [19 empty] | -1.2208 | [5 empty] | 2.072 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018833 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1035.9985 | [19 empty] | -1.2208 | [5 empty] | 2.072 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018833 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1036.0675 | [19 empty] | -1.2208 | [5 empty] | 2.0721 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018834 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1036.0675 | [19 empty] | -1.2208 | [5 empty] | 2.0721 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018834 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1036.1364 | [19 empty] | -1.2208 | [5 empty] | 2.0723 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018835 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1036.1364 | [19 empty] | -1.2208 | [5 empty] | 2.0723 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018835 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:33 | 20 | 20 |  | 1036.2054 | [19 empty] | -1.2208 | [5 empty] | 2.0724 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018836 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.2054 | [19 empty] | -1.2208 | [5 empty] | 2.0724 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018836 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.2744 | [19 empty] | -1.2208 | [5 empty] | 2.0725 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018838 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.3434 | [19 empty] | -1.2208 | [5 empty] | 2.0727 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018839 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.3434 | [19 empty] | -1.2208 | [5 empty] | 2.0727 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018839 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.3434 | [19 empty] | -1.2207 | [5 empty] | 2.0727 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018839 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.4123 | [19 empty] | -1.2208 | [5 empty] | 2.0728 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001884 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.4123 | [19 empty] | -1.2207 | [5 empty] | 2.0728 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001884 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.4813 | [19 empty] | -1.2207 | [5 empty] | 2.073 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018841 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.4813 | [19 empty] | -1.2208 | [5 empty] | 2.073 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018841 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:34 | 20 | 20 |  | 1036.5503 | [19 empty] | -1.2208 | [5 empty] | 2.0731 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018843 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.5503 | [19 empty] | -1.2208 | [5 empty] | 2.0731 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018843 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.6193 | [19 empty] | -1.2208 | [5 empty] | 2.0732 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018844 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.6193 | [19 empty] | -1.2208 | [5 empty] | 2.0732 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018844 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.6882 | [19 empty] | -1.2208 | [5 empty] | 2.0734 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018845 | -1.4558 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.6882 | [19 empty] | -1.2208 | [5 empty] | 2.0734 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018845 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.6882 | [19 empty] | -1.2208 | [5 empty] | 2.0734 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018845 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.6882 | [19 empty] | -1.2208 | [5 empty] | 2.0734 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018845 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.7572 | [19 empty] | -1.2207 | [5 empty] | 2.0735 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018847 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.8262 | [19 empty] | -1.2208 | [5 empty] | 2.0737 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018848 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:35 | 20 | 20 |  | 1036.8262 | [19 empty] | -1.2207 | [5 empty] | 2.0737 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018848 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1036.8262 | [19 empty] | -1.2207 | [5 empty] | 2.0737 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018848 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1036.8952 | [19 empty] | -1.2208 | [5 empty] | 2.0738 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018849 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1036.8952 | [19 empty] | -1.2208 | [5 empty] | 2.0738 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018849 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1036.9641 | [19 empty] | -1.2208 | [5 empty] | 2.0739 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1036.9641 | [19 empty] | -1.2208 | [5 empty] | 2.0739 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1036.9641 | [19 empty] | -1.2208 | [5 empty] | 2.0739 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1037.0331 | [19 empty] | -1.2208 | [5 empty] | 2.0741 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018852 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1037.0331 | [19 empty] | -1.2208 | [5 empty] | 2.0741 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018852 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1037.1021 | [19 empty] | -1.2208 | [5 empty] | 2.0742 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018853 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1037.1021 | [19 empty] | -1.2208 | [5 empty] | 2.0742 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018853 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:36 | 20 | 20 |  | 1037.1021 | [19 empty] | -1.2208 | [5 empty] | 2.0742 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018853 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.1021 | [19 empty] | -1.2208 | [5 empty] | 2.0742 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018853 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.1711 | [19 empty] | -1.2207 | [5 empty] | 2.0743 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018854 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.1711 | [19 empty] | -1.2207 | [5 empty] | 2.0743 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018854 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.1711 | [19 empty] | -1.2208 | [5 empty] | 2.0743 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018854 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.24 | [19 empty] | -1.2208 | [5 empty] | 2.0745 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018855 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.24 | [19 empty] | -1.2207 | [5 empty] | 2.0745 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018855 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.309 | [19 empty] | -1.2207 | [5 empty] | 2.0746 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018857 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.309 | [19 empty] | -1.2208 | [5 empty] | 2.0746 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018857 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.309 | [19 empty] | -1.2208 | [5 empty] | 2.0746 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018857 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:37 | 20 | 20 |  | 1037.309 | [19 empty] | -1.2208 | [5 empty] | 2.0746 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018857 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.378 | [19 empty] | -1.2207 | [5 empty] | 2.0748 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018858 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.378 | [19 empty] | -1.2207 | [5 empty] | 2.0748 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018858 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.378 | [19 empty] | -1.2208 | [5 empty] | 2.0748 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018858 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.447 | [19 empty] | -1.2208 | [5 empty] | 2.0749 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018859 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.447 | [19 empty] | -1.2208 | [5 empty] | 2.0749 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018859 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.447 | [19 empty] | -1.2207 | [5 empty] | 2.0749 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018859 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.5159 | [19 empty] | -1.2208 | [5 empty] | 2.075 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.5159 | [19 empty] | -1.2207 | [5 empty] | 2.075 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001886 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.5159 | [19 empty] | -1.2208 | [5 empty] | 2.075 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.001886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:38 | 20 | 20 |  | 1037.5159 | [19 empty] | -1.2208 | [5 empty] | 2.075 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.5849 | [19 empty] | -1.2208 | [5 empty] | 2.0752 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018862 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.5849 | [19 empty] | -1.2208 | [5 empty] | 2.0752 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018862 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.5849 | [19 empty] | -1.2208 | [5 empty] | 2.0752 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018862 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.5849 | [19 empty] | -1.2208 | [5 empty] | 2.0752 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018862 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.6539 | [19 empty] | -1.2207 | [5 empty] | 2.0753 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018863 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.6539 | [19 empty] | -1.2208 | [5 empty] | 2.0753 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018863 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.6539 | [19 empty] | -1.2208 | [5 empty] | 2.0753 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018863 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.6539 | [19 empty] | -1.2208 | [5 empty] | 2.0753 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018863 | -1.4557 | -1.4365 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.6539 | [19 empty] | -1.2208 | [5 empty] | 2.0753 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018863 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.6539 | [19 empty] | -1.2208 | [5 empty] | 2.0753 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018863 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:39 | 20 | 20 |  | 1037.7229 | [19 empty] | -1.2207 | [5 empty] | 2.0754 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018864 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.7229 | [19 empty] | -1.2208 | [5 empty] | 2.0754 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018864 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.7918 | [19 empty] | -1.2208 | [5 empty] | 2.0756 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018866 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.7918 | [19 empty] | -1.2207 | [5 empty] | 2.0756 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018866 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.7918 | [19 empty] | -1.2207 | [5 empty] | 2.0756 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018866 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.7918 | [19 empty] | -1.2208 | [5 empty] | 2.0756 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018866 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.8608 | [19 empty] | -1.2208 | [5 empty] | 2.0757 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018867 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.8608 | [19 empty] | -1.2208 | [5 empty] | 2.0757 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018867 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.8608 | [19 empty] | -1.2208 | [5 empty] | 2.0757 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018867 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.8608 | [19 empty] | -1.2208 | [5 empty] | 2.0757 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018867 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:40 | 20 | 20 |  | 1037.8608 | [19 empty] | -1.2208 | [5 empty] | 2.0757 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018867 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9298 | [19 empty] | -1.2208 | [5 empty] | 2.0759 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018868 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9298 | [19 empty] | -1.2208 | [5 empty] | 2.0759 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018868 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9298 | [19 empty] | -1.2208 | [5 empty] | 2.0759 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018868 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9988 | [19 empty] | -1.2208 | [5 empty] | 2.076 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018869 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9988 | [19 empty] | -1.2208 | [5 empty] | 2.076 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018869 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9988 | [19 empty] | -1.2208 | [5 empty] | 2.076 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018869 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9988 | [19 empty] | -1.2208 | [5 empty] | 2.076 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018869 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9988 | [19 empty] | -1.2208 | [5 empty] | 2.076 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018869 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1037.9988 | [19 empty] | -1.2208 | [5 empty] | 2.076 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018869 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:41 | 20 | 20 |  | 1038.0677 | [19 empty] | -1.2208 | [5 empty] | 2.0761 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018871 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.0677 | [19 empty] | -1.2208 | [5 empty] | 2.0761 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018871 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.0677 | [19 empty] | -1.2208 | [5 empty] | 2.0761 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018871 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.1367 | [19 empty] | -1.2208 | [5 empty] | 2.0763 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018872 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.2057 | [19 empty] | -1.2208 | [5 empty] | 2.0764 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018873 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:42 | 20 | 20 |  | 1038.2057 | [19 empty] | -1.2208 | [5 empty] | 2.0764 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018873 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2057 | [19 empty] | -1.2208 | [5 empty] | 2.0764 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018873 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2057 | [19 empty] | -1.2208 | [5 empty] | 2.0764 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018873 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2057 | [19 empty] | -1.2208 | [5 empty] | 2.0764 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018873 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2057 | [19 empty] | -1.2207 | [5 empty] | 2.0764 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018873 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2747 | [19 empty] | -1.2208 | [5 empty] | 2.0765 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018874 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2747 | [19 empty] | -1.2208 | [5 empty] | 2.0765 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018874 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2747 | [19 empty] | -1.2208 | [5 empty] | 2.0765 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018874 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.2747 | [19 empty] | -1.2208 | [5 empty] | 2.0765 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018874 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:43 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.3436 | [19 empty] | -1.2208 | [5 empty] | 2.0767 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018876 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.4126 | [19 empty] | -1.2208 | [5 empty] | 2.0768 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018877 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.4126 | [19 empty] | -1.2208 | [5 empty] | 2.0768 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018877 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.4126 | [19 empty] | -1.2208 | [5 empty] | 2.0768 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018877 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.4126 | [19 empty] | -1.2208 | [5 empty] | 2.0768 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018877 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.4126 | [19 empty] | -1.2208 | [5 empty] | 2.0768 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018877 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:44 | 20 | 20 |  | 1038.4126 | [19 empty] | -1.2208 | [5 empty] | 2.0768 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018877 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.4816 | [19 empty] | -1.2208 | [5 empty] | 2.077 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018878 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.4816 | [19 empty] | -1.2208 | [5 empty] | 2.077 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018878 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.4816 | [19 empty] | -1.2208 | [5 empty] | 2.077 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018878 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.4816 | [19 empty] | -1.2208 | [5 empty] | 2.077 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018878 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.4816 | [19 empty] | -1.2207 | [5 empty] | 2.077 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018878 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.4816 | [19 empty] | -1.2208 | [5 empty] | 2.077 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018878 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2207 | [5 empty] | 2.0771 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2208 | [5 empty] | 2.0771 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2208 | [5 empty] | 2.0771 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:45 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2208 | [5 empty] | 2.0771 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2208 | [5 empty] | 2.0771 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2208 | [5 empty] | 2.0771 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.5506 | [19 empty] | -1.2207 | [5 empty] | 2.0771 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018879 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6195 | [19 empty] | -1.2207 | [5 empty] | 2.0772 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018881 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6195 | [19 empty] | -1.2207 | [5 empty] | 2.0772 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018881 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6195 | [19 empty] | -1.2207 | [5 empty] | 2.0772 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018881 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6195 | [19 empty] | -1.2208 | [5 empty] | 2.0772 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018881 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6195 | [19 empty] | -1.2208 | [5 empty] | 2.0772 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018881 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6195 | [19 empty] | -1.2208 | [5 empty] | 2.0772 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018881 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:46 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.6885 | [19 empty] | -1.2208 | [5 empty] | 2.0774 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018882 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.7575 | [19 empty] | -1.2207 | [5 empty] | 2.0775 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018883 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:47 | 20 | 20 |  | 1038.7575 | [19 empty] | -1.2208 | [5 empty] | 2.0775 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.7575 | [19 empty] | -1.2208 | [5 empty] | 2.0775 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.7575 | [19 empty] | -1.2208 | [5 empty] | 2.0775 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.7575 | [19 empty] | -1.2208 | [5 empty] | 2.0775 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018883 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2208 | [5 empty] | 2.0777 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2208 | [5 empty] | 2.0777 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2208 | [5 empty] | 2.0777 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2208 | [5 empty] | 2.0777 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2207 | [5 empty] | 2.0777 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2208 | [5 empty] | 2.0777 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2207 | [5 empty] | 2.0777 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:48 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2207 | [5 empty] | 2.0777 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8265 | [19 empty] | -1.2207 | [5 empty] | 2.0777 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018885 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2199 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2207 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.8954 | [19 empty] | -1.2208 | [5 empty] | 2.0778 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018886 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:49 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2208 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2208 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2207 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2207 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2208 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2208 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2207 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2207 | [5 empty] | 2.0779 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1038.9644 | [19 empty] | -1.2207 | [5 empty] | 2.0779 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018887 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2207 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2208 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:50 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2207 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2208 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2207 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2208 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2207 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2207 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2207 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2208 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.0334 | [19 empty] | -1.2208 | [5 empty] | 2.0781 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018888 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.1023 | [19 empty] | -1.2207 | [5 empty] | 2.0782 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001889 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:51 | 20 | 20 |  | 1039.1023 | [19 empty] | -1.2207 | [5 empty] | 2.0782 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001889 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1023 | [19 empty] | -1.2207 | [5 empty] | 2.0782 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.001889 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1023 | [19 empty] | -1.2207 | [5 empty] | 2.0782 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.001889 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2208 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2207 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2208 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2208 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2208 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2207 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.1713 | [19 empty] | -1.2207 | [5 empty] | 2.0783 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018891 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:52 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2208 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2207 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2207 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2208 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2207 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2208 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2208 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2207 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.2403 | [19 empty] | -1.2208 | [5 empty] | 2.0785 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018892 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2207 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2208 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:53 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2208 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2208 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2207 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2207 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3093 | [19 empty] | -1.2207 | [5 empty] | 2.0786 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018893 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2208 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:54 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2208 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2208 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.3782 | [19 empty] | -1.2207 | [5 empty] | 2.0788 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018895 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:55 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2208 | [5 empty] | 2.0789 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2208 | [5 empty] | 2.0789 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4555 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.4472 | [19 empty] | -1.2207 | [5 empty] | 2.0789 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018896 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:56 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2208 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4555 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2208 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:57 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2208 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2208 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5162 | [19 empty] | -1.2207 | [5 empty] | 2.079 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018897 | -1.4556 | -1.4363 | -1.4363 | -1.4363 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2208 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:58 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2208 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2208 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 11:59 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.5852 | [19 empty] | -1.2207 | [5 empty] | 2.0792 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018898 | -1.4556 | -1.4363 | -1.4363 | -1.4363 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2208 | [5 empty] | 2.0793 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4557 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:00 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.6541 | [19 empty] | -1.2207 | [5 empty] | 2.0793 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.00189 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4555 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2208 | [5 empty] | 2.0794 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4364 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:01 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2198 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4556 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2207 | [5 empty] | 2.0794 | -1.2197 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4555 | -1.4363 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 1039.7231 | [19 empty] | -1.2204 | [5 empty] | 2.0794 | -1.2194 | 0.90444 | 17.6664 | 0.23773 | -0.0018901 | -1.4553 | -1.436 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 980.267 | [19 empty] | -1.2177 | [5 empty] | 1.9605 | -1.2167 | 0.90444 | 17.6664 | 0.23773 | -0.001781 | -1.4527 | -1.4334 | -1.4334 | -1.4334 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 728.5097 | [19 empty] | -1.2173 | [5 empty] | 1.457 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0013189 | -1.4527 | -1.4334 | -1.4334 | -1.4334 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:02 | 20 | 20 |  | 684.5039 | [19 empty] | -1.2173 | [5 empty] | 1.369 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012381 | -1.4528 | -1.4335 | -1.4335 | -1.4335 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 684.9868 | [19 empty] | -1.2173 | [5 empty] | 1.37 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001239 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 682.6416 | [19 empty] | -1.2173 | [5 empty] | 1.3653 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012347 | -1.4529 | -1.4336 | -1.4336 | -1.4336 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 683.1244 | [19 empty] | -1.2174 | [5 empty] | 1.3662 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012355 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 683.5383 | [19 empty] | -1.2174 | [5 empty] | 1.3671 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012363 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 683.8832 | [19 empty] | -1.2174 | [5 empty] | 1.3678 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012369 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 684.1591 | [19 empty] | -1.2174 | [5 empty] | 1.3683 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012374 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 684.435 | [19 empty] | -1.2174 | [5 empty] | 1.3689 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001238 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 684.6419 | [19 empty] | -1.2174 | [5 empty] | 1.3693 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012383 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 684.8488 | [19 empty] | -1.2174 | [5 empty] | 1.3697 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012387 | -1.453 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:03 | 20 | 20 |  | 685.0557 | [19 empty] | -1.2174 | [5 empty] | 1.3701 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012391 | -1.453 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 685.1937 | [19 empty] | -1.2174 | [5 empty] | 1.3704 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012393 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 685.3316 | [19 empty] | -1.2174 | [5 empty] | 1.3707 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012396 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 685.4696 | [19 empty] | -1.2174 | [5 empty] | 1.3709 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012399 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 685.6075 | [19 empty] | -1.2174 | [5 empty] | 1.3712 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012401 | -1.453 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 685.7455 | [19 empty] | -1.2174 | [5 empty] | 1.3715 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012404 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 685.8834 | [19 empty] | -1.2174 | [5 empty] | 1.3718 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012406 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 686.0214 | [19 empty] | -1.2174 | [5 empty] | 1.372 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012409 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 686.1593 | [19 empty] | -1.2174 | [5 empty] | 1.3723 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012411 | -1.453 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 686.2973 | [19 empty] | -1.2174 | [5 empty] | 1.3726 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012414 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 686.4352 | [19 empty] | -1.2174 | [5 empty] | 1.3729 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012416 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:04 | 20 | 20 |  | 686.5042 | [19 empty] | -1.2174 | [5 empty] | 1.373 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012418 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 686.6421 | [19 empty] | -1.2174 | [5 empty] | 1.3733 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001242 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 686.7801 | [19 empty] | -1.2173 | [5 empty] | 1.3736 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012423 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 686.8491 | [19 empty] | -1.2173 | [5 empty] | 1.3737 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012424 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 686.987 | [19 empty] | -1.2174 | [5 empty] | 1.374 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012426 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 687.125 | [19 empty] | -1.2174 | [5 empty] | 1.3742 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012429 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 687.1939 | [19 empty] | -1.2174 | [5 empty] | 1.3744 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.001243 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 687.2629 | [19 empty] | -1.2174 | [5 empty] | 1.3745 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012431 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 687.4009 | [19 empty] | -1.2174 | [5 empty] | 1.3748 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012434 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 687.4698 | [19 empty] | -1.2174 | [5 empty] | 1.3749 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012435 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:05 | 20 | 20 |  | 687.6078 | [19 empty] | -1.2174 | [5 empty] | 1.3752 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012438 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 687.6768 | [19 empty] | -1.2174 | [5 empty] | 1.3754 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012439 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 687.7457 | [19 empty] | -1.2174 | [5 empty] | 1.3755 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001244 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 687.8837 | [19 empty] | -1.2174 | [5 empty] | 1.3758 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012443 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 687.9527 | [19 empty] | -1.2174 | [5 empty] | 1.3759 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012444 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 688.0216 | [19 empty] | -1.2174 | [5 empty] | 1.376 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012445 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 688.0906 | [19 empty] | -1.2174 | [5 empty] | 1.3762 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012447 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 688.2286 | [19 empty] | -1.2174 | [5 empty] | 1.3765 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012449 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 688.2975 | [19 empty] | -1.2173 | [5 empty] | 1.3766 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001245 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 688.3665 | [19 empty] | -1.2173 | [5 empty] | 1.3767 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012452 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:06 | 20 | 20 |  | 688.4355 | [19 empty] | -1.2173 | [5 empty] | 1.3769 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012453 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.5045 | [19 empty] | -1.2174 | [5 empty] | 1.377 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012454 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.5734 | [19 empty] | -1.2173 | [5 empty] | 1.3771 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012455 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.6424 | [19 empty] | -1.2173 | [5 empty] | 1.3773 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012457 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.7114 | [19 empty] | -1.2174 | [5 empty] | 1.3774 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012458 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.7804 | [19 empty] | -1.2174 | [5 empty] | 1.3776 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012459 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.9183 | [19 empty] | -1.2174 | [5 empty] | 1.3778 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012462 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 688.9183 | [19 empty] | -1.2174 | [5 empty] | 1.3778 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012462 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 689.0563 | [19 empty] | -1.2174 | [5 empty] | 1.3781 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012464 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 689.1252 | [19 empty] | -1.2174 | [5 empty] | 1.3783 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012466 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 689.1252 | [19 empty] | -1.2174 | [5 empty] | 1.3783 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012466 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:07 | 20 | 20 |  | 689.1942 | [19 empty] | -1.2174 | [5 empty] | 1.3784 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012467 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.2632 | [19 empty] | -1.2173 | [5 empty] | 1.3785 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012468 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.3322 | [19 empty] | -1.2173 | [5 empty] | 1.3787 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012469 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.4011 | [19 empty] | -1.2173 | [5 empty] | 1.3788 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012471 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.4701 | [19 empty] | -1.2173 | [5 empty] | 1.3789 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012472 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.5391 | [19 empty] | -1.2174 | [5 empty] | 1.3791 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012473 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.6081 | [19 empty] | -1.2174 | [5 empty] | 1.3792 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012474 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.677 | [19 empty] | -1.2174 | [5 empty] | 1.3794 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012476 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.746 | [19 empty] | -1.2174 | [5 empty] | 1.3795 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012477 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.746 | [19 empty] | -1.2173 | [5 empty] | 1.3795 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012477 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:08 | 20 | 20 |  | 689.815 | [19 empty] | -1.2173 | [5 empty] | 1.3796 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012478 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 689.8839 | [19 empty] | -1.2174 | [5 empty] | 1.3798 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001248 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 689.9529 | [19 empty] | -1.2173 | [5 empty] | 1.3799 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012481 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.0219 | [19 empty] | -1.2173 | [5 empty] | 1.38 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012482 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.0219 | [19 empty] | -1.2174 | [5 empty] | 1.38 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012482 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.0909 | [19 empty] | -1.2174 | [5 empty] | 1.3802 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012483 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.1598 | [19 empty] | -1.2174 | [5 empty] | 1.3803 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012485 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.2288 | [19 empty] | -1.2174 | [5 empty] | 1.3805 | -1.2165 | 0.90444 | 17.6664 | 0.23773 | -0.0012486 | -1.4529 | -1.4337 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.2978 | [19 empty] | -1.2174 | [5 empty] | 1.3806 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012487 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.3668 | [19 empty] | -1.2174 | [5 empty] | 1.3807 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012488 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:09 | 20 | 20 |  | 690.3668 | [19 empty] | -1.2174 | [5 empty] | 1.3807 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012488 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.4357 | [19 empty] | -1.2174 | [5 empty] | 1.3809 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001249 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.5047 | [19 empty] | -1.2174 | [5 empty] | 1.381 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012491 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.5737 | [19 empty] | -1.2174 | [5 empty] | 1.3811 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012492 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.5737 | [19 empty] | -1.2174 | [5 empty] | 1.3811 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012492 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.6427 | [19 empty] | -1.2173 | [5 empty] | 1.3813 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012493 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.7116 | [19 empty] | -1.2173 | [5 empty] | 1.3814 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012495 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.7806 | [19 empty] | -1.2174 | [5 empty] | 1.3816 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012496 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.7806 | [19 empty] | -1.2174 | [5 empty] | 1.3816 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012496 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.8496 | [19 empty] | -1.2173 | [5 empty] | 1.3817 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012497 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.9186 | [19 empty] | -1.2174 | [5 empty] | 1.3818 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012499 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:10 | 20 | 20 |  | 690.9875 | [19 empty] | -1.2174 | [5 empty] | 1.382 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.00125 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 690.9875 | [19 empty] | -1.2174 | [5 empty] | 1.382 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.00125 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.0565 | [19 empty] | -1.2174 | [5 empty] | 1.3821 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012501 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.1255 | [19 empty] | -1.2174 | [5 empty] | 1.3823 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012502 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.1945 | [19 empty] | -1.2174 | [5 empty] | 1.3824 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012504 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.1945 | [19 empty] | -1.2174 | [5 empty] | 1.3824 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012504 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.2634 | [19 empty] | -1.2174 | [5 empty] | 1.3825 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012505 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.3324 | [19 empty] | -1.2174 | [5 empty] | 1.3827 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012506 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.3324 | [19 empty] | -1.2173 | [5 empty] | 1.3827 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012506 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.4014 | [19 empty] | -1.2173 | [5 empty] | 1.3828 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012507 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:11 | 20 | 20 |  | 691.4704 | [19 empty] | -1.2173 | [5 empty] | 1.3829 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012509 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.5393 | [19 empty] | -1.2174 | [5 empty] | 1.3831 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001251 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.5393 | [19 empty] | -1.2173 | [5 empty] | 1.3831 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001251 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.6083 | [19 empty] | -1.2174 | [5 empty] | 1.3832 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012511 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.6083 | [19 empty] | -1.2174 | [5 empty] | 1.3832 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012511 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.6773 | [19 empty] | -1.2174 | [5 empty] | 1.3834 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012512 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.7463 | [19 empty] | -1.2174 | [5 empty] | 1.3835 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012514 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.7463 | [19 empty] | -1.2174 | [5 empty] | 1.3835 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012514 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.8152 | [19 empty] | -1.2174 | [5 empty] | 1.3836 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012515 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.8152 | [19 empty] | -1.2174 | [5 empty] | 1.3836 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012515 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.8842 | [19 empty] | -1.2174 | [5 empty] | 1.3838 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012516 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:12 | 20 | 20 |  | 691.8842 | [19 empty] | -1.2174 | [5 empty] | 1.3838 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012516 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 691.9532 | [19 empty] | -1.2174 | [5 empty] | 1.3839 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012518 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.0222 | [19 empty] | -1.2174 | [5 empty] | 1.384 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012519 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.0222 | [19 empty] | -1.2174 | [5 empty] | 1.384 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012519 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.0911 | [19 empty] | -1.2173 | [5 empty] | 1.3842 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001252 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.1601 | [19 empty] | -1.2174 | [5 empty] | 1.3843 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012521 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.1601 | [19 empty] | -1.2174 | [5 empty] | 1.3843 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012521 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.2291 | [19 empty] | -1.2173 | [5 empty] | 1.3845 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012523 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.2291 | [19 empty] | -1.2174 | [5 empty] | 1.3845 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012523 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.2981 | [19 empty] | -1.2174 | [5 empty] | 1.3846 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012524 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:13 | 20 | 20 |  | 692.2981 | [19 empty] | -1.2174 | [5 empty] | 1.3846 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012524 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.367 | [19 empty] | -1.2174 | [5 empty] | 1.3847 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012525 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.367 | [19 empty] | -1.2174 | [5 empty] | 1.3847 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012525 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.436 | [19 empty] | -1.2173 | [5 empty] | 1.3849 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012526 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.436 | [19 empty] | -1.2174 | [5 empty] | 1.3849 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012526 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.436 | [19 empty] | -1.2174 | [5 empty] | 1.3849 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012526 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.505 | [19 empty] | -1.2174 | [5 empty] | 1.385 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012528 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.505 | [19 empty] | -1.2173 | [5 empty] | 1.385 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012528 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.574 | [19 empty] | -1.2174 | [5 empty] | 1.3851 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012529 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.574 | [19 empty] | -1.2173 | [5 empty] | 1.3851 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012529 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:14 | 20 | 20 |  | 692.6429 | [19 empty] | -1.2173 | [5 empty] | 1.3853 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001253 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.6429 | [19 empty] | -1.2174 | [5 empty] | 1.3853 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001253 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.6429 | [19 empty] | -1.2173 | [5 empty] | 1.3853 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001253 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.7119 | [19 empty] | -1.2174 | [5 empty] | 1.3854 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012531 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.7119 | [19 empty] | -1.2173 | [5 empty] | 1.3854 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012531 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.7809 | [19 empty] | -1.2174 | [5 empty] | 1.3856 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012533 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.7809 | [19 empty] | -1.2173 | [5 empty] | 1.3856 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012533 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.8499 | [19 empty] | -1.2174 | [5 empty] | 1.3857 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012534 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.8499 | [19 empty] | -1.2174 | [5 empty] | 1.3857 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012534 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.8499 | [19 empty] | -1.2174 | [5 empty] | 1.3857 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012534 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:15 | 20 | 20 |  | 692.9188 | [19 empty] | -1.2174 | [5 empty] | 1.3858 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012535 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 692.9188 | [19 empty] | -1.2174 | [5 empty] | 1.3858 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012535 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 692.9878 | [19 empty] | -1.2174 | [5 empty] | 1.386 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012537 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 692.9878 | [19 empty] | -1.2174 | [5 empty] | 1.386 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012537 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.0568 | [19 empty] | -1.2174 | [5 empty] | 1.3861 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012538 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.0568 | [19 empty] | -1.2174 | [5 empty] | 1.3861 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012538 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.0568 | [19 empty] | -1.2174 | [5 empty] | 1.3861 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012538 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.1258 | [19 empty] | -1.2174 | [5 empty] | 1.3863 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012539 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.1258 | [19 empty] | -1.2174 | [5 empty] | 1.3863 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012539 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.1947 | [19 empty] | -1.2174 | [5 empty] | 1.3864 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001254 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.1947 | [19 empty] | -1.2174 | [5 empty] | 1.3864 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001254 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:16 | 20 | 20 |  | 693.1947 | [19 empty] | -1.2174 | [5 empty] | 1.3864 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001254 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.2637 | [19 empty] | -1.2173 | [5 empty] | 1.3865 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012542 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.2637 | [19 empty] | -1.2173 | [5 empty] | 1.3865 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012542 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.3327 | [19 empty] | -1.2173 | [5 empty] | 1.3867 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012543 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.3327 | [19 empty] | -1.2174 | [5 empty] | 1.3867 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012543 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.3327 | [19 empty] | -1.2174 | [5 empty] | 1.3867 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012543 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.4017 | [19 empty] | -1.2174 | [5 empty] | 1.3868 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012544 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.4017 | [19 empty] | -1.2173 | [5 empty] | 1.3868 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012544 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.4706 | [19 empty] | -1.2174 | [5 empty] | 1.3869 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012545 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.4706 | [19 empty] | -1.2174 | [5 empty] | 1.3869 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012545 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:17 | 20 | 20 |  | 693.4706 | [19 empty] | -1.2174 | [5 empty] | 1.3869 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012545 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.5396 | [19 empty] | -1.2174 | [5 empty] | 1.3871 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012547 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.5396 | [19 empty] | -1.2174 | [5 empty] | 1.3871 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012547 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.6086 | [19 empty] | -1.2173 | [5 empty] | 1.3872 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012548 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.6086 | [19 empty] | -1.2174 | [5 empty] | 1.3872 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012548 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.6086 | [19 empty] | -1.2174 | [5 empty] | 1.3872 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012548 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.6776 | [19 empty] | -1.2173 | [5 empty] | 1.3874 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012549 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.6776 | [19 empty] | -1.2173 | [5 empty] | 1.3874 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012549 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.6776 | [19 empty] | -1.2173 | [5 empty] | 1.3874 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012549 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.7465 | [19 empty] | -1.2173 | [5 empty] | 1.3875 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001255 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.7465 | [19 empty] | -1.2173 | [5 empty] | 1.3875 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001255 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:18 | 20 | 20 |  | 693.7465 | [19 empty] | -1.2173 | [5 empty] | 1.3875 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001255 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.8155 | [19 empty] | -1.2173 | [5 empty] | 1.3876 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012552 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.8155 | [19 empty] | -1.2173 | [5 empty] | 1.3876 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012552 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.8845 | [19 empty] | -1.2173 | [5 empty] | 1.3878 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012553 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.8845 | [19 empty] | -1.2173 | [5 empty] | 1.3878 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012553 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.8845 | [19 empty] | -1.2173 | [5 empty] | 1.3878 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012553 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.9535 | [19 empty] | -1.2174 | [5 empty] | 1.3879 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012554 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.9535 | [19 empty] | -1.2173 | [5 empty] | 1.3879 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012554 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 693.9535 | [19 empty] | -1.2174 | [5 empty] | 1.3879 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012554 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 694.0224 | [19 empty] | -1.2173 | [5 empty] | 1.388 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012556 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:19 | 20 | 20 |  | 694.0224 | [19 empty] | -1.2174 | [5 empty] | 1.388 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012556 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.0224 | [19 empty] | -1.2173 | [5 empty] | 1.388 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012556 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.0914 | [19 empty] | -1.2174 | [5 empty] | 1.3882 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012557 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.0914 | [19 empty] | -1.2173 | [5 empty] | 1.3882 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012557 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.1604 | [19 empty] | -1.2174 | [5 empty] | 1.3883 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012558 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.1604 | [19 empty] | -1.2173 | [5 empty] | 1.3883 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012558 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2173 | [5 empty] | 1.3885 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2174 | [5 empty] | 1.3885 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.2293 | [19 empty] | -1.2174 | [5 empty] | 1.3885 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012559 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2173 | [5 empty] | 1.3886 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:20 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2173 | [5 empty] | 1.3886 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.2983 | [19 empty] | -1.2174 | [5 empty] | 1.3886 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012561 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2173 | [5 empty] | 1.3887 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.3673 | [19 empty] | -1.2173 | [5 empty] | 1.3887 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012562 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2173 | [5 empty] | 1.3889 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2173 | [5 empty] | 1.3889 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2173 | [5 empty] | 1.3889 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.4363 | [19 empty] | -1.2173 | [5 empty] | 1.3889 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012563 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2173 | [5 empty] | 1.389 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2173 | [5 empty] | 1.389 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2173 | [5 empty] | 1.389 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:21 | 20 | 20 |  | 694.5052 | [19 empty] | -1.2173 | [5 empty] | 1.389 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012564 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2173 | [5 empty] | 1.3891 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2174 | [5 empty] | 1.3891 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.5742 | [19 empty] | -1.2173 | [5 empty] | 1.3891 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012566 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2173 | [5 empty] | 1.3893 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2173 | [5 empty] | 1.3893 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2173 | [5 empty] | 1.3893 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2174 | [5 empty] | 1.3893 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.6432 | [19 empty] | -1.2173 | [5 empty] | 1.3893 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012567 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2173 | [5 empty] | 1.3894 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:22 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2174 | [5 empty] | 1.3894 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2173 | [5 empty] | 1.3894 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.7122 | [19 empty] | -1.2173 | [5 empty] | 1.3894 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012568 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2174 | [5 empty] | 1.3896 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.7811 | [19 empty] | -1.2174 | [5 empty] | 1.3896 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012569 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2173 | [5 empty] | 1.3897 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2173 | [5 empty] | 1.3897 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2173 | [5 empty] | 1.3897 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.8501 | [19 empty] | -1.2173 | [5 empty] | 1.3897 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012571 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2174 | [5 empty] | 1.3898 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:23 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2173 | [5 empty] | 1.3898 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2173 | [5 empty] | 1.3898 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2174 | [5 empty] | 1.3898 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 694.9191 | [19 empty] | -1.2173 | [5 empty] | 1.3898 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012572 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 694.9881 | [19 empty] | -1.2173 | [5 empty] | 1.39 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 694.9881 | [19 empty] | -1.2173 | [5 empty] | 1.39 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 694.9881 | [19 empty] | -1.2173 | [5 empty] | 1.39 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012573 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 695.057 | [19 empty] | -1.2173 | [5 empty] | 1.3901 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 695.057 | [19 empty] | -1.2173 | [5 empty] | 1.3901 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 695.057 | [19 empty] | -1.2173 | [5 empty] | 1.3901 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 695.057 | [19 empty] | -1.2173 | [5 empty] | 1.3901 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012574 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:24 | 20 | 20 |  | 695.126 | [19 empty] | -1.2173 | [5 empty] | 1.3903 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.126 | [19 empty] | -1.2173 | [5 empty] | 1.3903 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.126 | [19 empty] | -1.2173 | [5 empty] | 1.3903 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.126 | [19 empty] | -1.2173 | [5 empty] | 1.3903 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.126 | [19 empty] | -1.2173 | [5 empty] | 1.3903 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012576 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.195 | [19 empty] | -1.2173 | [5 empty] | 1.3904 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.195 | [19 empty] | -1.2173 | [5 empty] | 1.3904 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.195 | [19 empty] | -1.2173 | [5 empty] | 1.3904 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012577 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.264 | [19 empty] | -1.2173 | [5 empty] | 1.3905 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.264 | [19 empty] | -1.2173 | [5 empty] | 1.3905 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:25 | 20 | 20 |  | 695.264 | [19 empty] | -1.2173 | [5 empty] | 1.3905 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.264 | [19 empty] | -1.2173 | [5 empty] | 1.3905 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012578 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2173 | [5 empty] | 1.3907 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2173 | [5 empty] | 1.3907 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2173 | [5 empty] | 1.3907 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2173 | [5 empty] | 1.3907 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2174 | [5 empty] | 1.3907 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.3329 | [19 empty] | -1.2173 | [5 empty] | 1.3907 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001258 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.4019 | [19 empty] | -1.2173 | [5 empty] | 1.3908 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.4019 | [19 empty] | -1.2174 | [5 empty] | 1.3908 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.4019 | [19 empty] | -1.2173 | [5 empty] | 1.3908 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012581 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:26 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2173 | [5 empty] | 1.3909 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2174 | [5 empty] | 1.3909 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2173 | [5 empty] | 1.3909 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2174 | [5 empty] | 1.3909 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.4709 | [19 empty] | -1.2174 | [5 empty] | 1.3909 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012582 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2173 | [5 empty] | 1.3911 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2174 | [5 empty] | 1.3911 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2173 | [5 empty] | 1.3911 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.5399 | [19 empty] | -1.2173 | [5 empty] | 1.3911 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012583 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:27 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6088 | [19 empty] | -1.2173 | [5 empty] | 1.3912 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012585 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:28 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.6778 | [19 empty] | -1.2173 | [5 empty] | 1.3914 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012586 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:29 | 20 | 20 |  | 695.7468 | [19 empty] | -1.2173 | [5 empty] | 1.3915 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012587 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8158 | [19 empty] | -1.2173 | [5 empty] | 1.3916 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012588 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8158 | [19 empty] | -1.2173 | [5 empty] | 1.3916 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012588 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8158 | [19 empty] | -1.2173 | [5 empty] | 1.3916 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012588 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8158 | [19 empty] | -1.2173 | [5 empty] | 1.3916 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012588 | -1.4528 | -1.4335 | -1.4335 | -1.4335 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2174 | [5 empty] | 1.3918 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:30 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.8847 | [19 empty] | -1.2173 | [5 empty] | 1.3918 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001259 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2173 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2173 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2173 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2172 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2173 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2173 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2173 | [5 empty] | 1.3919 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:31 | 20 | 20 |  | 695.9537 | [19 empty] | -1.2172 | [5 empty] | 1.3919 | -1.2162 | 0.90444 | 17.6664 | 0.23773 | -0.0012591 | -1.4527 | -1.4334 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2172 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2173 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2173 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2173 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2173 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2173 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0227 | [19 empty] | -1.2173 | [5 empty] | 1.392 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012592 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2173 | [5 empty] | 1.3922 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2173 | [5 empty] | 1.3922 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2173 | [5 empty] | 1.3922 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:32 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2173 | [5 empty] | 1.3922 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.0917 | [19 empty] | -1.2173 | [5 empty] | 1.3922 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012593 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2174 | [5 empty] | 1.3923 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:33 | 20 | 20 |  | 696.1606 | [19 empty] | -1.2173 | [5 empty] | 1.3923 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012595 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2173 | [5 empty] | 1.3925 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2173 | [5 empty] | 1.3925 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2173 | [5 empty] | 1.3925 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2174 | [5 empty] | 1.3925 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2296 | [19 empty] | -1.2174 | [5 empty] | 1.3925 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012596 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:34 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.2986 | [19 empty] | -1.2173 | [5 empty] | 1.3926 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012597 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2174 | [5 empty] | 1.3927 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:35 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2173 | [5 empty] | 1.3929 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.3676 | [19 empty] | -1.2173 | [5 empty] | 1.3927 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012599 | -1.4528 | -1.4335 | -1.4335 | -1.4335 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2173 | [5 empty] | 1.3929 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2173 | [5 empty] | 1.3929 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2173 | [5 empty] | 1.3929 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.4365 | [19 empty] | -1.2173 | [5 empty] | 1.3929 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.00126 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2173 | [5 empty] | 1.393 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2173 | [5 empty] | 1.393 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2173 | [5 empty] | 1.393 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:36 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2173 | [5 empty] | 1.393 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5055 | [19 empty] | -1.2173 | [5 empty] | 1.393 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012601 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2174 | [5 empty] | 1.3931 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:37 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.5745 | [19 empty] | -1.2173 | [5 empty] | 1.3931 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012602 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2173 | [5 empty] | 1.3933 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2173 | [5 empty] | 1.3933 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2173 | [5 empty] | 1.3933 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2173 | [5 empty] | 1.3933 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.6435 | [19 empty] | -1.2173 | [5 empty] | 1.3933 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012604 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:38 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2172 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4527 | -1.4334 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7124 | [19 empty] | -1.2173 | [5 empty] | 1.3934 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012605 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:39 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.7814 | [19 empty] | -1.2173 | [5 empty] | 1.3936 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012606 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2173 | [5 empty] | 1.3937 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2173 | [5 empty] | 1.3937 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:40 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2173 | [5 empty] | 1.3937 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2173 | [5 empty] | 1.3937 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.8504 | [19 empty] | -1.2173 | [5 empty] | 1.3937 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012607 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4527 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2172 | [5 empty] | 1.3938 | -1.2162 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4527 | -1.4334 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2172 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4527 | -1.4334 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4527 | -1.4334 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:41 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9194 | [19 empty] | -1.2173 | [5 empty] | 1.3938 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012609 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2173 | [5 empty] | 1.394 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2174 | [5 empty] | 1.394 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2173 | [5 empty] | 1.394 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2173 | [5 empty] | 1.394 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2173 | [5 empty] | 1.394 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 696.9883 | [19 empty] | -1.2173 | [5 empty] | 1.394 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.001261 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:42 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2174 | [5 empty] | 1.3941 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.0573 | [19 empty] | -1.2173 | [5 empty] | 1.3941 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012611 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2174 | [5 empty] | 1.3943 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4529 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2173 | [5 empty] | 1.3943 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2173 | [5 empty] | 1.3943 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2173 | [5 empty] | 1.3943 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:43 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2173 | [5 empty] | 1.3943 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2173 | [5 empty] | 1.3943 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2174 | [5 empty] | 1.3943 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1263 | [19 empty] | -1.2173 | [5 empty] | 1.3943 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012612 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2173 | [5 empty] | 1.3944 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2173 | [5 empty] | 1.3944 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2174 | [5 empty] | 1.3944 | -1.2164 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4528 | -1.4336 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2173 | [5 empty] | 1.3944 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2173 | [5 empty] | 1.3944 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2173 | [5 empty] | 1.3944 | -1.2163 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4528 | -1.4335 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:44 | 20 | 20 |  | 697.1953 | [19 empty] | -1.2164 | [5 empty] | 1.3944 | -1.2154 | 0.90444 | 17.6664 | 0.23773 | -0.0012614 | -1.4519 | -1.4326 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 588.7672 | [19 empty] | -1.2154 | [5 empty] | 1.1775 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0010624 | -1.4511 | -1.4318 | -1.4318 | -1.4318 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 513.5849 | [19 empty] | -1.2154 | [5 empty] | 1.0272 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092435 | -1.4512 | -1.4319 | -1.4319 | -1.4319 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 513.102 | [19 empty] | -1.2155 | [5 empty] | 1.0262 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092346 | -1.4513 | -1.432 | -1.432 | -1.432 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 513.4469 | [19 empty] | -1.2155 | [5 empty] | 1.0269 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.0009241 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 513.7918 | [19 empty] | -1.2155 | [5 empty] | 1.0276 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092473 | -1.4513 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 513.9987 | [19 empty] | -1.2155 | [5 empty] | 1.028 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092511 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 514.2056 | [19 empty] | -1.2155 | [5 empty] | 1.0284 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092549 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 514.4126 | [19 empty] | -1.2155 | [5 empty] | 1.0288 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092587 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 514.5505 | [19 empty] | -1.2155 | [5 empty] | 1.0291 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092612 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:45 | 20 | 20 |  | 514.6885 | [19 empty] | -1.2155 | [5 empty] | 1.0294 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092638 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 514.8264 | [19 empty] | -1.2155 | [5 empty] | 1.0297 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092663 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 514.9644 | [19 empty] | -1.2155 | [5 empty] | 1.0299 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092688 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.0333 | [19 empty] | -1.2155 | [5 empty] | 1.0301 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092701 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.1713 | [19 empty] | -1.2155 | [5 empty] | 1.0303 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092726 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.2403 | [19 empty] | -1.2155 | [5 empty] | 1.0305 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092739 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.3092 | [19 empty] | -1.2155 | [5 empty] | 1.0306 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092752 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.4472 | [19 empty] | -1.2155 | [5 empty] | 1.0309 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092777 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.5162 | [19 empty] | -1.2155 | [5 empty] | 1.031 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.0009279 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.5851 | [19 empty] | -1.2154 | [5 empty] | 1.0312 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092802 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:46 | 20 | 20 |  | 515.6541 | [19 empty] | -1.2155 | [5 empty] | 1.0313 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092815 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 515.7231 | [19 empty] | -1.2155 | [5 empty] | 1.0314 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092828 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 515.7921 | [19 empty] | -1.2155 | [5 empty] | 1.0316 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.0009284 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 515.861 | [19 empty] | -1.2155 | [5 empty] | 1.0317 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092853 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 515.93 | [19 empty] | -1.2155 | [5 empty] | 1.0319 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092865 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 515.999 | [19 empty] | -1.2154 | [5 empty] | 1.032 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092878 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 516.1369 | [19 empty] | -1.2155 | [5 empty] | 1.0323 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092903 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 516.1369 | [19 empty] | -1.2154 | [5 empty] | 1.0323 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092903 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 516.2059 | [19 empty] | -1.2155 | [5 empty] | 1.0324 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092916 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 516.2749 | [19 empty] | -1.2154 | [5 empty] | 1.0325 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092929 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 516.3439 | [19 empty] | -1.2155 | [5 empty] | 1.0327 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092941 | -1.4513 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:47 | 20 | 20 |  | 516.4128 | [19 empty] | -1.2155 | [5 empty] | 1.0328 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092954 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.4818 | [19 empty] | -1.2155 | [5 empty] | 1.033 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092967 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.5508 | [19 empty] | -1.2155 | [5 empty] | 1.0331 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092979 | -1.4513 | -1.4321 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.5508 | [19 empty] | -1.2155 | [5 empty] | 1.0331 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092979 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.6198 | [19 empty] | -1.2155 | [5 empty] | 1.0332 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092992 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.6887 | [19 empty] | -1.2155 | [5 empty] | 1.0334 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093005 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.7577 | [19 empty] | -1.2155 | [5 empty] | 1.0335 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093017 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.7577 | [19 empty] | -1.2155 | [5 empty] | 1.0335 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093017 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.8267 | [19 empty] | -1.2155 | [5 empty] | 1.0337 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.0009303 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.8267 | [19 empty] | -1.2155 | [5 empty] | 1.0337 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.0009303 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:48 | 20 | 20 |  | 516.8957 | [19 empty] | -1.2155 | [5 empty] | 1.0338 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093043 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 516.9646 | [19 empty] | -1.2155 | [5 empty] | 1.0339 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093055 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.0336 | [19 empty] | -1.2155 | [5 empty] | 1.0341 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093068 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.0336 | [19 empty] | -1.2155 | [5 empty] | 1.0341 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093068 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.1026 | [19 empty] | -1.2155 | [5 empty] | 1.0342 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093081 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.1716 | [19 empty] | -1.2155 | [5 empty] | 1.0343 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093093 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.1716 | [19 empty] | -1.2155 | [5 empty] | 1.0343 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093093 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.2405 | [19 empty] | -1.2155 | [5 empty] | 1.0345 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093106 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.3095 | [19 empty] | -1.2154 | [5 empty] | 1.0346 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093119 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.3095 | [19 empty] | -1.2154 | [5 empty] | 1.0346 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093119 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2154 | [5 empty] | 1.0348 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:49 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2155 | [5 empty] | 1.0348 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2155 | [5 empty] | 1.0348 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.4474 | [19 empty] | -1.2155 | [5 empty] | 1.0349 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093144 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.5164 | [19 empty] | -1.2155 | [5 empty] | 1.035 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093157 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.5164 | [19 empty] | -1.2154 | [5 empty] | 1.035 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093157 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.5854 | [19 empty] | -1.2155 | [5 empty] | 1.0352 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093169 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.5854 | [19 empty] | -1.2155 | [5 empty] | 1.0352 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093169 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.6544 | [19 empty] | -1.2155 | [5 empty] | 1.0353 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093182 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.6544 | [19 empty] | -1.2155 | [5 empty] | 1.0353 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093182 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.7233 | [19 empty] | -1.2155 | [5 empty] | 1.0354 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093195 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:50 | 20 | 20 |  | 517.7233 | [19 empty] | -1.2155 | [5 empty] | 1.0354 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093195 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 517.7923 | [19 empty] | -1.2154 | [5 empty] | 1.0356 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093207 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 519.5167 | [19 empty] | -1.2154 | [5 empty] | 1.039 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093524 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 515.2403 | [19 empty] | -1.2155 | [5 empty] | 1.0305 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00092739 | -1.4513 | -1.432 | -1.432 | -1.432 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.6198 | [19 empty] | -1.2154 | [5 empty] | 1.0332 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092992 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.6198 | [19 empty] | -1.2154 | [5 empty] | 1.0332 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00092992 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.6887 | [19 empty] | -1.2154 | [5 empty] | 1.0334 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093005 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.6887 | [19 empty] | -1.2155 | [5 empty] | 1.0334 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093005 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.7577 | [19 empty] | -1.2155 | [5 empty] | 1.0335 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093017 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.7577 | [19 empty] | -1.2154 | [5 empty] | 1.0335 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093017 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:51 | 20 | 20 |  | 516.7577 | [19 empty] | -1.2154 | [5 empty] | 1.0335 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093017 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 516.8267 | [19 empty] | -1.2155 | [5 empty] | 1.0337 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.0009303 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 516.8267 | [19 empty] | -1.2154 | [5 empty] | 1.0337 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009303 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 516.8957 | [19 empty] | -1.2155 | [5 empty] | 1.0338 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093043 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 516.8957 | [19 empty] | -1.2154 | [5 empty] | 1.0338 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093043 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 516.9646 | [19 empty] | -1.2155 | [5 empty] | 1.0339 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093055 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 516.9646 | [19 empty] | -1.2154 | [5 empty] | 1.0339 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093055 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 517.0336 | [19 empty] | -1.2154 | [5 empty] | 1.0341 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093068 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 517.0336 | [19 empty] | -1.2154 | [5 empty] | 1.0341 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093068 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 517.1026 | [19 empty] | -1.2154 | [5 empty] | 1.0342 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093081 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 517.1026 | [19 empty] | -1.2154 | [5 empty] | 1.0342 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093081 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:52 | 20 | 20 |  | 517.1026 | [19 empty] | -1.2154 | [5 empty] | 1.0342 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093081 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.1716 | [19 empty] | -1.2154 | [5 empty] | 1.0343 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093093 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.1716 | [19 empty] | -1.2154 | [5 empty] | 1.0343 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093093 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.2405 | [19 empty] | -1.2154 | [5 empty] | 1.0345 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093106 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.2405 | [19 empty] | -1.2154 | [5 empty] | 1.0345 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093106 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.2405 | [19 empty] | -1.2154 | [5 empty] | 1.0345 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093106 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.3095 | [19 empty] | -1.2154 | [5 empty] | 1.0346 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093119 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.3095 | [19 empty] | -1.2154 | [5 empty] | 1.0346 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093119 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2155 | [5 empty] | 1.0348 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2154 | [5 empty] | 1.0348 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:53 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2154 | [5 empty] | 1.0348 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.3785 | [19 empty] | -1.2154 | [5 empty] | 1.0348 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093131 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.4474 | [19 empty] | -1.2154 | [5 empty] | 1.0349 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093144 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.4474 | [19 empty] | -1.2154 | [5 empty] | 1.0349 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093144 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.4474 | [19 empty] | -1.2154 | [5 empty] | 1.0349 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093144 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.5164 | [19 empty] | -1.2155 | [5 empty] | 1.035 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093157 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.5164 | [19 empty] | -1.2155 | [5 empty] | 1.035 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093157 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.5854 | [19 empty] | -1.2154 | [5 empty] | 1.0352 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093169 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.5854 | [19 empty] | -1.2154 | [5 empty] | 1.0352 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093169 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.5854 | [19 empty] | -1.2155 | [5 empty] | 1.0352 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093169 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:54 | 20 | 20 |  | 517.6544 | [19 empty] | -1.2154 | [5 empty] | 1.0353 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093182 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.6544 | [19 empty] | -1.2155 | [5 empty] | 1.0353 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093182 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.7233 | [19 empty] | -1.2154 | [5 empty] | 1.0354 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093195 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.7233 | [19 empty] | -1.2154 | [5 empty] | 1.0354 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093195 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.7923 | [19 empty] | -1.2154 | [5 empty] | 1.0356 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093207 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.7923 | [19 empty] | -1.2154 | [5 empty] | 1.0356 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093207 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.7923 | [19 empty] | -1.2154 | [5 empty] | 1.0356 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093207 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.8613 | [19 empty] | -1.2154 | [5 empty] | 1.0357 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009322 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.8613 | [19 empty] | -1.2154 | [5 empty] | 1.0357 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009322 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.9303 | [19 empty] | -1.2154 | [5 empty] | 1.0359 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093233 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.9303 | [19 empty] | -1.2154 | [5 empty] | 1.0359 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093233 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:55 | 20 | 20 |  | 517.9303 | [19 empty] | -1.2154 | [5 empty] | 1.0359 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093233 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 517.9992 | [19 empty] | -1.2154 | [5 empty] | 1.036 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093245 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 517.9992 | [19 empty] | -1.2154 | [5 empty] | 1.036 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093245 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.0682 | [19 empty] | -1.2154 | [5 empty] | 1.0361 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093258 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.0682 | [19 empty] | -1.2154 | [5 empty] | 1.0361 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093258 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.0682 | [19 empty] | -1.2154 | [5 empty] | 1.0361 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093258 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.1372 | [19 empty] | -1.2154 | [5 empty] | 1.0363 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093271 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.1372 | [19 empty] | -1.2154 | [5 empty] | 1.0363 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093271 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.2062 | [19 empty] | -1.2154 | [5 empty] | 1.0364 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093283 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.2062 | [19 empty] | -1.2154 | [5 empty] | 1.0364 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093283 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:56 | 20 | 20 |  | 518.2062 | [19 empty] | -1.2154 | [5 empty] | 1.0364 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093283 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.2062 | [19 empty] | -1.2154 | [5 empty] | 1.0364 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093283 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.2751 | [19 empty] | -1.2154 | [5 empty] | 1.0366 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093296 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.2751 | [19 empty] | -1.2154 | [5 empty] | 1.0366 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093296 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.2751 | [19 empty] | -1.2154 | [5 empty] | 1.0366 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093296 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.3441 | [19 empty] | -1.2154 | [5 empty] | 1.0367 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093309 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.3441 | [19 empty] | -1.2154 | [5 empty] | 1.0367 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093309 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.4131 | [19 empty] | -1.2154 | [5 empty] | 1.0368 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093321 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.4131 | [19 empty] | -1.2154 | [5 empty] | 1.0368 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093321 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.4131 | [19 empty] | -1.2154 | [5 empty] | 1.0368 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093321 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:57 | 20 | 20 |  | 518.4821 | [19 empty] | -1.2154 | [5 empty] | 1.037 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093334 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.4821 | [19 empty] | -1.2154 | [5 empty] | 1.037 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093334 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.4821 | [19 empty] | -1.2154 | [5 empty] | 1.037 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093334 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.4821 | [19 empty] | -1.2155 | [5 empty] | 1.037 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093334 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.551 | [19 empty] | -1.2154 | [5 empty] | 1.0371 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093347 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.551 | [19 empty] | -1.2154 | [5 empty] | 1.0371 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093347 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.551 | [19 empty] | -1.2154 | [5 empty] | 1.0371 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093347 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.62 | [19 empty] | -1.2154 | [5 empty] | 1.0372 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093359 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.62 | [19 empty] | -1.2155 | [5 empty] | 1.0372 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093359 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.62 | [19 empty] | -1.2154 | [5 empty] | 1.0372 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093359 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.689 | [19 empty] | -1.2154 | [5 empty] | 1.0374 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093372 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:58 | 20 | 20 |  | 518.689 | [19 empty] | -1.2154 | [5 empty] | 1.0374 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093372 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.689 | [19 empty] | -1.2154 | [5 empty] | 1.0374 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093372 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.689 | [19 empty] | -1.2154 | [5 empty] | 1.0374 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093372 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.758 | [19 empty] | -1.2154 | [5 empty] | 1.0375 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093385 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.758 | [19 empty] | -1.2154 | [5 empty] | 1.0375 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093385 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.758 | [19 empty] | -1.2154 | [5 empty] | 1.0375 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093385 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.8269 | [19 empty] | -1.2154 | [5 empty] | 1.0377 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093397 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.8269 | [19 empty] | -1.2154 | [5 empty] | 1.0377 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093397 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.8269 | [19 empty] | -1.2154 | [5 empty] | 1.0377 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093397 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.8269 | [19 empty] | -1.2154 | [5 empty] | 1.0377 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093397 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 12:59 | 20 | 20 |  | 518.8959 | [19 empty] | -1.2154 | [5 empty] | 1.0378 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009341 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 518.8959 | [19 empty] | -1.2154 | [5 empty] | 1.0378 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009341 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 518.8959 | [19 empty] | -1.2154 | [5 empty] | 1.0378 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009341 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 518.9649 | [19 empty] | -1.2154 | [5 empty] | 1.0379 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093423 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 518.9649 | [19 empty] | -1.2154 | [5 empty] | 1.0379 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093423 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 518.9649 | [19 empty] | -1.2154 | [5 empty] | 1.0379 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093423 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 518.9649 | [19 empty] | -1.2154 | [5 empty] | 1.0379 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093423 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 519.0339 | [19 empty] | -1.2154 | [5 empty] | 1.0381 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093435 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 519.0339 | [19 empty] | -1.2154 | [5 empty] | 1.0381 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093435 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 519.0339 | [19 empty] | -1.2154 | [5 empty] | 1.0381 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093435 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 519.1028 | [19 empty] | -1.2154 | [5 empty] | 1.0382 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093448 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:00 | 20 | 20 |  | 519.1028 | [19 empty] | -1.2154 | [5 empty] | 1.0382 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093448 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.1028 | [19 empty] | -1.2154 | [5 empty] | 1.0382 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093448 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.1028 | [19 empty] | -1.2154 | [5 empty] | 1.0382 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093448 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.1028 | [19 empty] | -1.2154 | [5 empty] | 1.0382 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093448 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.1718 | [19 empty] | -1.2154 | [5 empty] | 1.0383 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093461 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.1718 | [19 empty] | -1.2154 | [5 empty] | 1.0383 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093461 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.1718 | [19 empty] | -1.2154 | [5 empty] | 1.0383 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093461 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.2408 | [19 empty] | -1.2154 | [5 empty] | 1.0385 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093473 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.2408 | [19 empty] | -1.2154 | [5 empty] | 1.0385 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093473 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.2408 | [19 empty] | -1.2154 | [5 empty] | 1.0385 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093473 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:01 | 20 | 20 |  | 519.2408 | [19 empty] | -1.2154 | [5 empty] | 1.0385 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093473 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.2408 | [19 empty] | -1.2154 | [5 empty] | 1.0385 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093473 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3098 | [19 empty] | -1.2155 | [5 empty] | 1.0386 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093486 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3098 | [19 empty] | -1.2154 | [5 empty] | 1.0386 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093486 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3098 | [19 empty] | -1.2154 | [5 empty] | 1.0386 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093486 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3098 | [19 empty] | -1.2154 | [5 empty] | 1.0386 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093486 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3098 | [19 empty] | -1.2154 | [5 empty] | 1.0386 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093486 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3787 | [19 empty] | -1.2154 | [5 empty] | 1.0388 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093499 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3787 | [19 empty] | -1.2154 | [5 empty] | 1.0388 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093499 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.3787 | [19 empty] | -1.2154 | [5 empty] | 1.0388 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093499 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:02 | 20 | 20 |  | 519.4477 | [19 empty] | -1.2154 | [5 empty] | 1.0389 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093511 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.4477 | [19 empty] | -1.2154 | [5 empty] | 1.0389 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093511 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.4477 | [19 empty] | -1.2154 | [5 empty] | 1.0389 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093511 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.4477 | [19 empty] | -1.2154 | [5 empty] | 1.0389 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093511 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5167 | [19 empty] | -1.2154 | [5 empty] | 1.039 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093524 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5167 | [19 empty] | -1.2154 | [5 empty] | 1.039 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093524 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5167 | [19 empty] | -1.2154 | [5 empty] | 1.039 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093524 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5857 | [19 empty] | -1.2154 | [5 empty] | 1.0392 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093537 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5857 | [19 empty] | -1.2154 | [5 empty] | 1.0392 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093537 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5857 | [19 empty] | -1.2154 | [5 empty] | 1.0392 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093537 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:03 | 20 | 20 |  | 519.5857 | [19 empty] | -1.2154 | [5 empty] | 1.0392 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093537 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.5857 | [19 empty] | -1.2154 | [5 empty] | 1.0392 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093537 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.6546 | [19 empty] | -1.2154 | [5 empty] | 1.0393 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093549 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.6546 | [19 empty] | -1.2154 | [5 empty] | 1.0393 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093549 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.6546 | [19 empty] | -1.2154 | [5 empty] | 1.0393 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093549 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.6546 | [19 empty] | -1.2154 | [5 empty] | 1.0393 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093549 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.7236 | [19 empty] | -1.2155 | [5 empty] | 1.0394 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093562 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.7236 | [19 empty] | -1.2154 | [5 empty] | 1.0394 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093562 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.7236 | [19 empty] | -1.2154 | [5 empty] | 1.0394 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093562 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.7236 | [19 empty] | -1.2154 | [5 empty] | 1.0394 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093562 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.7236 | [19 empty] | -1.2154 | [5 empty] | 1.0394 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093562 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:04 | 20 | 20 |  | 519.7236 | [19 empty] | -1.2154 | [5 empty] | 1.0394 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093562 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.7926 | [19 empty] | -1.2154 | [5 empty] | 1.0396 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093574 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.7926 | [19 empty] | -1.2154 | [5 empty] | 1.0396 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093574 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.7926 | [19 empty] | -1.2154 | [5 empty] | 1.0396 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093574 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.7926 | [19 empty] | -1.2154 | [5 empty] | 1.0396 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093574 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.8616 | [19 empty] | -1.2154 | [5 empty] | 1.0397 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093587 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.8616 | [19 empty] | -1.2155 | [5 empty] | 1.0397 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093587 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.8616 | [19 empty] | -1.2155 | [5 empty] | 1.0397 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093587 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.8616 | [19 empty] | -1.2154 | [5 empty] | 1.0397 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093587 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.8616 | [19 empty] | -1.2154 | [5 empty] | 1.0397 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093587 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:05 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:06 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9305 | [19 empty] | -1.2154 | [5 empty] | 1.0399 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.000936 | -1.4512 | -1.4319 | -1.4319 | -1.4319 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2155 | [5 empty] | 1.04 | -1.2145 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4513 | -1.432 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:07 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 519.9995 | [19 empty] | -1.2154 | [5 empty] | 1.04 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093612 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:08 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:09 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.0685 | [19 empty] | -1.2154 | [5 empty] | 1.0401 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093625 | -1.4512 | -1.4319 | -1.4319 | -1.4319 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:10 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.1375 | [19 empty] | -1.2154 | [5 empty] | 1.0403 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093638 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:11 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2064 | [19 empty] | -1.2154 | [5 empty] | 1.0404 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.0009365 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.2754 | [19 empty] | -1.2154 | [5 empty] | 1.0406 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093663 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:12 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:13 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.3444 | [19 empty] | -1.2154 | [5 empty] | 1.0407 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093676 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2154 | [5 empty] | 1.0408 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2154 | [5 empty] | 1.0408 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2154 | [5 empty] | 1.0408 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2154 | [5 empty] | 1.0408 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2154 | [5 empty] | 1.0408 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4134 | [19 empty] | -1.2154 | [5 empty] | 1.0408 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093688 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2153 | [5 empty] | 1.041 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:14 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.4823 | [19 empty] | -1.2154 | [5 empty] | 1.041 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093701 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:15 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.5513 | [19 empty] | -1.2154 | [5 empty] | 1.0411 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093714 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:16 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2154 | [5 empty] | 1.0412 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2153 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2153 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:17 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2153 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2153 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2153 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6203 | [19 empty] | -1.2153 | [5 empty] | 1.0412 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093726 | -1.4511 | -1.4318 | -1.4318 | -1.4318 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2154 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2153 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2153 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2154 | [5 empty] | 1.0414 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2154 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2154 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2153 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:18 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2154 | [5 empty] | 1.0414 | -1.2144 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4512 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2154 | [5 empty] | 1.0415 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4511 | -1.4319 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2153 | [5 empty] | 1.0415 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 520.6893 | [19 empty] | -1.2154 | [5 empty] | 1.0414 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093739 | -1.4511 | -1.4319 | -1.4319 | -1.4319 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2153 | [5 empty] | 1.0415 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2153 | [5 empty] | 1.0415 | -1.2143 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4511 | -1.4318 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 520.7582 | [19 empty] | -1.2138 | [5 empty] | 1.0415 | -1.2128 | 0.90444 | 17.6664 | 0.23773 | -0.00093752 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 363.5651 | [19 empty] | -1.2134 | [5 empty] | 0.72713 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00064898 | -1.4495 | -1.4302 | -1.4302 | -1.4302 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 341.4243 | [19 empty] | -1.2134 | [5 empty] | 0.68285 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060834 | -1.4496 | -1.4303 | -1.4303 | -1.4303 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 339.6999 | [19 empty] | -1.2134 | [5 empty] | 0.6794 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060518 | -1.4495 | -1.4302 | -1.4302 | -1.4302 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:19 | 20 | 20 |  | 340.1827 | [19 empty] | -1.2135 | [5 empty] | 0.68037 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060606 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 340.5276 | [19 empty] | -1.2134 | [5 empty] | 0.68106 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006067 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 340.8035 | [19 empty] | -1.2134 | [5 empty] | 0.68161 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006072 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.0104 | [19 empty] | -1.2134 | [5 empty] | 0.68202 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060758 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.2173 | [19 empty] | -1.2135 | [5 empty] | 0.68243 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060796 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.4243 | [19 empty] | -1.2135 | [5 empty] | 0.68285 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060834 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.5622 | [19 empty] | -1.2134 | [5 empty] | 0.68312 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006086 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.7002 | [19 empty] | -1.2135 | [5 empty] | 0.6834 | -1.2125 | 0.90444 | 17.6664 | 0.23773 | -0.00060885 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.8381 | [19 empty] | -1.2134 | [5 empty] | 0.68368 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006091 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 341.9071 | [19 empty] | -1.2134 | [5 empty] | 0.68381 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060923 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 342.045 | [19 empty] | -1.2135 | [5 empty] | 0.68409 | -1.2125 | 0.90444 | 17.6664 | 0.23773 | -0.00060948 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:20 | 20 | 20 |  | 342.183 | [19 empty] | -1.2134 | [5 empty] | 0.68437 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060974 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.252 | [19 empty] | -1.2135 | [5 empty] | 0.6845 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060986 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.3209 | [19 empty] | -1.2135 | [5 empty] | 0.68464 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00060999 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.4589 | [19 empty] | -1.2134 | [5 empty] | 0.68492 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061024 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.5279 | [19 empty] | -1.2134 | [5 empty] | 0.68506 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061037 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.5968 | [19 empty] | -1.2135 | [5 empty] | 0.68519 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061049 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.7348 | [19 empty] | -1.2135 | [5 empty] | 0.68547 | -1.2125 | 0.90444 | 17.6664 | 0.23773 | -0.00061075 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.8038 | [19 empty] | -1.2135 | [5 empty] | 0.68561 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061087 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.8727 | [19 empty] | -1.2135 | [5 empty] | 0.68575 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.000611 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 342.9417 | [19 empty] | -1.2134 | [5 empty] | 0.68588 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061113 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:21 | 20 | 20 |  | 343.0107 | [19 empty] | -1.2134 | [5 empty] | 0.68602 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061125 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.0797 | [19 empty] | -1.2134 | [5 empty] | 0.68616 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061138 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.1486 | [19 empty] | -1.2134 | [5 empty] | 0.6863 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061151 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.2176 | [19 empty] | -1.2134 | [5 empty] | 0.68644 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061163 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.2866 | [19 empty] | -1.2134 | [5 empty] | 0.68657 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061176 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.3556 | [19 empty] | -1.2135 | [5 empty] | 0.68671 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061189 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.4245 | [19 empty] | -1.2134 | [5 empty] | 0.68685 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061201 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.4935 | [19 empty] | -1.2134 | [5 empty] | 0.68699 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061214 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.4935 | [19 empty] | -1.2134 | [5 empty] | 0.68699 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061214 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.5625 | [19 empty] | -1.2134 | [5 empty] | 0.68712 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061227 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:22 | 20 | 20 |  | 343.6315 | [19 empty] | -1.2134 | [5 empty] | 0.68726 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061239 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.7004 | [19 empty] | -1.2134 | [5 empty] | 0.6874 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061252 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.7694 | [19 empty] | -1.2134 | [5 empty] | 0.68754 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061265 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.7694 | [19 empty] | -1.2134 | [5 empty] | 0.68754 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061265 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.8384 | [19 empty] | -1.2134 | [5 empty] | 0.68768 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061277 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.9074 | [19 empty] | -1.2134 | [5 empty] | 0.68781 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006129 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.9074 | [19 empty] | -1.2134 | [5 empty] | 0.68781 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006129 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 343.9763 | [19 empty] | -1.2134 | [5 empty] | 0.68795 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061303 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 344.0453 | [19 empty] | -1.2134 | [5 empty] | 0.68809 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061315 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 344.1143 | [19 empty] | -1.2134 | [5 empty] | 0.68823 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061328 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 344.1143 | [19 empty] | -1.2134 | [5 empty] | 0.68823 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061328 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:23 | 20 | 20 |  | 344.1833 | [19 empty] | -1.2134 | [5 empty] | 0.68837 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061341 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.2522 | [19 empty] | -1.2134 | [5 empty] | 0.6885 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061353 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.3212 | [19 empty] | -1.2134 | [5 empty] | 0.68864 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061366 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.3902 | [19 empty] | -1.2134 | [5 empty] | 0.68878 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061379 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.3902 | [19 empty] | -1.2134 | [5 empty] | 0.68878 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061379 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.4592 | [19 empty] | -1.2134 | [5 empty] | 0.68892 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061391 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.4592 | [19 empty] | -1.2134 | [5 empty] | 0.68892 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061391 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.5281 | [19 empty] | -1.2134 | [5 empty] | 0.68906 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061404 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.5281 | [19 empty] | -1.2134 | [5 empty] | 0.68906 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061404 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.5971 | [19 empty] | -1.2134 | [5 empty] | 0.68919 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061417 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:24 | 20 | 20 |  | 344.5971 | [19 empty] | -1.2134 | [5 empty] | 0.68919 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061417 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.6661 | [19 empty] | -1.2134 | [5 empty] | 0.68933 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061429 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.6661 | [19 empty] | -1.2134 | [5 empty] | 0.68933 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061429 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.7351 | [19 empty] | -1.2134 | [5 empty] | 0.68947 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061442 | -1.4495 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.804 | [19 empty] | -1.2134 | [5 empty] | 0.68961 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061455 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.804 | [19 empty] | -1.2134 | [5 empty] | 0.68961 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061455 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.873 | [19 empty] | -1.2134 | [5 empty] | 0.68975 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061467 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.873 | [19 empty] | -1.2135 | [5 empty] | 0.68975 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061467 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.942 | [19 empty] | -1.2134 | [5 empty] | 0.68988 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006148 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 344.942 | [19 empty] | -1.2134 | [5 empty] | 0.68988 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006148 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:25 | 20 | 20 |  | 345.0109 | [19 empty] | -1.2134 | [5 empty] | 0.69002 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061493 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.0109 | [19 empty] | -1.2134 | [5 empty] | 0.69002 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061493 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.0799 | [19 empty] | -1.2134 | [5 empty] | 0.69016 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061505 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.0799 | [19 empty] | -1.2134 | [5 empty] | 0.69016 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061505 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.1489 | [19 empty] | -1.2134 | [5 empty] | 0.6903 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061518 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.2179 | [19 empty] | -1.2134 | [5 empty] | 0.69044 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061531 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.2179 | [19 empty] | -1.2134 | [5 empty] | 0.69044 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061531 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.2868 | [19 empty] | -1.2135 | [5 empty] | 0.69057 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061543 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.2868 | [19 empty] | -1.2134 | [5 empty] | 0.69057 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061543 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.2868 | [19 empty] | -1.2135 | [5 empty] | 0.69057 | -1.2125 | 0.90444 | 17.6664 | 0.23773 | -0.00061543 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.3558 | [19 empty] | -1.2134 | [5 empty] | 0.69071 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061556 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:26 | 20 | 20 |  | 345.4248 | [19 empty] | -1.2133 | [5 empty] | 0.69085 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061569 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.4248 | [19 empty] | -1.2133 | [5 empty] | 0.69085 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061569 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.4938 | [19 empty] | -1.2134 | [5 empty] | 0.69099 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061581 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.4938 | [19 empty] | -1.2135 | [5 empty] | 0.69099 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061581 | -1.4496 | -1.4303 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.5627 | [19 empty] | -1.2134 | [5 empty] | 0.69113 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061594 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.5627 | [19 empty] | -1.2134 | [5 empty] | 0.69113 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061594 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.6317 | [19 empty] | -1.2134 | [5 empty] | 0.69126 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061607 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.6317 | [19 empty] | -1.2134 | [5 empty] | 0.69126 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061607 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.7007 | [19 empty] | -1.2133 | [5 empty] | 0.6914 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061619 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.7007 | [19 empty] | -1.2134 | [5 empty] | 0.6914 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061619 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:27 | 20 | 20 |  | 345.7697 | [19 empty] | -1.2134 | [5 empty] | 0.69154 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061632 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.7697 | [19 empty] | -1.2134 | [5 empty] | 0.69154 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061632 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.7697 | [19 empty] | -1.2134 | [5 empty] | 0.69154 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061632 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.8386 | [19 empty] | -1.2134 | [5 empty] | 0.69168 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061645 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.8386 | [19 empty] | -1.2134 | [5 empty] | 0.69168 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061645 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.9076 | [19 empty] | -1.2134 | [5 empty] | 0.69182 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061657 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.9076 | [19 empty] | -1.2134 | [5 empty] | 0.69182 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061657 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.9766 | [19 empty] | -1.2134 | [5 empty] | 0.69195 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006167 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 345.9766 | [19 empty] | -1.2134 | [5 empty] | 0.69195 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.0006167 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 346.0456 | [19 empty] | -1.2134 | [5 empty] | 0.69209 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061683 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 346.0456 | [19 empty] | -1.2134 | [5 empty] | 0.69209 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061683 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:28 | 20 | 20 |  | 346.0456 | [19 empty] | -1.2134 | [5 empty] | 0.69209 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061683 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.1145 | [19 empty] | -1.2134 | [5 empty] | 0.69223 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061695 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.1145 | [19 empty] | -1.2134 | [5 empty] | 0.69223 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061695 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.1835 | [19 empty] | -1.2133 | [5 empty] | 0.69237 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061708 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.1835 | [19 empty] | -1.2134 | [5 empty] | 0.69237 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061708 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.1835 | [19 empty] | -1.2133 | [5 empty] | 0.69237 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061708 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.2525 | [19 empty] | -1.2134 | [5 empty] | 0.6925 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006172 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.3215 | [19 empty] | -1.2134 | [5 empty] | 0.69264 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061733 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.3215 | [19 empty] | -1.2134 | [5 empty] | 0.69264 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061733 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.3215 | [19 empty] | -1.2134 | [5 empty] | 0.69264 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061733 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:29 | 20 | 20 |  | 346.3904 | [19 empty] | -1.2134 | [5 empty] | 0.69278 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061746 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.3904 | [19 empty] | -1.2133 | [5 empty] | 0.69278 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061746 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.4594 | [19 empty] | -1.2133 | [5 empty] | 0.69292 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061758 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.4594 | [19 empty] | -1.2133 | [5 empty] | 0.69292 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061758 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.4594 | [19 empty] | -1.2133 | [5 empty] | 0.69292 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061758 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.5284 | [19 empty] | -1.2134 | [5 empty] | 0.69306 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061771 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.5284 | [19 empty] | -1.2134 | [5 empty] | 0.69306 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061771 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.5974 | [19 empty] | -1.2134 | [5 empty] | 0.69319 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061784 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.5974 | [19 empty] | -1.2134 | [5 empty] | 0.69319 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061784 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.5974 | [19 empty] | -1.2134 | [5 empty] | 0.69319 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061784 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:30 | 20 | 20 |  | 346.5974 | [19 empty] | -1.2134 | [5 empty] | 0.69319 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061784 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.6663 | [19 empty] | -1.2134 | [5 empty] | 0.69333 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061796 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.6663 | [19 empty] | -1.2134 | [5 empty] | 0.69333 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061796 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.7353 | [19 empty] | -1.2134 | [5 empty] | 0.69347 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061809 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.7353 | [19 empty] | -1.2134 | [5 empty] | 0.69347 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061809 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.7353 | [19 empty] | -1.2134 | [5 empty] | 0.69347 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061809 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.8043 | [19 empty] | -1.2134 | [5 empty] | 0.69361 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061822 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.8043 | [19 empty] | -1.2134 | [5 empty] | 0.69361 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061822 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.8733 | [19 empty] | -1.2134 | [5 empty] | 0.69375 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061834 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.8733 | [19 empty] | -1.2133 | [5 empty] | 0.69375 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061834 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:31 | 20 | 20 |  | 346.9422 | [19 empty] | -1.2134 | [5 empty] | 0.69388 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061847 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 346.9422 | [19 empty] | -1.2133 | [5 empty] | 0.69388 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061847 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.0112 | [19 empty] | -1.2133 | [5 empty] | 0.69402 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006186 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.0112 | [19 empty] | -1.2133 | [5 empty] | 0.69402 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006186 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.0112 | [19 empty] | -1.2133 | [5 empty] | 0.69402 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006186 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.0802 | [19 empty] | -1.2133 | [5 empty] | 0.69416 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061872 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.0802 | [19 empty] | -1.2133 | [5 empty] | 0.69416 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061872 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.1492 | [19 empty] | -1.2134 | [5 empty] | 0.6943 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061885 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.1492 | [19 empty] | -1.2134 | [5 empty] | 0.6943 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061885 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.1492 | [19 empty] | -1.2133 | [5 empty] | 0.6943 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061885 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.2181 | [19 empty] | -1.2133 | [5 empty] | 0.69444 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061898 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:32 | 20 | 20 |  | 347.2181 | [19 empty] | -1.2134 | [5 empty] | 0.69444 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061898 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.2871 | [19 empty] | -1.2133 | [5 empty] | 0.69457 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006191 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.2871 | [19 empty] | -1.2133 | [5 empty] | 0.69457 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006191 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.2871 | [19 empty] | -1.2133 | [5 empty] | 0.69457 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006191 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.3561 | [19 empty] | -1.2133 | [5 empty] | 0.69471 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061923 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.3561 | [19 empty] | -1.2133 | [5 empty] | 0.69471 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061923 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.3561 | [19 empty] | -1.2133 | [5 empty] | 0.69471 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061923 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.3561 | [19 empty] | -1.2133 | [5 empty] | 0.69471 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061923 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.4251 | [19 empty] | -1.2134 | [5 empty] | 0.69485 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061936 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.4251 | [19 empty] | -1.2133 | [5 empty] | 0.69485 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061936 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:33 | 20 | 20 |  | 347.4251 | [19 empty] | -1.2133 | [5 empty] | 0.69485 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061936 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.494 | [19 empty] | -1.2134 | [5 empty] | 0.69499 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061948 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.494 | [19 empty] | -1.2134 | [5 empty] | 0.69499 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061948 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.494 | [19 empty] | -1.2133 | [5 empty] | 0.69499 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061948 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.494 | [19 empty] | -1.2134 | [5 empty] | 0.69499 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061948 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.563 | [19 empty] | -1.2133 | [5 empty] | 0.69513 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061961 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.563 | [19 empty] | -1.2134 | [5 empty] | 0.69513 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061961 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.563 | [19 empty] | -1.2134 | [5 empty] | 0.69513 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061961 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.563 | [19 empty] | -1.2134 | [5 empty] | 0.69513 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061961 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.632 | [19 empty] | -1.2133 | [5 empty] | 0.69526 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061974 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:34 | 20 | 20 |  | 347.632 | [19 empty] | -1.2134 | [5 empty] | 0.69526 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061974 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.632 | [19 empty] | -1.2134 | [5 empty] | 0.69526 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061974 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.632 | [19 empty] | -1.2133 | [5 empty] | 0.69526 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061974 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.701 | [19 empty] | -1.2133 | [5 empty] | 0.6954 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061986 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.701 | [19 empty] | -1.2134 | [5 empty] | 0.6954 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061986 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.7699 | [19 empty] | -1.2134 | [5 empty] | 0.69554 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061999 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.7699 | [19 empty] | -1.2134 | [5 empty] | 0.69554 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00061999 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.7699 | [19 empty] | -1.2134 | [5 empty] | 0.69554 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061999 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.7699 | [19 empty] | -1.2133 | [5 empty] | 0.69554 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061999 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.7699 | [19 empty] | -1.2134 | [5 empty] | 0.69554 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00061999 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.8389 | [19 empty] | -1.2133 | [5 empty] | 0.69568 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062012 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:35 | 20 | 20 |  | 347.8389 | [19 empty] | -1.2134 | [5 empty] | 0.69568 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062012 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.8389 | [19 empty] | -1.2133 | [5 empty] | 0.69568 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062012 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.8389 | [19 empty] | -1.2133 | [5 empty] | 0.69568 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062012 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.8389 | [19 empty] | -1.2133 | [5 empty] | 0.69568 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062012 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9079 | [19 empty] | -1.2134 | [5 empty] | 0.69582 | -1.2124 | 0.90444 | 17.6664 | 0.23773 | -0.00062024 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9079 | [19 empty] | -1.2134 | [5 empty] | 0.69582 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062024 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9079 | [19 empty] | -1.2133 | [5 empty] | 0.69582 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062024 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9079 | [19 empty] | -1.2133 | [5 empty] | 0.69582 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062024 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9769 | [19 empty] | -1.2133 | [5 empty] | 0.69595 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062037 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9769 | [19 empty] | -1.2134 | [5 empty] | 0.69595 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062037 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:36 | 20 | 20 |  | 347.9769 | [19 empty] | -1.2133 | [5 empty] | 0.69595 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062037 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 347.9769 | [19 empty] | -1.2133 | [5 empty] | 0.69595 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062037 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 347.9769 | [19 empty] | -1.2133 | [5 empty] | 0.69595 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062037 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.0458 | [19 empty] | -1.2133 | [5 empty] | 0.69609 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006205 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.0458 | [19 empty] | -1.2133 | [5 empty] | 0.69609 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006205 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.0458 | [19 empty] | -1.2133 | [5 empty] | 0.69609 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006205 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.0458 | [19 empty] | -1.2133 | [5 empty] | 0.69609 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006205 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.1148 | [19 empty] | -1.2133 | [5 empty] | 0.69623 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062062 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.1148 | [19 empty] | -1.2133 | [5 empty] | 0.69623 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062062 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.1838 | [19 empty] | -1.2133 | [5 empty] | 0.69637 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062075 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.1838 | [19 empty] | -1.2134 | [5 empty] | 0.69637 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062075 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:37 | 20 | 20 |  | 348.1838 | [19 empty] | -1.2133 | [5 empty] | 0.69637 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062075 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.1838 | [19 empty] | -1.2133 | [5 empty] | 0.69637 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062075 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.2528 | [19 empty] | -1.2133 | [5 empty] | 0.69651 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062088 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.3217 | [19 empty] | -1.2133 | [5 empty] | 0.69664 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.000621 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:38 | 20 | 20 |  | 348.3217 | [19 empty] | -1.2133 | [5 empty] | 0.69664 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.000621 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3217 | [19 empty] | -1.2133 | [5 empty] | 0.69664 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.000621 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.3907 | [19 empty] | -1.2133 | [5 empty] | 0.69678 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062113 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.4597 | [19 empty] | -1.2133 | [5 empty] | 0.69692 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062126 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:39 | 20 | 20 |  | 348.4597 | [19 empty] | -1.2133 | [5 empty] | 0.69692 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062126 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.4597 | [19 empty] | -1.2133 | [5 empty] | 0.69692 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062126 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.4597 | [19 empty] | -1.2133 | [5 empty] | 0.69692 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062126 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.4597 | [19 empty] | -1.2133 | [5 empty] | 0.69692 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062126 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.4597 | [19 empty] | -1.2133 | [5 empty] | 0.69692 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062126 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5287 | [19 empty] | -1.2133 | [5 empty] | 0.69706 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062138 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5287 | [19 empty] | -1.2133 | [5 empty] | 0.69706 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062138 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5287 | [19 empty] | -1.2134 | [5 empty] | 0.69706 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062138 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5287 | [19 empty] | -1.2133 | [5 empty] | 0.69706 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062138 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5976 | [19 empty] | -1.2133 | [5 empty] | 0.6972 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062151 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5976 | [19 empty] | -1.2133 | [5 empty] | 0.6972 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062151 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:40 | 20 | 20 |  | 348.5976 | [19 empty] | -1.2133 | [5 empty] | 0.6972 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062151 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.5976 | [19 empty] | -1.2133 | [5 empty] | 0.6972 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062151 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.5976 | [19 empty] | -1.2133 | [5 empty] | 0.6972 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062151 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2133 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2133 | [5 empty] | 0.69733 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2133 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2134 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4495 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2133 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2133 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2134 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4302 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:41 | 20 | 20 |  | 348.6666 | [19 empty] | -1.2133 | [5 empty] | 0.69733 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062164 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.7356 | [19 empty] | -1.2133 | [5 empty] | 0.69747 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062176 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.8046 | [19 empty] | -1.2133 | [5 empty] | 0.69761 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062189 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.8046 | [19 empty] | -1.2133 | [5 empty] | 0.69761 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062189 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:42 | 20 | 20 |  | 348.8046 | [19 empty] | -1.2133 | [5 empty] | 0.69761 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062189 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8046 | [19 empty] | -1.2133 | [5 empty] | 0.69761 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062189 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8046 | [19 empty] | -1.2133 | [5 empty] | 0.69761 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062189 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.8735 | [19 empty] | -1.2133 | [5 empty] | 0.69775 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062202 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:43 | 20 | 20 |  | 348.9425 | [19 empty] | -1.2133 | [5 empty] | 0.69789 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062214 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 348.9425 | [19 empty] | -1.2133 | [5 empty] | 0.69789 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062214 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 348.9425 | [19 empty] | -1.2133 | [5 empty] | 0.69789 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062214 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 348.9425 | [19 empty] | -1.2133 | [5 empty] | 0.69789 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062214 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 348.9425 | [19 empty] | -1.2132 | [5 empty] | 0.69789 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062214 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 348.9425 | [19 empty] | -1.2133 | [5 empty] | 0.69789 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062214 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 349.0115 | [19 empty] | -1.2133 | [5 empty] | 0.69802 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062227 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 349.0115 | [19 empty] | -1.2133 | [5 empty] | 0.69802 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062227 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 349.0115 | [19 empty] | -1.2133 | [5 empty] | 0.69802 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062227 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 349.0115 | [19 empty] | -1.2133 | [5 empty] | 0.69802 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062227 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:44 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2133 | [5 empty] | 0.69816 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0115 | [19 empty] | -1.2133 | [5 empty] | 0.69802 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062227 | -1.4494 | -1.4301 | -1.4301 | -1.4301 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0115 | [19 empty] | -1.2133 | [5 empty] | 0.69802 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062227 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2132 | [5 empty] | 0.69816 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2133 | [5 empty] | 0.69816 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2132 | [5 empty] | 0.69816 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2133 | [5 empty] | 0.69816 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2133 | [5 empty] | 0.69816 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.0805 | [19 empty] | -1.2132 | [5 empty] | 0.69816 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006224 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.1494 | [19 empty] | -1.2133 | [5 empty] | 0.6983 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062252 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:45 | 20 | 20 |  | 349.1494 | [19 empty] | -1.2133 | [5 empty] | 0.6983 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062252 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.1494 | [19 empty] | -1.2133 | [5 empty] | 0.6983 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062252 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.1494 | [19 empty] | -1.2133 | [5 empty] | 0.6983 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062252 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.1494 | [19 empty] | -1.2133 | [5 empty] | 0.6983 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062252 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2133 | [5 empty] | 0.69844 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2133 | [5 empty] | 0.69844 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2133 | [5 empty] | 0.69844 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2133 | [5 empty] | 0.69844 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2132 | [5 empty] | 0.69844 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2133 | [5 empty] | 0.69844 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2184 | [19 empty] | -1.2133 | [5 empty] | 0.69844 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062265 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:46 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2133 | [5 empty] | 0.69857 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2132 | [5 empty] | 0.69857 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2133 | [5 empty] | 0.69857 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2133 | [5 empty] | 0.69857 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2133 | [5 empty] | 0.69857 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2133 | [5 empty] | 0.69857 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.2874 | [19 empty] | -1.2133 | [5 empty] | 0.69857 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.00062278 | -1.4493 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.3563 | [19 empty] | -1.2133 | [5 empty] | 0.69871 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006229 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.3563 | [19 empty] | -1.2132 | [5 empty] | 0.69871 | -1.2122 | 0.90444 | 17.6664 | 0.23773 | -0.0006229 | -1.4493 | -1.43 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 349.3563 | [19 empty] | -1.2133 | [5 empty] | 0.69871 | -1.2123 | 0.90444 | 17.6664 | 0.23773 | -0.0006229 | -1.4494 | -1.4301 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:47 | 20 | 20 |  | 352.1843 | [19 empty] | -1.2115 | [5 empty] | 0.70437 | -1.2105 | 0.90444 | 17.6664 | 0.23773 | -0.00062809 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 187.1971 | [19 empty] | -1.2112 | [5 empty] | 0.37439 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00032525 | -1.4476 | -1.4283 | -1.4283 | -1.4283 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 169.1257 | [19 empty] | -1.2112 | [5 empty] | 0.33825 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029208 | -1.4476 | -1.4283 | -1.4283 | -1.4283 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 170.2983 | [19 empty] | -1.2113 | [5 empty] | 0.3406 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029423 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 170.9191 | [19 empty] | -1.2113 | [5 empty] | 0.34184 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029537 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 171.3329 | [19 empty] | -1.2113 | [5 empty] | 0.34267 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029613 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 171.7468 | [19 empty] | -1.2113 | [5 empty] | 0.34349 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 172.0227 | [19 empty] | -1.2113 | [5 empty] | 0.34405 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2113 | [5 empty] | 0.3446 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 172.5055 | [19 empty] | -1.2113 | [5 empty] | 0.34501 | -1.2103 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 172.7124 | [19 empty] | -1.2113 | [5 empty] | 0.34542 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4476 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:48 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2112 | [5 empty] | 0.34584 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.0573 | [19 empty] | -1.2113 | [5 empty] | 0.34611 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2112 | [5 empty] | 0.34653 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2113 | [5 empty] | 0.3468 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2113 | [5 empty] | 0.34708 | -1.2103 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.678 | [19 empty] | -1.2113 | [5 empty] | 0.34736 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.816 | [19 empty] | -1.2113 | [5 empty] | 0.34763 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030069 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 173.885 | [19 empty] | -1.2113 | [5 empty] | 0.34777 | -1.2103 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2113 | [5 empty] | 0.34805 | -1.2103 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2113 | [5 empty] | 0.34832 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:49 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2113 | [5 empty] | 0.34846 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2113 | [5 empty] | 0.34874 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4476 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2112 | [5 empty] | 0.34887 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4476 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2112 | [5 empty] | 0.34901 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2112 | [5 empty] | 0.34929 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2112 | [5 empty] | 0.34943 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2112 | [5 empty] | 0.34956 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2112 | [5 empty] | 0.34984 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2113 | [5 empty] | 0.34998 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 175.0575 | [19 empty] | -1.2113 | [5 empty] | 0.35012 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030297 | -1.4477 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:50 | 20 | 20 |  | 175.1955 | [19 empty] | -1.2112 | [5 empty] | 0.35039 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030322 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.2645 | [19 empty] | -1.2112 | [5 empty] | 0.35053 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030335 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.3334 | [19 empty] | -1.2112 | [5 empty] | 0.35067 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030348 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.4024 | [19 empty] | -1.2112 | [5 empty] | 0.3508 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0003036 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.4024 | [19 empty] | -1.2113 | [5 empty] | 0.3508 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0003036 | -1.4476 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.4714 | [19 empty] | -1.2112 | [5 empty] | 0.35094 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030373 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.6093 | [19 empty] | -1.2112 | [5 empty] | 0.35122 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030398 | -1.4475 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.6783 | [19 empty] | -1.2112 | [5 empty] | 0.35136 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030411 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.6783 | [19 empty] | -1.2112 | [5 empty] | 0.35136 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030411 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.7473 | [19 empty] | -1.2112 | [5 empty] | 0.35149 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030424 | -1.4475 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.8163 | [19 empty] | -1.2112 | [5 empty] | 0.35163 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030436 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:51 | 20 | 20 |  | 175.8852 | [19 empty] | -1.2112 | [5 empty] | 0.35177 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030449 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 175.9542 | [19 empty] | -1.2112 | [5 empty] | 0.35191 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030462 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 175.9542 | [19 empty] | -1.2112 | [5 empty] | 0.35191 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030462 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.0232 | [19 empty] | -1.2112 | [5 empty] | 0.35205 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030474 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.0922 | [19 empty] | -1.2112 | [5 empty] | 0.35218 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030487 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.1611 | [19 empty] | -1.2112 | [5 empty] | 0.35232 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.000305 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.2301 | [19 empty] | -1.2112 | [5 empty] | 0.35246 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030512 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.2301 | [19 empty] | -1.2113 | [5 empty] | 0.35246 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030512 | -1.4476 | -1.4284 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.3681 | [19 empty] | -1.2112 | [5 empty] | 0.35274 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030538 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.3681 | [19 empty] | -1.2112 | [5 empty] | 0.35274 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030538 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:52 | 20 | 20 |  | 176.437 | [19 empty] | -1.2112 | [5 empty] | 0.35287 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0003055 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.506 | [19 empty] | -1.2112 | [5 empty] | 0.35301 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030563 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.506 | [19 empty] | -1.2111 | [5 empty] | 0.35301 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030563 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.575 | [19 empty] | -1.2112 | [5 empty] | 0.35315 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030576 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.644 | [19 empty] | -1.2112 | [5 empty] | 0.35329 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030588 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.7129 | [19 empty] | -1.2111 | [5 empty] | 0.35343 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030601 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.7129 | [19 empty] | -1.2112 | [5 empty] | 0.35343 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030601 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.7819 | [19 empty] | -1.2112 | [5 empty] | 0.35356 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030613 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.8509 | [19 empty] | -1.2112 | [5 empty] | 0.3537 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030626 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.9198 | [19 empty] | -1.2112 | [5 empty] | 0.35384 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030639 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:53 | 20 | 20 |  | 176.9198 | [19 empty] | -1.2112 | [5 empty] | 0.35384 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030639 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 176.9888 | [19 empty] | -1.2112 | [5 empty] | 0.35398 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030651 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 176.9888 | [19 empty] | -1.2112 | [5 empty] | 0.35398 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030651 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.0578 | [19 empty] | -1.2111 | [5 empty] | 0.35412 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030664 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.1268 | [19 empty] | -1.2111 | [5 empty] | 0.35425 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030677 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.1268 | [19 empty] | -1.2112 | [5 empty] | 0.35425 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030677 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.1957 | [19 empty] | -1.2112 | [5 empty] | 0.35439 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030689 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.1957 | [19 empty] | -1.2112 | [5 empty] | 0.35439 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030689 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.2647 | [19 empty] | -1.2112 | [5 empty] | 0.35453 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030702 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.3337 | [19 empty] | -1.2112 | [5 empty] | 0.35467 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030715 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.3337 | [19 empty] | -1.2112 | [5 empty] | 0.35467 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030715 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:54 | 20 | 20 |  | 177.4027 | [19 empty] | -1.2112 | [5 empty] | 0.35481 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030727 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.4716 | [19 empty] | -1.2112 | [5 empty] | 0.35494 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.0003074 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.4716 | [19 empty] | -1.2112 | [5 empty] | 0.35494 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.0003074 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.4716 | [19 empty] | -1.2112 | [5 empty] | 0.35494 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.0003074 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.5406 | [19 empty] | -1.2112 | [5 empty] | 0.35508 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030753 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.5406 | [19 empty] | -1.2112 | [5 empty] | 0.35508 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030753 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.6096 | [19 empty] | -1.2111 | [5 empty] | 0.35522 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030765 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.6786 | [19 empty] | -1.2112 | [5 empty] | 0.35536 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030778 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.6786 | [19 empty] | -1.2112 | [5 empty] | 0.35536 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030778 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.7475 | [19 empty] | -1.2112 | [5 empty] | 0.3555 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030791 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:55 | 20 | 20 |  | 177.7475 | [19 empty] | -1.2112 | [5 empty] | 0.3555 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030791 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.8165 | [19 empty] | -1.2112 | [5 empty] | 0.35563 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030803 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.8165 | [19 empty] | -1.2112 | [5 empty] | 0.35563 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030803 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.8855 | [19 empty] | -1.2112 | [5 empty] | 0.35577 | -1.2102 | 0.90444 | 17.6664 | 0.23773 | -0.00030816 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.8855 | [19 empty] | -1.2111 | [5 empty] | 0.35577 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030816 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.8855 | [19 empty] | -1.2112 | [5 empty] | 0.35577 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030816 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.9545 | [19 empty] | -1.2111 | [5 empty] | 0.35591 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030829 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.9545 | [19 empty] | -1.2111 | [5 empty] | 0.35591 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030829 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 177.9545 | [19 empty] | -1.2112 | [5 empty] | 0.35591 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030829 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 178.0234 | [19 empty] | -1.2112 | [5 empty] | 0.35605 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030841 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:56 | 20 | 20 |  | 178.0234 | [19 empty] | -1.2112 | [5 empty] | 0.35605 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030841 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.0924 | [19 empty] | -1.2112 | [5 empty] | 0.35618 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030854 | -1.4476 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.0924 | [19 empty] | -1.2112 | [5 empty] | 0.35618 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030854 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.1614 | [19 empty] | -1.2112 | [5 empty] | 0.35632 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030867 | -1.4475 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.1614 | [19 empty] | -1.2112 | [5 empty] | 0.35632 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030867 | -1.4475 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.1614 | [19 empty] | -1.2111 | [5 empty] | 0.35632 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030867 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.2304 | [19 empty] | -1.2111 | [5 empty] | 0.35646 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030879 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.2304 | [19 empty] | -1.2111 | [5 empty] | 0.35646 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030879 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.2304 | [19 empty] | -1.2111 | [5 empty] | 0.35646 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030879 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.2993 | [19 empty] | -1.2111 | [5 empty] | 0.3566 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030892 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.2304 | [19 empty] | -1.2112 | [5 empty] | 0.35646 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030879 | -1.4475 | -1.4282 | -1.4282 | -1.4282 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:57 | 20 | 20 |  | 178.2993 | [19 empty] | -1.2111 | [5 empty] | 0.3566 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030892 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.3683 | [19 empty] | -1.2111 | [5 empty] | 0.35674 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030905 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.3683 | [19 empty] | -1.2111 | [5 empty] | 0.35674 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030905 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.3683 | [19 empty] | -1.2111 | [5 empty] | 0.35674 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030905 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.3683 | [19 empty] | -1.2111 | [5 empty] | 0.35674 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030905 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.4373 | [19 empty] | -1.2111 | [5 empty] | 0.35687 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030917 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.4373 | [19 empty] | -1.2112 | [5 empty] | 0.35687 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030917 | -1.4475 | -1.4283 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.4373 | [19 empty] | -1.2111 | [5 empty] | 0.35687 | -1.2101 | 0.90444 | 17.6664 | 0.23773 | -0.00030917 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 178.4373 | [19 empty] | -1.211 | [5 empty] | 0.35687 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00030917 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 170.9191 | [19 empty] | -1.211 | [5 empty] | 0.34184 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029537 | -1.4474 | -1.4281 | -1.4281 | -1.4281 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 13:58 | 20 | 20 |  | 170.9191 | [19 empty] | -1.211 | [5 empty] | 0.34184 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029537 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 170.988 | [19 empty] | -1.2111 | [5 empty] | 0.34198 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.0002955 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.057 | [19 empty] | -1.211 | [5 empty] | 0.34211 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029563 | -1.4474 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.126 | [19 empty] | -1.2111 | [5 empty] | 0.34225 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029575 | -1.4474 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.126 | [19 empty] | -1.211 | [5 empty] | 0.34225 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029575 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.195 | [19 empty] | -1.211 | [5 empty] | 0.34239 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029588 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.195 | [19 empty] | -1.211 | [5 empty] | 0.34239 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029588 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.2639 | [19 empty] | -1.2111 | [5 empty] | 0.34253 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029601 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.3329 | [19 empty] | -1.211 | [5 empty] | 0.34267 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029613 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.3329 | [19 empty] | -1.211 | [5 empty] | 0.34267 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029613 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 13:59 | 20 | 20 |  | 171.3329 | [19 empty] | -1.211 | [5 empty] | 0.34267 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029613 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.3329 | [19 empty] | -1.2111 | [5 empty] | 0.34267 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029613 | -1.4475 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.4019 | [19 empty] | -1.211 | [5 empty] | 0.3428 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029626 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.4019 | [19 empty] | -1.211 | [5 empty] | 0.3428 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029626 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.4709 | [19 empty] | -1.2111 | [5 empty] | 0.34294 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029639 | -1.4474 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.4709 | [19 empty] | -1.211 | [5 empty] | 0.34294 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029639 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.4709 | [19 empty] | -1.211 | [5 empty] | 0.34294 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029639 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.5398 | [19 empty] | -1.211 | [5 empty] | 0.34308 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029651 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.5398 | [19 empty] | -1.211 | [5 empty] | 0.34308 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029651 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.5398 | [19 empty] | -1.211 | [5 empty] | 0.34308 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029651 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.6088 | [19 empty] | -1.211 | [5 empty] | 0.34322 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:00 | 20 | 20 |  | 171.6088 | [19 empty] | -1.211 | [5 empty] | 0.34322 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.6088 | [19 empty] | -1.211 | [5 empty] | 0.34322 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029664 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.6778 | [19 empty] | -1.211 | [5 empty] | 0.34336 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.6778 | [19 empty] | -1.211 | [5 empty] | 0.34336 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.6778 | [19 empty] | -1.211 | [5 empty] | 0.34336 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029677 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.7468 | [19 empty] | -1.211 | [5 empty] | 0.34349 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.7468 | [19 empty] | -1.211 | [5 empty] | 0.34349 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.7468 | [19 empty] | -1.211 | [5 empty] | 0.34349 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029689 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.8157 | [19 empty] | -1.211 | [5 empty] | 0.34363 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.8157 | [19 empty] | -1.211 | [5 empty] | 0.34363 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:01 | 20 | 20 |  | 171.8157 | [19 empty] | -1.211 | [5 empty] | 0.34363 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029702 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 171.8847 | [19 empty] | -1.211 | [5 empty] | 0.34377 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 171.8847 | [19 empty] | -1.211 | [5 empty] | 0.34377 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 171.8847 | [19 empty] | -1.211 | [5 empty] | 0.34377 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029715 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 171.9537 | [19 empty] | -1.211 | [5 empty] | 0.34391 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 171.9537 | [19 empty] | -1.211 | [5 empty] | 0.34391 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 171.9537 | [19 empty] | -1.211 | [5 empty] | 0.34391 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029727 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 172.0227 | [19 empty] | -1.211 | [5 empty] | 0.34405 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 172.0227 | [19 empty] | -1.211 | [5 empty] | 0.34405 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002974 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 172.0916 | [19 empty] | -1.211 | [5 empty] | 0.34418 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 172.0916 | [19 empty] | -1.211 | [5 empty] | 0.34418 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:02 | 20 | 20 |  | 172.0916 | [19 empty] | -1.211 | [5 empty] | 0.34418 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029753 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.1606 | [19 empty] | -1.211 | [5 empty] | 0.34432 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.1606 | [19 empty] | -1.211 | [5 empty] | 0.34432 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.1606 | [19 empty] | -1.211 | [5 empty] | 0.34432 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.1606 | [19 empty] | -1.211 | [5 empty] | 0.34432 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029765 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.2296 | [19 empty] | -1.211 | [5 empty] | 0.34446 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.2296 | [19 empty] | -1.211 | [5 empty] | 0.34446 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029778 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.2986 | [19 empty] | -1.211 | [5 empty] | 0.3446 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.2986 | [19 empty] | -1.2111 | [5 empty] | 0.3446 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4474 | -1.4282 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.2986 | [19 empty] | -1.211 | [5 empty] | 0.3446 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:03 | 20 | 20 |  | 172.2986 | [19 empty] | -1.211 | [5 empty] | 0.3446 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.2986 | [19 empty] | -1.211 | [5 empty] | 0.3446 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029791 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.3675 | [19 empty] | -1.211 | [5 empty] | 0.34474 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.3675 | [19 empty] | -1.211 | [5 empty] | 0.34474 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.3675 | [19 empty] | -1.211 | [5 empty] | 0.34474 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.3675 | [19 empty] | -1.211 | [5 empty] | 0.34474 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.3675 | [19 empty] | -1.211 | [5 empty] | 0.34474 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029803 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.4365 | [19 empty] | -1.211 | [5 empty] | 0.34487 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.4365 | [19 empty] | -1.211 | [5 empty] | 0.34487 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029816 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.5055 | [19 empty] | -1.211 | [5 empty] | 0.34501 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:04 | 20 | 20 |  | 172.5055 | [19 empty] | -1.211 | [5 empty] | 0.34501 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.5055 | [19 empty] | -1.211 | [5 empty] | 0.34501 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.5055 | [19 empty] | -1.211 | [5 empty] | 0.34501 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029829 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.5744 | [19 empty] | -1.211 | [5 empty] | 0.34515 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.5744 | [19 empty] | -1.211 | [5 empty] | 0.34515 | -1.21 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.5744 | [19 empty] | -1.211 | [5 empty] | 0.34515 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.5744 | [19 empty] | -1.2109 | [5 empty] | 0.34515 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029841 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.6434 | [19 empty] | -1.211 | [5 empty] | 0.34529 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2109 | [5 empty] | 0.34529 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2109 | [5 empty] | 0.34529 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:05 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2109 | [5 empty] | 0.34529 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2109 | [5 empty] | 0.34529 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.6434 | [19 empty] | -1.2109 | [5 empty] | 0.34529 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029854 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7124 | [19 empty] | -1.2109 | [5 empty] | 0.34542 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7124 | [19 empty] | -1.2109 | [5 empty] | 0.34542 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7124 | [19 empty] | -1.211 | [5 empty] | 0.34542 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7124 | [19 empty] | -1.211 | [5 empty] | 0.34542 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029867 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7814 | [19 empty] | -1.2109 | [5 empty] | 0.34556 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029879 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7814 | [19 empty] | -1.2109 | [5 empty] | 0.34556 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029879 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.7814 | [19 empty] | -1.211 | [5 empty] | 0.34556 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029879 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.8503 | [19 empty] | -1.211 | [5 empty] | 0.3457 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:06 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2109 | [5 empty] | 0.3457 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2109 | [5 empty] | 0.3457 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.8503 | [19 empty] | -1.2109 | [5 empty] | 0.3457 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029892 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9193 | [19 empty] | -1.211 | [5 empty] | 0.34584 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2109 | [5 empty] | 0.34584 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2109 | [5 empty] | 0.34584 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2109 | [5 empty] | 0.34584 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2109 | [5 empty] | 0.34584 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9193 | [19 empty] | -1.2109 | [5 empty] | 0.34584 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029904 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2109 | [5 empty] | 0.34598 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:07 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2109 | [5 empty] | 0.34598 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2109 | [5 empty] | 0.34598 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2109 | [5 empty] | 0.34598 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2109 | [5 empty] | 0.34598 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 172.9883 | [19 empty] | -1.2109 | [5 empty] | 0.34598 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029917 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.0573 | [19 empty] | -1.2109 | [5 empty] | 0.34611 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.0573 | [19 empty] | -1.211 | [5 empty] | 0.34611 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.0573 | [19 empty] | -1.211 | [5 empty] | 0.34611 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.0573 | [19 empty] | -1.211 | [5 empty] | 0.34611 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002993 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.1262 | [19 empty] | -1.211 | [5 empty] | 0.34625 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4473 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2109 | [5 empty] | 0.34625 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:08 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2109 | [5 empty] | 0.34625 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2109 | [5 empty] | 0.34625 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1262 | [19 empty] | -1.2109 | [5 empty] | 0.34625 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029942 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2109 | [5 empty] | 0.34639 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2109 | [5 empty] | 0.34639 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2109 | [5 empty] | 0.34639 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2109 | [5 empty] | 0.34639 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.1952 | [19 empty] | -1.2109 | [5 empty] | 0.34639 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029955 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2109 | [5 empty] | 0.34653 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2109 | [5 empty] | 0.34653 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:09 | 20 | 20 |  | 173.2642 | [19 empty] | -1.2109 | [5 empty] | 0.34653 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029968 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2109 | [5 empty] | 0.34667 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2109 | [5 empty] | 0.34667 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2109 | [5 empty] | 0.34667 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2109 | [5 empty] | 0.34667 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.3332 | [19 empty] | -1.2109 | [5 empty] | 0.34667 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0002998 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.4021 | [19 empty] | -1.211 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2109 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2109 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.4021 | [19 empty] | -1.211 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:10 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2109 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2109 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4021 | [19 empty] | -1.211 | [5 empty] | 0.3468 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4474 | -1.4281 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2109 | [5 empty] | 0.34694 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4021 | [19 empty] | -1.2109 | [5 empty] | 0.3468 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00029993 | -1.4473 | -1.428 | -1.428 | -1.428 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2109 | [5 empty] | 0.34694 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2109 | [5 empty] | 0.34694 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2109 | [5 empty] | 0.34694 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2109 | [5 empty] | 0.34694 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.4711 | [19 empty] | -1.2109 | [5 empty] | 0.34694 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030006 | -1.4473 | -1.428 | -1.428 | -1.428 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:11 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.5401 | [19 empty] | -1.2109 | [5 empty] | 0.34708 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030018 | -1.4473 | -1.428 | -1.428 | -1.428 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:12 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.6091 | [19 empty] | -1.2109 | [5 empty] | 0.34722 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030031 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.678 | [19 empty] | -1.2109 | [5 empty] | 0.34736 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.678 | [19 empty] | -1.2109 | [5 empty] | 0.34736 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.678 | [19 empty] | -1.2109 | [5 empty] | 0.34736 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.678 | [19 empty] | -1.2109 | [5 empty] | 0.34736 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.678 | [19 empty] | -1.2109 | [5 empty] | 0.34736 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.678 | [19 empty] | -1.2109 | [5 empty] | 0.34736 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030044 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.747 | [19 empty] | -1.2108 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:13 | 20 | 20 |  | 173.747 | [19 empty] | -1.2109 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.747 | [19 empty] | -1.2109 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.747 | [19 empty] | -1.2109 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.747 | [19 empty] | -1.2109 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.747 | [19 empty] | -1.2109 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.747 | [19 empty] | -1.2109 | [5 empty] | 0.34749 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030056 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.816 | [19 empty] | -1.2109 | [5 empty] | 0.34763 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030069 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.816 | [19 empty] | -1.2108 | [5 empty] | 0.34763 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030069 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.816 | [19 empty] | -1.2108 | [5 empty] | 0.34763 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030069 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.885 | [19 empty] | -1.2109 | [5 empty] | 0.34777 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.885 | [19 empty] | -1.2109 | [5 empty] | 0.34777 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:14 | 20 | 20 |  | 173.885 | [19 empty] | -1.2109 | [5 empty] | 0.34777 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.885 | [19 empty] | -1.2109 | [5 empty] | 0.34777 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.885 | [19 empty] | -1.2109 | [5 empty] | 0.34777 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030082 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2109 | [5 empty] | 0.34791 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2109 | [5 empty] | 0.34791 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2108 | [5 empty] | 0.34791 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2109 | [5 empty] | 0.34791 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2109 | [5 empty] | 0.34791 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2109 | [5 empty] | 0.34791 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 173.9539 | [19 empty] | -1.2109 | [5 empty] | 0.34791 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030094 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:15 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2109 | [5 empty] | 0.34805 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2109 | [5 empty] | 0.34805 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2108 | [5 empty] | 0.34805 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2109 | [5 empty] | 0.34805 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2109 | [5 empty] | 0.34805 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.211 | [5 empty] | 0.34805 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2109 | [5 empty] | 0.34805 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0229 | [19 empty] | -1.2109 | [5 empty] | 0.34805 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030107 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0919 | [19 empty] | -1.2109 | [5 empty] | 0.34818 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003012 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0919 | [19 empty] | -1.2109 | [5 empty] | 0.34818 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003012 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:16 | 20 | 20 |  | 174.0919 | [19 empty] | -1.2108 | [5 empty] | 0.34818 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003012 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.0919 | [19 empty] | -1.2109 | [5 empty] | 0.34818 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003012 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2108 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2108 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2108 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:17 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2108 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.1609 | [19 empty] | -1.2109 | [5 empty] | 0.34832 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030132 | -1.4472 | -1.4279 | -1.4279 | -1.4279 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:18 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2108 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2108 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2108 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2108 | [5 empty] | 0.34846 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2108 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:19 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4473 | -1.428 | -1.428 | -1.428 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2108 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2108 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2298 | [19 empty] | -1.2109 | [5 empty] | 0.34846 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030145 | -1.4472 | -1.4279 | -1.4279 | -1.4279 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2108 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2108 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:20 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | -1.4279 | -1.4279 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2108 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2108 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4473 | -1.428 | -1.428 | -1.428 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.2988 | [19 empty] | -1.2109 | [5 empty] | 0.3486 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030158 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2108 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:21 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2108 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2108 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2108 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.3678 | [19 empty] | -1.2109 | [5 empty] | 0.34874 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.0003017 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:22 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2108 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2108 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2109 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2108 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.4368 | [19 empty] | -1.2108 | [5 empty] | 0.34887 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030183 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2108 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:23 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2108 | [5 empty] | 0.34901 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4471 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2108 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2109 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2109 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2108 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2108 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2109 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5057 | [19 empty] | -1.2108 | [5 empty] | 0.34901 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030196 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:24 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4471 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2109 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2109 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.5747 | [19 empty] | -1.2108 | [5 empty] | 0.34915 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030208 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:25 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.6437 | [19 empty] | -1.2108 | [5 empty] | 0.34929 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030221 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:26 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2109 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2109 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7127 | [19 empty] | -1.2108 | [5 empty] | 0.34943 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030234 | -1.4471 | -1.4279 | -1.4279 | -1.4279 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2107 | [5 empty] | 0.34956 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4471 | -1.4278 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:27 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2109 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2109 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.7816 | [19 empty] | -1.2108 | [5 empty] | 0.34956 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030246 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.8506 | [19 empty] | -1.2108 | [5 empty] | 0.3497 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030259 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:28 | 20 | 20 |  | 174.8506 | [19 empty] | -1.2108 | [5 empty] | 0.3497 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030259 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.8506 | [19 empty] | -1.2108 | [5 empty] | 0.3497 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030259 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.8506 | [19 empty] | -1.2108 | [5 empty] | 0.3497 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030259 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.8506 | [19 empty] | -1.2108 | [5 empty] | 0.3497 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030259 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4471 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:29 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:30 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2109 | [5 empty] | 0.34984 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9196 | [19 empty] | -1.2108 | [5 empty] | 0.34984 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030272 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2108 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:31 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2099 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4473 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2108 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 175.0575 | [19 empty] | -1.2108 | [5 empty] | 0.35012 | -1.2097 | 0.90444 | 17.6664 | 0.23773 | -0.00030297 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 174.9886 | [19 empty] | -1.2109 | [5 empty] | 0.34998 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030284 | -1.4472 | -1.4279 | -1.4279 | -1.4279 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 175.0575 | [19 empty] | -1.2108 | [5 empty] | 0.35012 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030297 | -1.4472 | -1.4279 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 175.0575 | [19 empty] | -1.2109 | [5 empty] | 0.35012 | -1.2098 | 0.90444 | 17.6664 | 0.23773 | -0.00030297 | -1.4472 | -1.428 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 175.0575 | [19 empty] | -1.2096 | [5 empty] | 0.35012 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00030297 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 76.3549 | [19 empty] | -1.2096 | [5 empty] | 0.15271 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001218 | -1.4461 | -1.4268 | -1.4268 | -1.4268 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:32 | 20 | 20 |  | 65.181 | [19 empty] | -1.2095 | [5 empty] | 0.13036 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00010129 | -1.4461 | -1.4268 | -1.4268 | -1.4268 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 64.6292 | [19 empty] | -1.2095 | [5 empty] | 0.12926 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010027 | -1.4461 | -1.4268 | -1.4268 | -1.4268 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 65.25 | [19 empty] | -1.2095 | [5 empty] | 0.1305 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010141 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 65.6638 | [19 empty] | -1.2095 | [5 empty] | 0.13133 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010217 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 66.0087 | [19 empty] | -1.2095 | [5 empty] | 0.13202 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010281 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 66.3536 | [19 empty] | -1.2096 | [5 empty] | 0.13271 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010344 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 66.6984 | [19 empty] | -1.2096 | [5 empty] | 0.1334 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010407 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 66.9054 | [19 empty] | -1.2096 | [5 empty] | 0.13381 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00010445 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 67.1123 | [19 empty] | -1.2096 | [5 empty] | 0.13422 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00010483 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 67.3192 | [19 empty] | -1.2096 | [5 empty] | 0.13464 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010521 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:33 | 20 | 20 |  | 67.5261 | [19 empty] | -1.2096 | [5 empty] | 0.13505 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010559 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 67.6641 | [19 empty] | -1.2096 | [5 empty] | 0.13533 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010584 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 67.871 | [19 empty] | -1.2096 | [5 empty] | 0.13574 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010622 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.009 | [19 empty] | -1.2096 | [5 empty] | 0.13602 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010648 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.0779 | [19 empty] | -1.2096 | [5 empty] | 0.13616 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001066 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.2849 | [19 empty] | -1.2096 | [5 empty] | 0.13657 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010698 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.3538 | [19 empty] | -1.2096 | [5 empty] | 0.13671 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010711 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.4228 | [19 empty] | -1.2096 | [5 empty] | 0.13685 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010724 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.5608 | [19 empty] | -1.2096 | [5 empty] | 0.13712 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00010749 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.6297 | [19 empty] | -1.2096 | [5 empty] | 0.13726 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010762 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.6987 | [19 empty] | -1.2096 | [5 empty] | 0.1374 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010774 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:34 | 20 | 20 |  | 68.8367 | [19 empty] | -1.2096 | [5 empty] | 0.13767 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.000108 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 68.9056 | [19 empty] | -1.2096 | [5 empty] | 0.13781 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010812 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 68.9746 | [19 empty] | -1.2096 | [5 empty] | 0.13795 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010825 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.0436 | [19 empty] | -1.2096 | [5 empty] | 0.13809 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010838 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.1815 | [19 empty] | -1.2096 | [5 empty] | 0.13836 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010863 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.2505 | [19 empty] | -1.2096 | [5 empty] | 0.1385 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010876 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.2505 | [19 empty] | -1.2096 | [5 empty] | 0.1385 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010876 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.3195 | [19 empty] | -1.2096 | [5 empty] | 0.13864 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010888 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.3884 | [19 empty] | -1.2096 | [5 empty] | 0.13878 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010901 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.4574 | [19 empty] | -1.2096 | [5 empty] | 0.13891 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00010914 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:35 | 20 | 20 |  | 69.5264 | [19 empty] | -1.2096 | [5 empty] | 0.13905 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00010926 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.5954 | [19 empty] | -1.2096 | [5 empty] | 0.13919 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010939 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.6643 | [19 empty] | -1.2096 | [5 empty] | 0.13933 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010952 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.6643 | [19 empty] | -1.2096 | [5 empty] | 0.13933 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010952 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.7333 | [19 empty] | -1.2096 | [5 empty] | 0.13947 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010964 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.8023 | [19 empty] | -1.2096 | [5 empty] | 0.1396 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00010977 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.8023 | [19 empty] | -1.2096 | [5 empty] | 0.1396 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00010977 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.8713 | [19 empty] | -1.2096 | [5 empty] | 0.13974 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001099 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 69.9402 | [19 empty] | -1.2096 | [5 empty] | 0.13988 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011002 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 70.0092 | [19 empty] | -1.2096 | [5 empty] | 0.14002 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011015 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 70.0782 | [19 empty] | -1.2096 | [5 empty] | 0.14016 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011028 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:36 | 20 | 20 |  | 70.0782 | [19 empty] | -1.2096 | [5 empty] | 0.14016 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00011028 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.1472 | [19 empty] | -1.2097 | [5 empty] | 0.14029 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.0001104 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.2161 | [19 empty] | -1.2096 | [5 empty] | 0.14043 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00011053 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.2161 | [19 empty] | -1.2096 | [5 empty] | 0.14043 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011053 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.2851 | [19 empty] | -1.2096 | [5 empty] | 0.14057 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00011066 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.3541 | [19 empty] | -1.2096 | [5 empty] | 0.14071 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011078 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.3541 | [19 empty] | -1.2096 | [5 empty] | 0.14071 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011078 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.4231 | [19 empty] | -1.2096 | [5 empty] | 0.14085 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011091 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.4231 | [19 empty] | -1.2095 | [5 empty] | 0.14085 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011091 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.492 | [19 empty] | -1.2095 | [5 empty] | 0.14098 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011104 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:37 | 20 | 20 |  | 70.492 | [19 empty] | -1.2096 | [5 empty] | 0.14098 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011104 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.561 | [19 empty] | -1.2096 | [5 empty] | 0.14112 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011116 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.63 | [19 empty] | -1.2096 | [5 empty] | 0.14126 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011129 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.63 | [19 empty] | -1.2095 | [5 empty] | 0.14126 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011129 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.699 | [19 empty] | -1.2096 | [5 empty] | 0.1414 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011142 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.699 | [19 empty] | -1.2095 | [5 empty] | 0.1414 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011142 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.7679 | [19 empty] | -1.2095 | [5 empty] | 0.14154 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011154 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.7679 | [19 empty] | -1.2096 | [5 empty] | 0.14154 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011154 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.8369 | [19 empty] | -1.2096 | [5 empty] | 0.14167 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011167 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.8369 | [19 empty] | -1.2095 | [5 empty] | 0.14167 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011167 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:38 | 20 | 20 |  | 70.9059 | [19 empty] | -1.2096 | [5 empty] | 0.14181 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001118 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 70.9059 | [19 empty] | -1.2096 | [5 empty] | 0.14181 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001118 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 70.9059 | [19 empty] | -1.2096 | [5 empty] | 0.14181 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001118 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 70.9749 | [19 empty] | -1.2096 | [5 empty] | 0.14195 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011192 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 70.9749 | [19 empty] | -1.2096 | [5 empty] | 0.14195 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011192 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.0438 | [19 empty] | -1.2096 | [5 empty] | 0.14209 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011205 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.1128 | [19 empty] | -1.2096 | [5 empty] | 0.14223 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011218 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.1128 | [19 empty] | -1.2096 | [5 empty] | 0.14223 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011218 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.1818 | [19 empty] | -1.2096 | [5 empty] | 0.14236 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.0001123 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.1818 | [19 empty] | -1.2096 | [5 empty] | 0.14236 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001123 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.1818 | [19 empty] | -1.2096 | [5 empty] | 0.14236 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001123 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:39 | 20 | 20 |  | 71.2508 | [19 empty] | -1.2096 | [5 empty] | 0.1425 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011243 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.2508 | [19 empty] | -1.2096 | [5 empty] | 0.1425 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011243 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.3197 | [19 empty] | -1.2096 | [5 empty] | 0.14264 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011255 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.3197 | [19 empty] | -1.2096 | [5 empty] | 0.14264 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011255 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.3887 | [19 empty] | -1.2095 | [5 empty] | 0.14278 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011268 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.3887 | [19 empty] | -1.2096 | [5 empty] | 0.14278 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011268 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.4577 | [19 empty] | -1.2096 | [5 empty] | 0.14292 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011281 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.4577 | [19 empty] | -1.2096 | [5 empty] | 0.14292 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011281 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.5267 | [19 empty] | -1.2096 | [5 empty] | 0.14305 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011293 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.5267 | [19 empty] | -1.2096 | [5 empty] | 0.14305 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011293 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:40 | 20 | 20 |  | 71.5956 | [19 empty] | -1.2096 | [5 empty] | 0.14319 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011306 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.5956 | [19 empty] | -1.2096 | [5 empty] | 0.14319 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011306 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.6646 | [19 empty] | -1.2095 | [5 empty] | 0.14333 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011319 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.6646 | [19 empty] | -1.2096 | [5 empty] | 0.14333 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011319 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.7336 | [19 empty] | -1.2096 | [5 empty] | 0.14347 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011331 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.7336 | [19 empty] | -1.2096 | [5 empty] | 0.14347 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011331 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.8026 | [19 empty] | -1.2096 | [5 empty] | 0.14361 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011344 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.8026 | [19 empty] | -1.2095 | [5 empty] | 0.14361 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011344 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.8715 | [19 empty] | -1.2096 | [5 empty] | 0.14374 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011357 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.8715 | [19 empty] | -1.2096 | [5 empty] | 0.14374 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011357 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:41 | 20 | 20 |  | 71.8715 | [19 empty] | -1.2096 | [5 empty] | 0.14374 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011357 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 71.9405 | [19 empty] | -1.2095 | [5 empty] | 0.14388 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011369 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 71.9405 | [19 empty] | -1.2095 | [5 empty] | 0.14388 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011369 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 71.9405 | [19 empty] | -1.2096 | [5 empty] | 0.14388 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011369 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.0095 | [19 empty] | -1.2096 | [5 empty] | 0.14402 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011382 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.0095 | [19 empty] | -1.2095 | [5 empty] | 0.14402 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011382 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.0095 | [19 empty] | -1.2096 | [5 empty] | 0.14402 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011382 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.0785 | [19 empty] | -1.2096 | [5 empty] | 0.14416 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011395 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.0785 | [19 empty] | -1.2095 | [5 empty] | 0.14416 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011395 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.1474 | [19 empty] | -1.2096 | [5 empty] | 0.14429 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.1474 | [19 empty] | -1.2096 | [5 empty] | 0.14429 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011407 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:42 | 20 | 20 |  | 72.2164 | [19 empty] | -1.2096 | [5 empty] | 0.14443 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.2164 | [19 empty] | -1.2096 | [5 empty] | 0.14443 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001142 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.2854 | [19 empty] | -1.2095 | [5 empty] | 0.14457 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.2854 | [19 empty] | -1.2096 | [5 empty] | 0.14457 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.2854 | [19 empty] | -1.2095 | [5 empty] | 0.14457 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011433 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.3544 | [19 empty] | -1.2096 | [5 empty] | 0.14471 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.3544 | [19 empty] | -1.2095 | [5 empty] | 0.14471 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011445 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.4233 | [19 empty] | -1.2096 | [5 empty] | 0.14485 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.4233 | [19 empty] | -1.2096 | [5 empty] | 0.14485 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.4233 | [19 empty] | -1.2096 | [5 empty] | 0.14485 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011458 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:43 | 20 | 20 |  | 72.4923 | [19 empty] | -1.2096 | [5 empty] | 0.14498 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.4923 | [19 empty] | -1.2095 | [5 empty] | 0.14498 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.4923 | [19 empty] | -1.2095 | [5 empty] | 0.14498 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011471 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.5613 | [19 empty] | -1.2096 | [5 empty] | 0.14512 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.5613 | [19 empty] | -1.2097 | [5 empty] | 0.14512 | -1.2086 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.5613 | [19 empty] | -1.2095 | [5 empty] | 0.14512 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.5613 | [19 empty] | -1.2096 | [5 empty] | 0.14512 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011483 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.6303 | [19 empty] | -1.2095 | [5 empty] | 0.14526 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.6303 | [19 empty] | -1.2096 | [5 empty] | 0.14526 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011496 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.6992 | [19 empty] | -1.2096 | [5 empty] | 0.1454 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.6992 | [19 empty] | -1.2095 | [5 empty] | 0.1454 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:44 | 20 | 20 |  | 72.6992 | [19 empty] | -1.2095 | [5 empty] | 0.1454 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011509 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.7682 | [19 empty] | -1.2096 | [5 empty] | 0.14554 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.7682 | [19 empty] | -1.2096 | [5 empty] | 0.14554 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011521 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.8372 | [19 empty] | -1.2096 | [5 empty] | 0.14567 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.8372 | [19 empty] | -1.2095 | [5 empty] | 0.14567 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.8372 | [19 empty] | -1.2096 | [5 empty] | 0.14567 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011534 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.9062 | [19 empty] | -1.2095 | [5 empty] | 0.14581 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.9062 | [19 empty] | -1.2095 | [5 empty] | 0.14581 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011547 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.9751 | [19 empty] | -1.2096 | [5 empty] | 0.14595 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4462 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 72.9751 | [19 empty] | -1.2096 | [5 empty] | 0.14595 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011559 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:45 | 20 | 20 |  | 73.0441 | [19 empty] | -1.2096 | [5 empty] | 0.14609 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.0441 | [19 empty] | -1.2096 | [5 empty] | 0.14609 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.0441 | [19 empty] | -1.2096 | [5 empty] | 0.14609 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4461 | -1.4269 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.0441 | [19 empty] | -1.2095 | [5 empty] | 0.14609 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011572 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1131 | [19 empty] | -1.2096 | [5 empty] | 0.14623 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1131 | [19 empty] | -1.2096 | [5 empty] | 0.14623 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1131 | [19 empty] | -1.2095 | [5 empty] | 0.14623 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1131 | [19 empty] | -1.2095 | [5 empty] | 0.14623 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011585 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1821 | [19 empty] | -1.2095 | [5 empty] | 0.14636 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1821 | [19 empty] | -1.2096 | [5 empty] | 0.14636 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:46 | 20 | 20 |  | 73.1821 | [19 empty] | -1.2095 | [5 empty] | 0.14636 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011597 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.251 | [19 empty] | -1.2095 | [5 empty] | 0.1465 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.251 | [19 empty] | -1.2095 | [5 empty] | 0.1465 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.251 | [19 empty] | -1.2095 | [5 empty] | 0.1465 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.251 | [19 empty] | -1.2096 | [5 empty] | 0.1465 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.251 | [19 empty] | -1.2096 | [5 empty] | 0.1465 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.251 | [19 empty] | -1.2095 | [5 empty] | 0.1465 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001161 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.32 | [19 empty] | -1.2095 | [5 empty] | 0.14664 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.32 | [19 empty] | -1.2095 | [5 empty] | 0.14664 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.32 | [19 empty] | -1.2095 | [5 empty] | 0.14664 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:47 | 20 | 20 |  | 73.32 | [19 empty] | -1.2095 | [5 empty] | 0.14664 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011623 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.389 | [19 empty] | -1.2095 | [5 empty] | 0.14678 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.389 | [19 empty] | -1.2095 | [5 empty] | 0.14678 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.389 | [19 empty] | -1.2095 | [5 empty] | 0.14678 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.389 | [19 empty] | -1.2096 | [5 empty] | 0.14678 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011635 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.4579 | [19 empty] | -1.2096 | [5 empty] | 0.14692 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.4579 | [19 empty] | -1.2096 | [5 empty] | 0.14692 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.4579 | [19 empty] | -1.2095 | [5 empty] | 0.14692 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011648 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.5269 | [19 empty] | -1.2095 | [5 empty] | 0.14705 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.5269 | [19 empty] | -1.2095 | [5 empty] | 0.14705 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.5269 | [19 empty] | -1.2095 | [5 empty] | 0.14705 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011661 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:48 | 20 | 20 |  | 73.5959 | [19 empty] | -1.2095 | [5 empty] | 0.14719 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.5959 | [19 empty] | -1.2095 | [5 empty] | 0.14719 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.5959 | [19 empty] | -1.2095 | [5 empty] | 0.14719 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011673 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.6649 | [19 empty] | -1.2095 | [5 empty] | 0.14733 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011686 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.6649 | [19 empty] | -1.2095 | [5 empty] | 0.14733 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011686 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.6649 | [19 empty] | -1.2095 | [5 empty] | 0.14733 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011686 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.7338 | [19 empty] | -1.2095 | [5 empty] | 0.14747 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.7338 | [19 empty] | -1.2095 | [5 empty] | 0.14747 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011699 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.8028 | [19 empty] | -1.2095 | [5 empty] | 0.14761 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.8028 | [19 empty] | -1.2095 | [5 empty] | 0.14761 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:49 | 20 | 20 |  | 73.8028 | [19 empty] | -1.2095 | [5 empty] | 0.14761 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.8028 | [19 empty] | -1.2095 | [5 empty] | 0.14761 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.8028 | [19 empty] | -1.2095 | [5 empty] | 0.14761 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011711 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.8718 | [19 empty] | -1.2095 | [5 empty] | 0.14774 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.8718 | [19 empty] | -1.2095 | [5 empty] | 0.14774 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.8718 | [19 empty] | -1.2095 | [5 empty] | 0.14774 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.8718 | [19 empty] | -1.2095 | [5 empty] | 0.14774 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011724 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.9408 | [19 empty] | -1.2095 | [5 empty] | 0.14788 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.9408 | [19 empty] | -1.2095 | [5 empty] | 0.14788 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 73.9408 | [19 empty] | -1.2095 | [5 empty] | 0.14788 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011737 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 74.0097 | [19 empty] | -1.2095 | [5 empty] | 0.14802 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011749 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:50 | 20 | 20 |  | 74.0097 | [19 empty] | -1.2096 | [5 empty] | 0.14802 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011749 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.0787 | [19 empty] | -1.2095 | [5 empty] | 0.14816 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.0787 | [19 empty] | -1.2095 | [5 empty] | 0.14816 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.0787 | [19 empty] | -1.2095 | [5 empty] | 0.14816 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.0787 | [19 empty] | -1.2095 | [5 empty] | 0.14816 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.0787 | [19 empty] | -1.2095 | [5 empty] | 0.14816 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.0787 | [19 empty] | -1.2095 | [5 empty] | 0.14816 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011762 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.1477 | [19 empty] | -1.2095 | [5 empty] | 0.1483 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.1477 | [19 empty] | -1.2096 | [5 empty] | 0.1483 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.1477 | [19 empty] | -1.2096 | [5 empty] | 0.1483 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:51 | 20 | 20 |  | 74.1477 | [19 empty] | -1.2095 | [5 empty] | 0.1483 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.1477 | [19 empty] | -1.2095 | [5 empty] | 0.1483 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011775 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.2167 | [19 empty] | -1.2095 | [5 empty] | 0.14843 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011787 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.2167 | [19 empty] | -1.2095 | [5 empty] | 0.14843 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011787 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.2856 | [19 empty] | -1.2095 | [5 empty] | 0.14857 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.000118 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.2856 | [19 empty] | -1.2095 | [5 empty] | 0.14857 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.000118 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.2856 | [19 empty] | -1.2095 | [5 empty] | 0.14857 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.000118 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.3546 | [19 empty] | -1.2095 | [5 empty] | 0.14871 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.3546 | [19 empty] | -1.2095 | [5 empty] | 0.14871 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.3546 | [19 empty] | -1.2094 | [5 empty] | 0.14871 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:52 | 20 | 20 |  | 74.3546 | [19 empty] | -1.2095 | [5 empty] | 0.14871 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.3546 | [19 empty] | -1.2095 | [5 empty] | 0.14871 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.3546 | [19 empty] | -1.2095 | [5 empty] | 0.14871 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011813 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4236 | [19 empty] | -1.2095 | [5 empty] | 0.14885 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4236 | [19 empty] | -1.2095 | [5 empty] | 0.14885 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4236 | [19 empty] | -1.2095 | [5 empty] | 0.14885 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4236 | [19 empty] | -1.2095 | [5 empty] | 0.14885 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4236 | [19 empty] | -1.2095 | [5 empty] | 0.14885 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011825 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4926 | [19 empty] | -1.2095 | [5 empty] | 0.14899 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011838 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4926 | [19 empty] | -1.2095 | [5 empty] | 0.14899 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011838 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4926 | [19 empty] | -1.2095 | [5 empty] | 0.14899 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011838 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:53 | 20 | 20 |  | 74.4926 | [19 empty] | -1.2095 | [5 empty] | 0.14899 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011838 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.5615 | [19 empty] | -1.2095 | [5 empty] | 0.14912 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.5615 | [19 empty] | -1.2095 | [5 empty] | 0.14912 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.5615 | [19 empty] | -1.2095 | [5 empty] | 0.14912 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.5615 | [19 empty] | -1.2095 | [5 empty] | 0.14912 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.5615 | [19 empty] | -1.2095 | [5 empty] | 0.14912 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011851 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.6305 | [19 empty] | -1.2095 | [5 empty] | 0.14926 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011863 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.6305 | [19 empty] | -1.2095 | [5 empty] | 0.14926 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011863 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.6305 | [19 empty] | -1.2095 | [5 empty] | 0.14926 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011863 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.6305 | [19 empty] | -1.2095 | [5 empty] | 0.14926 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011863 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:54 | 20 | 20 |  | 74.6995 | [19 empty] | -1.2095 | [5 empty] | 0.1494 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011876 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.6995 | [19 empty] | -1.2095 | [5 empty] | 0.1494 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011876 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.7685 | [19 empty] | -1.2095 | [5 empty] | 0.14954 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.7685 | [19 empty] | -1.2096 | [5 empty] | 0.14954 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.7685 | [19 empty] | -1.2095 | [5 empty] | 0.14954 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.7685 | [19 empty] | -1.2095 | [5 empty] | 0.14954 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.7685 | [19 empty] | -1.2095 | [5 empty] | 0.14954 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.7685 | [19 empty] | -1.2095 | [5 empty] | 0.14954 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011889 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.8374 | [19 empty] | -1.2096 | [5 empty] | 0.14967 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.8374 | [19 empty] | -1.2095 | [5 empty] | 0.14967 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:55 | 20 | 20 |  | 74.8374 | [19 empty] | -1.2095 | [5 empty] | 0.14967 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.8374 | [19 empty] | -1.2095 | [5 empty] | 0.14967 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.8374 | [19 empty] | -1.2095 | [5 empty] | 0.14967 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.8374 | [19 empty] | -1.2095 | [5 empty] | 0.14967 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011901 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9064 | [19 empty] | -1.2095 | [5 empty] | 0.14981 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011914 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9064 | [19 empty] | -1.2095 | [5 empty] | 0.14981 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011914 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9064 | [19 empty] | -1.2095 | [5 empty] | 0.14981 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011914 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9754 | [19 empty] | -1.2095 | [5 empty] | 0.14995 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011926 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9754 | [19 empty] | -1.2095 | [5 empty] | 0.14995 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011926 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9754 | [19 empty] | -1.2094 | [5 empty] | 0.14995 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011926 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 74.9754 | [19 empty] | -1.2095 | [5 empty] | 0.14995 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011926 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:56 | 20 | 20 |  | 75.0444 | [19 empty] | -1.2095 | [5 empty] | 0.15009 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011939 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.0444 | [19 empty] | -1.2095 | [5 empty] | 0.15009 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011939 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.0444 | [19 empty] | -1.2095 | [5 empty] | 0.15009 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011939 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.0444 | [19 empty] | -1.2095 | [5 empty] | 0.15009 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011939 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.0444 | [19 empty] | -1.2095 | [5 empty] | 0.15009 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011939 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2096 | [5 empty] | 0.15023 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:57 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.1133 | [19 empty] | -1.2095 | [5 empty] | 0.15023 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011952 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.1823 | [19 empty] | -1.2095 | [5 empty] | 0.15036 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011964 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.1823 | [19 empty] | -1.2095 | [5 empty] | 0.15036 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011964 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.1823 | [19 empty] | -1.2095 | [5 empty] | 0.15036 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011964 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.2513 | [19 empty] | -1.2095 | [5 empty] | 0.1505 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011977 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.2513 | [19 empty] | -1.2095 | [5 empty] | 0.1505 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011977 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.2513 | [19 empty] | -1.2095 | [5 empty] | 0.1505 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011977 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.2513 | [19 empty] | -1.2095 | [5 empty] | 0.1505 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00011977 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:58 | 20 | 20 |  | 75.2513 | [19 empty] | -1.2096 | [5 empty] | 0.1505 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00011977 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3203 | [19 empty] | -1.2095 | [5 empty] | 0.15064 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001199 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3203 | [19 empty] | -1.2096 | [5 empty] | 0.15064 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001199 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3203 | [19 empty] | -1.2096 | [5 empty] | 0.15064 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001199 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3203 | [19 empty] | -1.2095 | [5 empty] | 0.15064 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001199 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3892 | [19 empty] | -1.2095 | [5 empty] | 0.15078 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012002 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3892 | [19 empty] | -1.2095 | [5 empty] | 0.15078 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012002 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3892 | [19 empty] | -1.2095 | [5 empty] | 0.15078 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012002 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3892 | [19 empty] | -1.2095 | [5 empty] | 0.15078 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012002 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.3892 | [19 empty] | -1.2095 | [5 empty] | 0.15078 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012002 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.4582 | [19 empty] | -1.2095 | [5 empty] | 0.15092 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012015 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 14:59 | 20 | 20 |  | 75.4582 | [19 empty] | -1.2095 | [5 empty] | 0.15092 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012015 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.4582 | [19 empty] | -1.2095 | [5 empty] | 0.15092 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012015 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.4582 | [19 empty] | -1.2095 | [5 empty] | 0.15092 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012015 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.4582 | [19 empty] | -1.2095 | [5 empty] | 0.15092 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012015 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5272 | [19 empty] | -1.2095 | [5 empty] | 0.15105 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5272 | [19 empty] | -1.2095 | [5 empty] | 0.15105 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5272 | [19 empty] | -1.2095 | [5 empty] | 0.15105 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5272 | [19 empty] | -1.2095 | [5 empty] | 0.15105 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5272 | [19 empty] | -1.2095 | [5 empty] | 0.15105 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5272 | [19 empty] | -1.2095 | [5 empty] | 0.15105 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012028 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:00 | 20 | 20 |  | 75.5962 | [19 empty] | -1.2095 | [5 empty] | 0.15119 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.5962 | [19 empty] | -1.2095 | [5 empty] | 0.15119 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.5962 | [19 empty] | -1.2095 | [5 empty] | 0.15119 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.5962 | [19 empty] | -1.2095 | [5 empty] | 0.15119 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.5962 | [19 empty] | -1.2096 | [5 empty] | 0.15119 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.5962 | [19 empty] | -1.2095 | [5 empty] | 0.15119 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001204 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.6651 | [19 empty] | -1.2095 | [5 empty] | 0.15133 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012053 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.6651 | [19 empty] | -1.2095 | [5 empty] | 0.15133 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012053 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.6651 | [19 empty] | -1.2095 | [5 empty] | 0.15133 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012053 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.6651 | [19 empty] | -1.2095 | [5 empty] | 0.15133 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012053 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:01 | 20 | 20 |  | 75.7341 | [19 empty] | -1.2095 | [5 empty] | 0.15147 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012066 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.7341 | [19 empty] | -1.2095 | [5 empty] | 0.15147 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012066 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.7341 | [19 empty] | -1.2095 | [5 empty] | 0.15147 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012066 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.7341 | [19 empty] | -1.2095 | [5 empty] | 0.15147 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012066 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.7341 | [19 empty] | -1.2095 | [5 empty] | 0.15147 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012066 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.7341 | [19 empty] | -1.2095 | [5 empty] | 0.15147 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012066 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:02 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8031 | [19 empty] | -1.2095 | [5 empty] | 0.15161 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012078 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8721 | [19 empty] | -1.2095 | [5 empty] | 0.15174 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012091 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8721 | [19 empty] | -1.2095 | [5 empty] | 0.15174 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012091 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8721 | [19 empty] | -1.2094 | [5 empty] | 0.15174 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012091 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8721 | [19 empty] | -1.2095 | [5 empty] | 0.15174 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012091 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.8721 | [19 empty] | -1.2095 | [5 empty] | 0.15174 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012091 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.941 | [19 empty] | -1.2095 | [5 empty] | 0.15188 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012104 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:03 | 20 | 20 |  | 75.941 | [19 empty] | -1.2095 | [5 empty] | 0.15188 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012104 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 75.941 | [19 empty] | -1.2095 | [5 empty] | 0.15188 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012104 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 75.941 | [19 empty] | -1.2095 | [5 empty] | 0.15188 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012104 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.01 | [19 empty] | -1.2095 | [5 empty] | 0.15202 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012116 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:04 | 20 | 20 |  | 76.079 | [19 empty] | -1.2095 | [5 empty] | 0.15216 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012129 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.079 | [19 empty] | -1.2096 | [5 empty] | 0.15216 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012129 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.079 | [19 empty] | -1.2096 | [5 empty] | 0.15216 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012129 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.079 | [19 empty] | -1.2095 | [5 empty] | 0.15216 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012129 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.079 | [19 empty] | -1.2094 | [5 empty] | 0.15216 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012129 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.079 | [19 empty] | -1.2095 | [5 empty] | 0.15216 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012129 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.148 | [19 empty] | -1.2095 | [5 empty] | 0.1523 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012142 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.148 | [19 empty] | -1.2095 | [5 empty] | 0.1523 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012142 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.148 | [19 empty] | -1.2095 | [5 empty] | 0.1523 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012142 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.148 | [19 empty] | -1.2095 | [5 empty] | 0.1523 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012142 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.148 | [19 empty] | -1.2095 | [5 empty] | 0.1523 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012142 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:05 | 20 | 20 |  | 76.148 | [19 empty] | -1.2094 | [5 empty] | 0.1523 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012142 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2169 | [19 empty] | -1.2095 | [5 empty] | 0.15243 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012154 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2169 | [19 empty] | -1.2095 | [5 empty] | 0.15243 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012154 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2169 | [19 empty] | -1.2095 | [5 empty] | 0.15243 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012154 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2169 | [19 empty] | -1.2095 | [5 empty] | 0.15243 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012154 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2169 | [19 empty] | -1.2095 | [5 empty] | 0.15243 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012154 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2859 | [19 empty] | -1.2095 | [5 empty] | 0.15257 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012167 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2859 | [19 empty] | -1.2095 | [5 empty] | 0.15257 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012167 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2859 | [19 empty] | -1.2095 | [5 empty] | 0.15257 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012167 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.2859 | [19 empty] | -1.2095 | [5 empty] | 0.15257 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012167 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:06 | 20 | 20 |  | 76.3549 | [19 empty] | -1.2095 | [5 empty] | 0.15271 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001218 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.3549 | [19 empty] | -1.2095 | [5 empty] | 0.15271 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.3549 | [19 empty] | -1.2095 | [5 empty] | 0.15271 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.3549 | [19 empty] | -1.2095 | [5 empty] | 0.15271 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.3549 | [19 empty] | -1.2095 | [5 empty] | 0.15271 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4239 | [19 empty] | -1.2095 | [5 empty] | 0.15285 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012192 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4239 | [19 empty] | -1.2095 | [5 empty] | 0.15285 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012192 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4239 | [19 empty] | -1.2095 | [5 empty] | 0.15285 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012192 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4239 | [19 empty] | -1.2095 | [5 empty] | 0.15285 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012192 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4239 | [19 empty] | -1.2095 | [5 empty] | 0.15285 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012192 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4928 | [19 empty] | -1.2095 | [5 empty] | 0.15299 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012205 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:07 | 20 | 20 |  | 76.4928 | [19 empty] | -1.2095 | [5 empty] | 0.15299 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012205 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.4928 | [19 empty] | -1.2095 | [5 empty] | 0.15299 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012205 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.4928 | [19 empty] | -1.2094 | [5 empty] | 0.15299 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012205 | -1.4459 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.4928 | [19 empty] | -1.2095 | [5 empty] | 0.15299 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012205 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2095 | [5 empty] | 0.15312 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2094 | [5 empty] | 0.15312 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.4459 | -1.4266 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2094 | [5 empty] | 0.15312 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2095 | [5 empty] | 0.15312 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2095 | [5 empty] | 0.15312 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2095 | [5 empty] | 0.15312 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:08 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2095 | [5 empty] | 0.15312 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2094 | [5 empty] | 0.15312 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.5618 | [19 empty] | -1.2094 | [5 empty] | 0.15312 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012218 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2095 | [5 empty] | 0.15326 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2094 | [5 empty] | 0.15326 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.4459 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2094 | [5 empty] | 0.15326 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.4459 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2095 | [5 empty] | 0.15326 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2094 | [5 empty] | 0.15326 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2094 | [5 empty] | 0.15326 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.4459 | -1.4266 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2094 | [5 empty] | 0.15326 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.4459 | -1.4266 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:09 | 20 | 20 |  | 76.6308 | [19 empty] | -1.2095 | [5 empty] | 0.15326 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.0001223 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.6998 | [19 empty] | -1.2095 | [5 empty] | 0.1534 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012243 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.6998 | [19 empty] | -1.2095 | [5 empty] | 0.1534 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012243 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.6998 | [19 empty] | -1.2095 | [5 empty] | 0.1534 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012243 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.6998 | [19 empty] | -1.2095 | [5 empty] | 0.1534 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012243 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.6998 | [19 empty] | -1.2095 | [5 empty] | 0.1534 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012243 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2094 | [5 empty] | 0.15354 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2095 | [5 empty] | 0.15354 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2095 | [5 empty] | 0.15354 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2095 | [5 empty] | 0.15354 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2095 | [5 empty] | 0.15354 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:10 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2094 | [5 empty] | 0.15354 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.4459 | -1.4266 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.7687 | [19 empty] | -1.2095 | [5 empty] | 0.15354 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012256 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2094 | [5 empty] | 0.15368 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2094 | [5 empty] | 0.15368 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.4459 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:11 | 20 | 20 |  | 76.8377 | [19 empty] | -1.2095 | [5 empty] | 0.15368 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012268 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2095 | [5 empty] | 0.15381 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2094 | [5 empty] | 0.15381 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.4459 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2095 | [5 empty] | 0.15381 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2094 | [5 empty] | 0.15381 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2095 | [5 empty] | 0.15381 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2095 | [5 empty] | 0.15381 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9067 | [19 empty] | -1.2095 | [5 empty] | 0.15381 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012281 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9757 | [19 empty] | -1.2095 | [5 empty] | 0.15395 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012294 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9757 | [19 empty] | -1.2095 | [5 empty] | 0.15395 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012294 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:12 | 20 | 20 |  | 76.9757 | [19 empty] | -1.2095 | [5 empty] | 0.15395 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012294 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2095 | [5 empty] | 0.15409 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2095 | [5 empty] | 0.15409 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2095 | [5 empty] | 0.15409 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.446 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2096 | [5 empty] | 0.15409 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2096 | [5 empty] | 0.15409 | -1.2085 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.4461 | -1.4268 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2095 | [5 empty] | 0.15409 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2095 | [5 empty] | 0.15409 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.0446 | [19 empty] | -1.2095 | [5 empty] | 0.15409 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012306 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2094 | [5 empty] | 0.15423 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2095 | [5 empty] | 0.15423 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:13 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2095 | [5 empty] | 0.15423 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2095 | [5 empty] | 0.15423 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2094 | [5 empty] | 0.15423 | -1.2083 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.4459 | -1.4266 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2095 | [5 empty] | 0.15423 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 77.1136 | [19 empty] | -1.2095 | [5 empty] | 0.15423 | -1.2084 | 0.90444 | 17.6664 | 0.23773 | -0.00012319 | -1.446 | -1.4267 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 77.1826 | [19 empty] | -1.2088 | [5 empty] | 0.15437 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -0.00012332 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 35.246 | [19 empty] | -1.2087 | [5 empty] | 0.070492 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.634e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 31.9352 | [19 empty] | -1.2088 | [5 empty] | 0.06387 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.0263e-05 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 34.2114 | [19 empty] | -1.2088 | [5 empty] | 0.068423 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4441e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 32.487 | [19 empty] | -1.2088 | [5 empty] | 0.064974 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.1276e-05 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:14 | 20 | 20 |  | 32.9699 | [19 empty] | -1.2088 | [5 empty] | 0.06594 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.2162e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 33.2458 | [19 empty] | -1.2088 | [5 empty] | 0.066492 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.2669e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 33.5217 | [19 empty] | -1.2089 | [5 empty] | 0.067043 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.3175e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 33.7286 | [19 empty] | -1.2088 | [5 empty] | 0.067457 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.3555e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 33.8665 | [19 empty] | -1.2089 | [5 empty] | 0.067733 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.3808e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 34.0045 | [19 empty] | -1.2089 | [5 empty] | 0.068009 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.4061e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 34.1424 | [19 empty] | -1.2089 | [5 empty] | 0.068285 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.4314e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 34.2804 | [19 empty] | -1.2088 | [5 empty] | 0.068561 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4568e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 34.3494 | [19 empty] | -1.2089 | [5 empty] | 0.068699 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.4694e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 34.4873 | [19 empty] | -1.2088 | [5 empty] | 0.068975 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4947e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:15 | 20 | 20 |  | 34.5563 | [19 empty] | -1.2088 | [5 empty] | 0.069113 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5074e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2089 | [5 empty] | 0.06925 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 34.7632 | [19 empty] | -1.2089 | [5 empty] | 0.069526 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.5454e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 34.9011 | [19 empty] | -1.2089 | [5 empty] | 0.069802 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.5707e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 35.0391 | [19 empty] | -1.2089 | [5 empty] | 0.070078 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.596e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 35.1081 | [19 empty] | -1.2089 | [5 empty] | 0.070216 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.6087e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 35.177 | [19 empty] | -1.2088 | [5 empty] | 0.070354 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6213e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 35.246 | [19 empty] | -1.2088 | [5 empty] | 0.070492 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.634e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 35.315 | [19 empty] | -1.2089 | [5 empty] | 0.07063 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.6467e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:16 | 20 | 20 |  | 35.315 | [19 empty] | -1.2088 | [5 empty] | 0.07063 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6467e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.384 | [19 empty] | -1.2089 | [5 empty] | 0.070768 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.4529 | [19 empty] | -1.2088 | [5 empty] | 0.070906 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.672e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.5219 | [19 empty] | -1.2088 | [5 empty] | 0.071044 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6847e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2089 | [5 empty] | 0.071182 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2089 | [5 empty] | 0.071182 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2089 | [5 empty] | 0.071182 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2089 | [5 empty] | 0.07132 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.7288 | [19 empty] | -1.2089 | [5 empty] | 0.071458 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.7226e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2088 | [5 empty] | 0.071596 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:17 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2089 | [5 empty] | 0.071596 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2089 | [5 empty] | 0.071734 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2089 | [5 empty] | 0.071734 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2089 | [5 empty] | 0.071872 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2088 | [5 empty] | 0.072009 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2089 | [5 empty] | 0.072009 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2089 | [5 empty] | 0.072147 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2089 | [5 empty] | 0.072285 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2089 | [5 empty] | 0.072423 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2089 | [5 empty] | 0.072561 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:18 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2089 | [5 empty] | 0.072699 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2089 | [5 empty] | 0.072699 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2089 | [5 empty] | 0.072837 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2089 | [5 empty] | 0.072837 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2088 | [5 empty] | 0.072975 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2089 | [5 empty] | 0.073113 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2089 | [5 empty] | 0.073113 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.6255 | [19 empty] | -1.2089 | [5 empty] | 0.073251 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.6255 | [19 empty] | -1.2089 | [5 empty] | 0.073251 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8872e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.6945 | [19 empty] | -1.2089 | [5 empty] | 0.073389 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:19 | 20 | 20 |  | 36.6945 | [19 empty] | -1.2089 | [5 empty] | 0.073389 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.8999e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.7635 | [19 empty] | -1.2089 | [5 empty] | 0.073527 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.7635 | [19 empty] | -1.2088 | [5 empty] | 0.073527 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.9125e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.8324 | [19 empty] | -1.2089 | [5 empty] | 0.073665 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.8324 | [19 empty] | -1.2089 | [5 empty] | 0.073665 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9252e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.9014 | [19 empty] | -1.2089 | [5 empty] | 0.073803 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9379e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.9704 | [19 empty] | -1.2089 | [5 empty] | 0.073941 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 36.9704 | [19 empty] | -1.2089 | [5 empty] | 0.073941 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9505e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 37.0394 | [19 empty] | -1.2089 | [5 empty] | 0.074079 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 37.0394 | [19 empty] | -1.2089 | [5 empty] | 0.074079 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:20 | 20 | 20 |  | 37.0394 | [19 empty] | -1.2089 | [5 empty] | 0.074079 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9632e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.1083 | [19 empty] | -1.2089 | [5 empty] | 0.074217 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9758e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.1773 | [19 empty] | -1.2088 | [5 empty] | 0.074355 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.1773 | [19 empty] | -1.2089 | [5 empty] | 0.074355 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -4.9885e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.2463 | [19 empty] | -1.2089 | [5 empty] | 0.074493 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.2463 | [19 empty] | -1.2089 | [5 empty] | 0.074493 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0012e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.3153 | [19 empty] | -1.2089 | [5 empty] | 0.074631 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.3153 | [19 empty] | -1.2089 | [5 empty] | 0.074631 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.3153 | [19 empty] | -1.2088 | [5 empty] | 0.074631 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.0138e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.3842 | [19 empty] | -1.2089 | [5 empty] | 0.074768 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0265e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:21 | 20 | 20 |  | 37.3842 | [19 empty] | -1.2089 | [5 empty] | 0.074768 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0265e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.4532 | [19 empty] | -1.2089 | [5 empty] | 0.074906 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0391e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.4532 | [19 empty] | -1.2089 | [5 empty] | 0.074906 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0391e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.5222 | [19 empty] | -1.2089 | [5 empty] | 0.075044 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0518e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.5222 | [19 empty] | -1.2088 | [5 empty] | 0.075044 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.0518e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.5912 | [19 empty] | -1.2089 | [5 empty] | 0.075182 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0645e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.5222 | [19 empty] | -1.2089 | [5 empty] | 0.075044 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0518e-05 | -1.4454 | -1.4262 | -1.4262 | -1.4262 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.5912 | [19 empty] | -1.2089 | [5 empty] | 0.075182 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0645e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.6601 | [19 empty] | -1.2088 | [5 empty] | 0.07532 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.0771e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.6601 | [19 empty] | -1.2089 | [5 empty] | 0.07532 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0771e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.6601 | [19 empty] | -1.2088 | [5 empty] | 0.07532 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.0771e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:22 | 20 | 20 |  | 37.7291 | [19 empty] | -1.2089 | [5 empty] | 0.075458 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0898e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.7291 | [19 empty] | -1.2089 | [5 empty] | 0.075458 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0898e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.7291 | [19 empty] | -1.2089 | [5 empty] | 0.075458 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0898e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.7291 | [19 empty] | -1.2089 | [5 empty] | 0.075458 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.0898e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.7981 | [19 empty] | -1.2089 | [5 empty] | 0.075596 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1024e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.8671 | [19 empty] | -1.2089 | [5 empty] | 0.075734 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1151e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.8671 | [19 empty] | -1.2089 | [5 empty] | 0.075734 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1151e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.8671 | [19 empty] | -1.2089 | [5 empty] | 0.075734 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1151e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.936 | [19 empty] | -1.2088 | [5 empty] | 0.075872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1278e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.936 | [19 empty] | -1.2088 | [5 empty] | 0.075872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1278e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:23 | 20 | 20 |  | 37.936 | [19 empty] | -1.2088 | [5 empty] | 0.075872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1278e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.005 | [19 empty] | -1.2089 | [5 empty] | 0.07601 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1404e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.005 | [19 empty] | -1.2089 | [5 empty] | 0.07601 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1404e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.005 | [19 empty] | -1.2089 | [5 empty] | 0.07601 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1404e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.005 | [19 empty] | -1.2088 | [5 empty] | 0.07601 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1404e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.074 | [19 empty] | -1.2088 | [5 empty] | 0.076148 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1531e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.074 | [19 empty] | -1.2089 | [5 empty] | 0.076148 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1531e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.074 | [19 empty] | -1.2089 | [5 empty] | 0.076148 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1531e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.2119 | [19 empty] | -1.2088 | [5 empty] | 0.076424 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1784e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.2119 | [19 empty] | -1.2088 | [5 empty] | 0.076424 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1784e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.2119 | [19 empty] | -1.2089 | [5 empty] | 0.076424 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.1784e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:24 | 20 | 20 |  | 38.2119 | [19 empty] | -1.2088 | [5 empty] | 0.076424 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1784e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.2809 | [19 empty] | -1.2089 | [5 empty] | 0.076562 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1911e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.2809 | [19 empty] | -1.2088 | [5 empty] | 0.076562 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1911e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.2809 | [19 empty] | -1.2088 | [5 empty] | 0.076562 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1911e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.2809 | [19 empty] | -1.2088 | [5 empty] | 0.076562 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.1911e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.3499 | [19 empty] | -1.2089 | [5 empty] | 0.0767 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2037e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.3499 | [19 empty] | -1.2089 | [5 empty] | 0.0767 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2037e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.4189 | [19 empty] | -1.2088 | [5 empty] | 0.076838 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2164e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.4189 | [19 empty] | -1.2088 | [5 empty] | 0.076838 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2164e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.4878 | [19 empty] | -1.2088 | [5 empty] | 0.076976 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2291e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:25 | 20 | 20 |  | 38.4878 | [19 empty] | -1.2088 | [5 empty] | 0.076976 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2291e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.4878 | [19 empty] | -1.2089 | [5 empty] | 0.076976 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2291e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.4878 | [19 empty] | -1.2088 | [5 empty] | 0.076976 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2291e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.5568 | [19 empty] | -1.2089 | [5 empty] | 0.077114 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2417e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.5568 | [19 empty] | -1.2089 | [5 empty] | 0.077114 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2417e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.5568 | [19 empty] | -1.2089 | [5 empty] | 0.077114 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2417e-05 | -1.4454 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.6258 | [19 empty] | -1.2089 | [5 empty] | 0.077252 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2544e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.6258 | [19 empty] | -1.2089 | [5 empty] | 0.077252 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2544e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.6258 | [19 empty] | -1.2088 | [5 empty] | 0.077252 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2544e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.6948 | [19 empty] | -1.2089 | [5 empty] | 0.07739 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.267e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:26 | 20 | 20 |  | 38.6948 | [19 empty] | -1.2089 | [5 empty] | 0.07739 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.267e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 38.6948 | [19 empty] | -1.2089 | [5 empty] | 0.07739 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.267e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 38.7637 | [19 empty] | -1.2089 | [5 empty] | 0.077527 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2797e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 38.7637 | [19 empty] | -1.2088 | [5 empty] | 0.077527 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2797e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 38.8327 | [19 empty] | -1.2089 | [5 empty] | 0.077665 | -1.2078 | 0.90444 | 17.6664 | 0.23773 | -5.2924e-05 | -1.4455 | -1.4262 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 38.8327 | [19 empty] | -1.2088 | [5 empty] | 0.077665 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -5.2924e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 34.1424 | [19 empty] | -1.2088 | [5 empty] | 0.068285 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4314e-05 | -1.4454 | -1.4261 | -1.4261 | -1.4261 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 34.2114 | [19 empty] | -1.2088 | [5 empty] | 0.068423 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4441e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 34.2804 | [19 empty] | -1.2088 | [5 empty] | 0.068561 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4568e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 34.2804 | [19 empty] | -1.2088 | [5 empty] | 0.068561 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4568e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 34.2804 | [19 empty] | -1.2088 | [5 empty] | 0.068561 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4568e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:27 | 20 | 20 |  | 34.2804 | [19 empty] | -1.2088 | [5 empty] | 0.068561 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4568e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.2804 | [19 empty] | -1.2088 | [5 empty] | 0.068561 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4568e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.3494 | [19 empty] | -1.2088 | [5 empty] | 0.068699 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4694e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.3494 | [19 empty] | -1.2088 | [5 empty] | 0.068699 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4694e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.3494 | [19 empty] | -1.2088 | [5 empty] | 0.068699 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4694e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.3494 | [19 empty] | -1.2088 | [5 empty] | 0.068699 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4694e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.3494 | [19 empty] | -1.2088 | [5 empty] | 0.068699 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4694e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:28 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2087 | [5 empty] | 0.068837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2087 | [5 empty] | 0.068837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4183 | [19 empty] | -1.2088 | [5 empty] | 0.068837 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4821e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4873 | [19 empty] | -1.2088 | [5 empty] | 0.068975 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4947e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4873 | [19 empty] | -1.2088 | [5 empty] | 0.068975 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4947e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.4873 | [19 empty] | -1.2088 | [5 empty] | 0.068975 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.4947e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:29 | 20 | 20 |  | 34.5563 | [19 empty] | -1.2088 | [5 empty] | 0.069113 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5074e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.5563 | [19 empty] | -1.2088 | [5 empty] | 0.069113 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5074e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.5563 | [19 empty] | -1.2088 | [5 empty] | 0.069113 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5074e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2088 | [5 empty] | 0.06925 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2088 | [5 empty] | 0.06925 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2088 | [5 empty] | 0.06925 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2088 | [5 empty] | 0.06925 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2088 | [5 empty] | 0.06925 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6252 | [19 empty] | -1.2088 | [5 empty] | 0.06925 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5201e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:30 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.6942 | [19 empty] | -1.2088 | [5 empty] | 0.069388 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5327e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.7632 | [19 empty] | -1.2088 | [5 empty] | 0.069526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5454e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.7632 | [19 empty] | -1.2088 | [5 empty] | 0.069526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5454e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.7632 | [19 empty] | -1.2088 | [5 empty] | 0.069526 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5454e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.7632 | [19 empty] | -1.2087 | [5 empty] | 0.069526 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5454e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.8322 | [19 empty] | -1.2087 | [5 empty] | 0.069664 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.558e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.8322 | [19 empty] | -1.2088 | [5 empty] | 0.069664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.558e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:31 | 20 | 20 |  | 34.8322 | [19 empty] | -1.2088 | [5 empty] | 0.069664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.558e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.8322 | [19 empty] | -1.2087 | [5 empty] | 0.069664 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.558e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.8322 | [19 empty] | -1.2088 | [5 empty] | 0.069664 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.558e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.8322 | [19 empty] | -1.2088 | [5 empty] | 0.069664 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.558e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9011 | [19 empty] | -1.2088 | [5 empty] | 0.069802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5707e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9011 | [19 empty] | -1.2088 | [5 empty] | 0.069802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5707e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9011 | [19 empty] | -1.2088 | [5 empty] | 0.069802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5707e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9011 | [19 empty] | -1.2088 | [5 empty] | 0.069802 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5707e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9011 | [19 empty] | -1.2088 | [5 empty] | 0.069802 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5707e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:32 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2087 | [5 empty] | 0.06994 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2087 | [5 empty] | 0.06994 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2087 | [5 empty] | 0.06994 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 34.9701 | [19 empty] | -1.2088 | [5 empty] | 0.06994 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.5834e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 35.0391 | [19 empty] | -1.2088 | [5 empty] | 0.070078 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.596e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 35.0391 | [19 empty] | -1.2088 | [5 empty] | 0.070078 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.596e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 35.0391 | [19 empty] | -1.2088 | [5 empty] | 0.070078 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.596e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:33 | 20 | 20 |  | 35.1081 | [19 empty] | -1.2088 | [5 empty] | 0.070216 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6087e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.1081 | [19 empty] | -1.2088 | [5 empty] | 0.070216 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6087e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.1081 | [19 empty] | -1.2087 | [5 empty] | 0.070216 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6087e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.177 | [19 empty] | -1.2088 | [5 empty] | 0.070354 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6213e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.177 | [19 empty] | -1.2087 | [5 empty] | 0.070354 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6213e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.177 | [19 empty] | -1.2087 | [5 empty] | 0.070354 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6213e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.177 | [19 empty] | -1.2088 | [5 empty] | 0.070354 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6213e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.246 | [19 empty] | -1.2087 | [5 empty] | 0.070492 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.634e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.246 | [19 empty] | -1.2087 | [5 empty] | 0.070492 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.634e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.246 | [19 empty] | -1.2087 | [5 empty] | 0.070492 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.634e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:34 | 20 | 20 |  | 35.246 | [19 empty] | -1.2087 | [5 empty] | 0.070492 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.634e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.315 | [19 empty] | -1.2088 | [5 empty] | 0.07063 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6467e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.315 | [19 empty] | -1.2088 | [5 empty] | 0.07063 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6467e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.315 | [19 empty] | -1.2088 | [5 empty] | 0.07063 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6467e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.315 | [19 empty] | -1.2088 | [5 empty] | 0.07063 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6467e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.384 | [19 empty] | -1.2087 | [5 empty] | 0.070768 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.384 | [19 empty] | -1.2088 | [5 empty] | 0.070768 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.384 | [19 empty] | -1.2088 | [5 empty] | 0.070768 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.384 | [19 empty] | -1.2088 | [5 empty] | 0.070768 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.384 | [19 empty] | -1.2087 | [5 empty] | 0.070768 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:35 | 20 | 20 |  | 35.384 | [19 empty] | -1.2088 | [5 empty] | 0.070768 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.384 | [19 empty] | -1.2088 | [5 empty] | 0.070768 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.384 | [19 empty] | -1.2088 | [5 empty] | 0.070768 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6593e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.4529 | [19 empty] | -1.2088 | [5 empty] | 0.070906 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.672e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.4529 | [19 empty] | -1.2088 | [5 empty] | 0.070906 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.672e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.4529 | [19 empty] | -1.2088 | [5 empty] | 0.070906 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.672e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.4529 | [19 empty] | -1.2088 | [5 empty] | 0.070906 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.672e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.4529 | [19 empty] | -1.2088 | [5 empty] | 0.070906 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.672e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.5219 | [19 empty] | -1.2088 | [5 empty] | 0.071044 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.6847e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.5219 | [19 empty] | -1.2087 | [5 empty] | 0.071044 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6847e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.5219 | [19 empty] | -1.2087 | [5 empty] | 0.071044 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6847e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:36 | 20 | 20 |  | 35.5219 | [19 empty] | -1.2088 | [5 empty] | 0.071044 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6847e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2088 | [5 empty] | 0.071182 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2087 | [5 empty] | 0.071182 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2087 | [5 empty] | 0.071182 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.5909 | [19 empty] | -1.2088 | [5 empty] | 0.071182 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.6973e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2087 | [5 empty] | 0.07132 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2087 | [5 empty] | 0.07132 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2088 | [5 empty] | 0.07132 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2088 | [5 empty] | 0.07132 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2088 | [5 empty] | 0.07132 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:37 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2088 | [5 empty] | 0.07132 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.6599 | [19 empty] | -1.2087 | [5 empty] | 0.07132 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.71e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7288 | [19 empty] | -1.2087 | [5 empty] | 0.071458 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7226e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7288 | [19 empty] | -1.2087 | [5 empty] | 0.071458 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7226e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7288 | [19 empty] | -1.2087 | [5 empty] | 0.071458 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7226e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2087 | [5 empty] | 0.071596 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2088 | [5 empty] | 0.071596 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2088 | [5 empty] | 0.071596 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2088 | [5 empty] | 0.071596 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.7978 | [19 empty] | -1.2088 | [5 empty] | 0.071596 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7353e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2088 | [5 empty] | 0.071734 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:38 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2088 | [5 empty] | 0.071734 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2088 | [5 empty] | 0.071734 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2087 | [5 empty] | 0.071734 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.8668 | [19 empty] | -1.2088 | [5 empty] | 0.071734 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.748e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:39 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2088 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:40 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2088 | [5 empty] | 0.072009 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 35.9358 | [19 empty] | -1.2087 | [5 empty] | 0.071872 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7606e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2087 | [5 empty] | 0.072009 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2088 | [5 empty] | 0.072009 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2088 | [5 empty] | 0.072009 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2088 | [5 empty] | 0.072009 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0047 | [19 empty] | -1.2088 | [5 empty] | 0.072009 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7733e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2088 | [5 empty] | 0.072147 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2088 | [5 empty] | 0.072147 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2088 | [5 empty] | 0.072147 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:41 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2088 | [5 empty] | 0.072147 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4454 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2087 | [5 empty] | 0.072147 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.0737 | [19 empty] | -1.2087 | [5 empty] | 0.072147 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7859e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:42 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.1427 | [19 empty] | -1.2087 | [5 empty] | 0.072285 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.7986e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:43 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2077 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.4261 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:44 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2088 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2117 | [19 empty] | -1.2087 | [5 empty] | 0.072423 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8113e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2088 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2088 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:45 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2088 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:46 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2088 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:47 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE!
2020-10-27 15:47 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:47 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:47 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4452 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE!
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2088 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2088 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:48 | 20 | 20 |  | 36.2806 | [19 empty] | -1.2087 | [5 empty] | 0.072561 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8239e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:49 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2088 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2088 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2088 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:50 | 20 | 20 |  | 36.3496 | [19 empty] | -1.2087 | [5 empty] | 0.072699 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8366e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE!
2020-10-27 15:50 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2088 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2088 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:51 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE!
2020-10-27 15:52 | 20 | 20 |  | 36.4186 | [19 empty] | -1.2087 | [5 empty] | 0.072837 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8492e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:52 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2086 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:53 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:54 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:55 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2075 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4452 | -1.4259 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:56 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2087 | [5 empty] | 0.073113 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:57 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2087 | [5 empty] | 0.073113 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4452 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:57 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE!
2020-10-27 15:57 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2087 | [5 empty] | 0.073113 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:57 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | -1.426 | -1.426 | #VALUE! | #VALUE!
2020-10-27 15:57 | 20 | 20 |  | 36.4876 | [19 empty] | -1.2087 | [5 empty] | 0.072975 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8619e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:57 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2087 | [5 empty] | 0.073113 | -1.2076 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4453 | -1.426 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:57 | 20 | 20 |  | 36.5565 | [19 empty] | -1.2083 | [5 empty] | 0.073113 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | -4.8746e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:57 | 20 | 20 |  | 11.4498 | [19 empty] | -1.2081 | [5 empty] | 0.0229 | -1.207 | 0.90444 | 17.6664 | 0.23773 | -2.6611e-06 | -1.4447 | -1.4254 | -1.4254 | -1.4254 | #VALUE! | #VALUE!
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4257 | -1.4257 | -1.4257 | #VALUE! | #VALUE!
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2084 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.445 | -1.4257 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2084 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.445 | -1.4257 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4257 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4257 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:58 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 15:59 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:00 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:01 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:02 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2072 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:03 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:04 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2083 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:05 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:06 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4449 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:07 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4256 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:08 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:09 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:10 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:11 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:12 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:13 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:14 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:15 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:16 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:17 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:18 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:19 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.2071 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:20 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:21 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:22 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:23 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:24 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:25 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2082 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4448 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:26 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:27 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:28 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4255 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:29 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:30 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:31 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:32 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:33 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:34 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:35 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:36 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:37 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:38 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.207 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:39 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2068 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:40 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:41 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2068 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:42 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2068 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:43 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2068 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:44 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.2081 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4447 | -1.4254 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #VALUE! | #N/A
2020-10-27 16:45 | 20 | 20 |  | 0 | [19 empty] | -1.208 | [5 empty] | 0 | -1.2069 | 0.90444 | 17.6664 | 0.23773 | 1.8355e-05 | -1.4446 | -1.4253 | #N/A |  | #N/A | #N/A
Note on AI10: Johnson, Ken: Linear Fit to data in Harned and Owen Table 11-4-1A for 0.01 m HCl
Note on AJ10: Millero Chem Oc Vol 8, table 43.5 for pure water

## Dcalcs
Dcalcs
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6
Drift Corr (mV/d)
0
Calibration period
Start | 2020-10-26 12:50
Stop | 2020-10-27 16:45
Drift Offset Calc (mV)
T Offset | 0.58157 | days
k0 Offset | 0 | mV
Checksum |  | Zero?
Avg(Vrs) | -1.2184 | -0.00061218
Avg(Vrs+D) | -1.2178
Start | -1.2064 | -1.4249
Stop | -1.2069 | -1.4253

## k2calcs
k2calcs |  |  |  | Vrs-Vc | [12 empty] | Vrs-Vc-k2T | [12 empty] | Vrs-Vc-k2T (Up Only)
|  |  |  | W | [12 empty] | Y | [12 empty] | AA
Calibration period
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6 |  |  |  |  | Start | 2020-10-26 12:50 |  | -1.2064 | [5 empty] | min P | 0 | USE k2
[5 empty] | Stop | 2020-10-27 16:45 |  | -1.2069 | [5 empty] | Max P | 1600 | -0.00096427
#VALUE!
|  |  |  | 1 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19 |  | Del k0 | Del k2 | 1 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19 |  |  |  | 1 | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19
|  |  |  | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19 | 21 | 50 - Ave | #DIV/0! | #DIV/0! | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19 | 21 |  |  |  | 3 | 5 | 7 | 9 | 11 | 13 | 15 | 17 | 19 | 21
| 200 |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20
Pressure (dBar) |  |  |  | 1 3: Temp; 3 5: Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets.; 19 21: #DIV/0! |  | #VALUE! |  | Temp | Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. | [7 empty] | #DIV/0! |  | #VALUE! |  | Temp | Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. | [7 empty] | #DIV/0! | k0 HCl | #VALUE!
|  | Pres | Pres (ave) | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20 | Vrs Ave | k0 | k2 | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20 | Ave | k0 | k2-residual | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20 | Ave | k0 | k2-residual
5 | 10 | 7.5 | 5.5869 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4448 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4256 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
10 | 20 | 15 | 11.4498 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4447 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4254 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4254 | #DIV/0! | #DIV/0! | #DIV/0!
20 | 30 | 25 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
30 | 50 | 40 | 36.2749 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4451 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4258 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4255 | #DIV/0! | #DIV/0! | #DIV/0!
50 | 75 | 62.5 | 72.6572 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4457 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4264 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4261 | #DIV/0! | #DIV/0! | #DIV/0!
75 | 100 | 87.5 | 76.1168 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.446 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4267 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4261 | #DIV/0! | #DIV/0! | #DIV/0!
100 | 150 | 125 | 142.1567 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4469 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4277 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
150 | 200 | 175 | 173.7176 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
200 | 300 | 250 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
300 | 400 | 350 | 350.3999 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
400 | 500 | 450 | 445.438 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
500 | 600 | 550 | 520.5893 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4499 | -1.45 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0!
600 | 700 | 650 | 694.5857 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 8: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 10: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0!
700 | 800 | 750 | 702.6007 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 8: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 10: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0!
800 | 1000 | 900 | 894.9109 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 8: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 10: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0!
1000 | 1200 | 1100 | 1040.6705 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 8: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 10: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [5 empty] | #DIV/0! | #DIV/0! | #DIV/0!
1200 | 1400 | 1300 | 1383.4765 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0!
1400 | 1600 | 1500 | 1499.025 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0!
1600 | 1800 | 1700 | 1734.1921 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0!
1800 | 2000 | 1900 | 1833.069 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0!
2000 | 2200 | 2100 | 2069.8367 | Temp 2: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 4: #DIV/0!; Note: For #DIV/0! errors, enter red formula -> Be sure to copy down or over for new spreadsheets. 6: #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0!
All data | [15 empty] | All data=biased
0 | 2200 | 1100 | 1078.8316 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -1.4542 | #DIV/0! | #DIV/0! | #DIV/0!
|  | Ave |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | StDev |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## fPcalcs
f(P) calcs |  |  |  | MBARI DuraFET DF#313 + VL#313 PEEK Ver 6
Vrs - Vc - k2*T = k6 * p^6 + k5 * p^5 + k4 * p^4 + k3 * p^3 + k2 * p^2 + k1 * p + k0
UP profiles ONLY
| k6 | k5 | k4 | k3 | k2 | k1 | k0 | R^2
2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
4 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
6 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
8 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
10 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
12 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
14 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
16 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
18 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
20 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Ave | [6 empty] | #VALUE! | #VALUE!
StDev | [6 empty] | #VALUE! | #VALUE!
Ave | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
| #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
By interpolating Vrs signals to P=0: | [6 empty] | k0
[5 empty] | Vrs-Vc |  | #VALUE!
[5 empty] | Vrs-Vc-k2T |  | #VALUE!
[5 empty] | Vrs-Vc-k2T (Up Only) |  | #VALUE!
| #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
#VALUE!

## Poff
SW Calibration | [24 empty] | K0 spec pH |  |  |  | 7 | Adjusted Time
[22 empty] | Start: |  |  | Colorimetric Analysis | Date |  | SW | Time (d) | offsets: | 7
Pump OFF | [23 empty] | #N/A |  |  |  |  | 0.29167 | 0.29167
[22 empty] | 0 |  | #N/A |  |  |  |  | 0.29167 | 0.29167
[29 empty] | 0.29167 | 0.29167
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6 | [21 empty] | Stop: | [6 empty] | 0.29167 | 0.29167
[24 empty] | #N/A |  |  |  |  | 0.29167 | 0.29167
[22 empty] | 0 |  | #N/A |  |  |  |  | 0.29167 | 0.29167 | [6 empty] | Ave | StDev | Dk0 (off-on)
[29 empty] | 0.29167 | 0.29167 |  |  |  |  | In range | Pump ON | #DIV/0! | #DIV/0! | #DIV/0!
[29 empty] | 0.29167 | 0.29167 | [5 empty] | Pump OFF: | #DIV/0! | #DIV/0!
[29 empty] | 0.29167 | 0.29167
[29 empty] | 0.29167 | 0.29167 | [7 empty] | Flyer finder
[29 empty] | 0.29167 | 0.29167 |  |  |  |  | All data | Max | #DIV/0! | #DIV/0! | #DIV/0!
[29 empty] | 0.29167 | 0.29167 | [5 empty] | Min | #DIV/0! | #DIV/0! | #DIV/0!
[25 empty] | 2nd order |  | Avg: | #DIV/0! | [7 empty] | Max | #DIV/0! | Ave
[25 empty] | = t^2 * | 0 | StDev: | #DIV/0! | Fit | [6 empty] | Min | #DIV/0! | #DIV/0!
[25 empty] | + t * | 0 | Slope: | #DIV/0! | [0,1,2]
[25 empty] | + int | 0 | Intercept: | #DIV/0! | 1 | [5 empty] | PTS: | 1
|  | Time | 1970 seconds | Pres | Temp | S | Sam ctr | Pwr ctr | Err ctr | Int Temp | Int Hum | Supp (V) | Current | fake pH | Bias V | Vrs | SDev Vrs | Vk | SDev Vk | Ik | Ib | T (K) | 1/TK | S | Ionic Strength | mCl | SO4 (tot) | pHTTSW tot | K(HSO4) | pHtris free | Debye Huckel Const | log activity coef | Beta SO4 | mHfree | aHfree x aCl | k0+k2T | k0 | Dk0 (off-on) | Time offset | 0.0003 | -1.3925 |  |  |  | LAB SPEC pH data during K0 run
0x2E00 | pH | 2019-03-11 23:19 | 1552346310 | -0.08 | 16.9425 | 33.1357 | 0 | 2 | 0 | 20.78 | 48.58 | 11.924 | 0.046 | 7.1234 | 3.659 | -0.9409 | 2.4e-05 | -0.97044 | 2.8e-05 | 7.838e-09 | 1.362e-09 | 290.0925 | 0.0034472 | 33.1357 | 0.68294 | 0.53523 | 0.026731 | #DIV/0! | 0.13096 | #DIV/0! | 0.50415 | -0.1342 | 7.6358 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9713 |  |  | K0 data | date | pH | date+time | final time | pH | std dev
0xFAA5 | pH | 2019-03-11 23:22 | 1552346491 | -0.08 | 16.5555 | 33.164 | 2 | 2 | 0 | 20.86 | 48.16 | 11.928 | 0.0462 | 7.1234 | 3.66 | -0.94055 | 1.7e-05 | -0.96972 | 1.9e-05 | 7.772e-09 | 8.853e-10 | 289.7055 | 0.0034518 | 33.164 | 0.68354 | 0.5357 | 0.026754 | #DIV/0! | 0.13303 | #DIV/0! | 0.50384 | -0.13405 | 7.5172 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9734
0xA74B | pH | 2019-03-11 23:25 | 1552346671 | -0.07 | 16.6014 | 33.1455 | 4 | 2 | 0 | 20.93 | 47.63 | 11.928 | 0.0462 | 7.1234 | 3.66 | -0.94033 | 2.7e-05 | -0.9699 | 3.2e-05 | 7.239e-09 | 1.163e-09 | 289.7514 | 0.0034512 | 33.1455 | 0.68315 | 0.53539 | 0.026739 | #DIV/0! | 0.13274 | #DIV/0! | 0.50388 | -0.13407 | 7.5333 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9755
0x74BC | pH | 2019-03-11 23:28 | 1552346851 | -0.08 | 16.6337 | 33.1385 | 6 | 2 | 0 | 20.96 | 47.42 | 11.928 | 0.0462 | 7.1234 | 3.66 | -0.94004 | 4.5e-05 | -0.97011 | 2.4e-05 | 7.479e-09 | 1.678e-09 | 289.7837 | 0.0034508 | 33.1385 | 0.683 | 0.53527 | 0.026734 | #DIV/0! | 0.13256 | #DIV/0! | 0.5039 | -0.13409 | 7.5439 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9775
0xDFFF | pH | 2019-03-11 23:31 | 1552347031 | -0.08 | 16.6583 | 33.1368 | 8 | 2 | 0 | 21 | 47.64 | 11.928 | 0.0462 | 7.1234 | 3.66 | -0.93987 | 3.4e-05 | -0.97046 | 3.5e-05 | 7.947e-09 | 2.326e-09 | 289.8083 | 0.0034506 | 33.1368 | 0.68296 | 0.53525 | 0.026732 | #DIV/0! | 0.13243 | #DIV/0! | 0.50392 | -0.1341 | 7.5514 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9796
0xBB36 | pH | 2019-03-11 23:34 | 1552347211 | -0.08 | 16.7119 | 33.1351 | 10 | 2 | 0 | 21.02 | 47.27 | 11.928 | 0.0462 | 7.1234 | 3.661 | -0.93958 | 3.1e-05 | -0.97078 | 2.3e-05 | 7.648e-09 | 1.251e-09 | 289.8619 | 0.0034499 | 33.1351 | 0.68293 | 0.53522 | 0.026731 | #DIV/0! | 0.13214 | #DIV/0! | 0.50396 | -0.13412 | 7.5675 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9817
0x690B | pH | 2019-03-11 23:37 | 1552347391 | -0.07 | 16.7524 | 33.1277 | 12 | 2 | 0 | 21.03 | 47.38 | 11.928 | 0.046 | 7.1234 | 3.661 | -0.93946 | 2.1e-05 | -0.97117 | 2.4e-05 | 7.242e-09 | 1.164e-09 | 289.9024 | 0.0034494 | 33.1277 | 0.68277 | 0.53509 | 0.026725 | #DIV/0! | 0.13192 | #DIV/0! | 0.504 | -0.13414 | 7.5805 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9838
0xF630 | pH | 2019-03-11 23:40 | 1552347571 | -0.08 | 16.804 | 33.1187 | 14 | 2 | 0 | 21.03 | 47.32 | 11.928 | 0.046 | 7.1234 | 3.661 | -0.93928 | 2.9e-05 | -0.9715 | 2.5e-05 | 7.85e-09 | 2.573e-09 | 289.954 | 0.0034488 | 33.1187 | 0.68258 | 0.53494 | 0.026718 | #DIV/0! | 0.13163 | #DIV/0! | 0.50404 | -0.13416 | 7.5971 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9859
0x2019 | pH | 2019-03-11 23:43 | 1552347751 | -0.08 | 16.8417 | 33.1231 | 16 | 2 | 0 | 21.03 | 47.1 | 11.928 | 0.0462 | 7.1234 | 3.661 | -0.93914 | 3.5e-05 | -0.97187 | 3.5e-05 | 7.706e-09 | 1.357e-09 | 289.9917 | 0.0034484 | 33.1231 | 0.68267 | 0.53502 | 0.026721 | #DIV/0! | 0.13145 | #DIV/0! | 0.50407 | -0.13417 | 7.6076 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.988
0x1298 | pH | 2019-03-11 23:46 | 1552347931 | -0.06 | 16.8738 | 33.1263 | 18 | 2 | 0 | 21.03 | 47.34 | 11.928 | 0.0462 | 7.1234 | 3.661 | -0.93911 | 2.5e-05 | -0.97228 | 2.3e-05 | 7.914e-09 | 1.464e-09 | 290.0238 | 0.003448 | 33.1263 | 0.68274 | 0.53507 | 0.026724 | #DIV/0! | 0.13129 | #DIV/0! | 0.50409 | -0.13418 | 7.6167 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9901
0x2F41 | pH | 2019-03-11 23:49 | 1552348110 | -0.07 | 16.8954 | 33.1216 | 20 | 2 | 0 | 21.04 | 47.62 | 11.932 | 0.0462 | 7.1234 | 3.661 | -0.9389 | 1.4e-05 | -0.97258 | 3.5e-05 | 7.335e-09 | 7.161e-10 | 290.0454 | 0.0034477 | 33.1216 | 0.68264 | 0.53499 | 0.02672 | #DIV/0! | 0.13117 | #DIV/0! | 0.50411 | -0.13419 | 7.6238 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9921
0x2921 | pH | 2019-03-11 23:52 | 1552348291 | -0.08 | 16.9546 | 33.123 | 22 | 2 | 0 | 21.06 | 47.41 | 11.928 | 0.046 | 7.1234 | 3.661 | -0.93879 | 1.9e-05 | -0.97289 | 1.9e-05 | 8.203e-09 | 3.057e-10 | 290.1046 | 0.003447 | 33.123 | 0.68267 | 0.53502 | 0.026721 | #DIV/0! | 0.13087 | #DIV/0! | 0.50416 | -0.13422 | 7.6412 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9942
0xFB7E | pH | 2019-03-11 23:55 | 1552348471 | -0.08 | 16.9946 | 33.1248 | 24 | 2 | 0 | 21.07 | 47.51 | 11.932 | 0.046 | 7.1234 | 3.661 | -0.93867 | 2e-05 | -0.97319 | 4.9e-05 | 7.366e-09 | 1.113e-09 | 290.1446 | 0.0034466 | 33.1248 | 0.68271 | 0.53505 | 0.026723 | #DIV/0! | 0.13067 | #DIV/0! | 0.50419 | -0.13423 | 7.6529 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9963
0x1532 | pH | 2019-03-11 23:58 | 1552348651 | -0.08 | 17.0538 | 33.1179 | 26 | 2 | 0 | 21.07 | 47.44 | 11.932 | 0.0462 | 7.1234 | 3.662 | -0.93864 | 3.3e-05 | -0.9736 | 2.9e-05 | 7.487e-09 | 6.219e-10 | 290.2038 | 0.0034459 | 33.1179 | 0.68256 | 0.53493 | 0.026717 | #DIV/0! | 0.13035 | #DIV/0! | 0.50423 | -0.13426 | 7.6715 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43535.9984
0x7D26 | pH | 2019-03-12 00:01 | 1552348831 | -0.08 | 17.0533 | 33.124 | 28 | 2 | 0 | 21.1 | 47.63 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93847 | 1.3e-05 | -0.97385 | 1.6e-05 | 7.463e-09 | 1.131e-09 | 290.2033 | 0.0034459 | 33.124 | 0.68269 | 0.53503 | 0.026722 | #DIV/0! | 0.13037 | #DIV/0! | 0.50423 | -0.13425 | 7.6705 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0005
0x46D3 | pH | 2019-03-12 00:04 | 1552349011 | -0.08 | 17.1105 | 33.1259 | 30 | 2 | 0 | 21.11 | 47.66 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93842 | 2.5e-05 | -0.97417 | 2.9e-05 | 7.6e-09 | 1.049e-09 | 290.2605 | 0.0034452 | 33.1259 | 0.68273 | 0.53506 | 0.026724 | #DIV/0! | 0.13008 | #DIV/0! | 0.50428 | -0.13427 | 7.6873 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0026
0x4009 | pH | 2019-03-12 00:07 | 1552349191 | -0.08 | 17.1659 | 33.1197 | 32 | 2 | 0 | 21.14 | 47.44 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93844 | 2.1e-05 | -0.97458 | 2.9e-05 | 7.812e-09 | 9.472e-10 | 290.3159 | 0.0034445 | 33.1197 | 0.6826 | 0.53496 | 0.026719 | #DIV/0! | 0.12979 | #DIV/0! | 0.50432 | -0.1343 | 7.7048 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0046
0xE663 | pH | 2019-03-12 00:10 | 1552349371 | -0.08 | 17.1781 | 33.1244 | 34 | 2 | 0 | 21.14 | 47.46 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93833 | 2.6e-05 | -0.97483 | 3.9e-05 | 7.429e-09 | 1.586e-09 | 290.3281 | 0.0034444 | 33.1244 | 0.6827 | 0.53504 | 0.026722 | #DIV/0! | 0.12974 | #DIV/0! | 0.50433 | -0.1343 | 7.7078 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0067
0xF81C | pH | 2019-03-12 00:13 | 1552349551 | -0.08 | 17.2084 | 33.1198 | 36 | 2 | 0 | 21.13 | 47.49 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93826 | 2.3e-05 | -0.97508 | 2.6e-05 | 7.574e-09 | 1.49e-09 | 290.3584 | 0.003444 | 33.1198 | 0.6826 | 0.53496 | 0.026719 | #DIV/0! | 0.12958 | #DIV/0! | 0.50436 | -0.13431 | 7.7175 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0088
0x8BA9 | pH | 2019-03-12 00:16 | 1552349731 | -0.08 | 17.2396 | 33.1247 | 38 | 2 | 0 | 21.13 | 47.51 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93822 | 2.1e-05 | -0.97544 | 2.3e-05 | 7.981e-09 | 1.169e-09 | 290.3896 | 0.0034436 | 33.1247 | 0.6827 | 0.53504 | 0.026723 | #DIV/0! | 0.12943 | #DIV/0! | 0.50438 | -0.13432 | 7.7262 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0109
0x598F | pH | 2019-03-12 00:19 | 1552349911 | -0.08 | 17.2598 | 33.1221 | 40 | 2 | 0 | 21.11 | 47.7 | 11.932 | 0.0458 | 7.1234 | 3.662 | -0.93813 | 1.7e-05 | -0.97568 | 2.5e-05 | 7.669e-09 | 1.203e-09 | 290.4098 | 0.0034434 | 33.1221 | 0.68265 | 0.535 | 0.02672 | #DIV/0! | 0.12932 | #DIV/0! | 0.5044 | -0.13433 | 7.7326 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.013
0x53D3 | pH | 2019-03-12 00:22 | 1552350091 | -0.08 | 17.3047 | 33.1261 | 42 | 2 | 0 | 21.11 | 47.66 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93815 | 1.9e-05 | -0.97607 | 1.8e-05 | 7.677e-09 | 2.175e-09 | 290.4547 | 0.0034429 | 33.1261 | 0.68273 | 0.53507 | 0.026724 | #DIV/0! | 0.12911 | #DIV/0! | 0.50443 | -0.13435 | 7.7455 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0151
0x7225 | pH | 2019-03-12 00:25 | 1552350271 | -0.09 | 17.3439 | 33.1231 | 44 | 2 | 0 | 21.1 | 47.68 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93797 | 2.1e-05 | -0.97626 | 2.9e-05 | 7.6e-09 | 8.076e-10 | 290.4939 | 0.0034424 | 33.1231 | 0.68267 | 0.53502 | 0.026721 | #DIV/0! | 0.1289 | #DIV/0! | 0.50446 | -0.13436 | 7.7577 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0171
0x8273 | pH | 2019-03-12 00:28 | 1552350450 | -0.09 | 17.355 | 33.1213 | 46 | 2 | 0 | 21.1 | 47.73 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93802 | 3e-05 | -0.97664 | 2.7e-05 | 7.734e-09 | 9.421e-10 | 290.505 | 0.0034423 | 33.1213 | 0.68263 | 0.53499 | 0.02672 | #DIV/0! | 0.12884 | #DIV/0! | 0.50447 | -0.13437 | 7.7613 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0192
0xB454 | pH | 2019-03-12 00:31 | 1552350630 | -0.09 | 17.4061 | 33.1273 | 48 | 2 | 0 | 21.11 | 47.68 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93793 | 3.2e-05 | -0.97688 | 2e-05 | 7.596e-09 | 1.505e-09 | 290.5561 | 0.0034417 | 33.1273 | 0.68276 | 0.53509 | 0.026725 | #DIV/0! | 0.1286 | #DIV/0! | 0.50451 | -0.13438 | 7.7759 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0213
0xA820 | pH | 2019-03-12 00:34 | 1552350811 | -0.09 | 17.4279 | 33.1254 | 50 | 2 | 0 | 21.11 | 47.81 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93789 | 1.9e-05 | -0.97719 | 3e-05 | 7.45e-09 | 1.494e-09 | 290.5779 | 0.0034414 | 33.1254 | 0.68272 | 0.53506 | 0.026723 | #DIV/0! | 0.12849 | #DIV/0! | 0.50453 | -0.13439 | 7.7827 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0234
0x887E | pH | 2019-03-12 00:37 | 1552350991 | -0.09 | 17.4676 | 33.1229 | 52 | 2 | 0 | 21.1 | 48.05 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.9379 | 3.6e-05 | -0.9775 | 1.9e-05 | 7.757e-09 | 5.706e-10 | 290.6176 | 0.0034409 | 33.1229 | 0.68267 | 0.53501 | 0.026721 | #DIV/0! | 0.12829 | #DIV/0! | 0.50456 | -0.13441 | 7.795 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0255
0x3FD5 | pH | 2019-03-12 00:40 | 1552351171 | -0.08 | 17.4897 | 33.1232 | 54 | 2 | 0 | 21.1 | 48.08 | 11.936 | 0.046 | 7.1234 | 3.662 | -0.93782 | 2.7e-05 | -0.97774 | 2.3e-05 | 7.8e-09 | 9.201e-10 | 290.6397 | 0.0034407 | 33.1232 | 0.68267 | 0.53502 | 0.026721 | #DIV/0! | 0.12818 | #DIV/0! | 0.50458 | -0.13442 | 7.8017 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0275
0x895B | pH | 2019-03-12 00:43 | 1552351351 | -0.09 | 17.52 | 33.124 | 56 | 2 | 0 | 21.1 | 48 | 11.932 | 0.0462 | 7.1234 | 3.662 | -0.93778 | 1.7e-05 | -0.97803 | 2.6e-05 | 7.806e-09 | 8.781e-10 | 290.67 | 0.0034403 | 33.124 | 0.68269 | 0.53503 | 0.026722 | #DIV/0! | 0.12803 | #DIV/0! | 0.5046 | -0.13443 | 7.8107 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0296
0xBF9A | pH | 2019-03-12 00:46 | 1552351531 | -0.09 | 17.55 | 33.1262 | 58 | 2 | 0 | 21.1 | 48.05 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93774 | 1.9e-05 | -0.97829 | 2.9e-05 | 7.501e-09 | 7.691e-10 | 290.7 | 0.00344 | 33.1262 | 0.68274 | 0.53507 | 0.026724 | #DIV/0! | 0.12789 | #DIV/0! | 0.50463 | -0.13444 | 7.8195 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0317
0x1DB2 | pH | 2019-03-12 00:49 | 1552351711 | -0.08 | 17.5961 | 33.1212 | 60 | 2 | 0 | 21.1 | 48.17 | 11.936 | 0.046 | 7.1234 | 3.662 | -0.93774 | 3.2e-05 | -0.97863 | 4e-05 | 8.167e-09 | 1.357e-09 | 290.7461 | 0.0034394 | 33.1212 | 0.68263 | 0.53499 | 0.02672 | #DIV/0! | 0.12765 | #DIV/0! | 0.50467 | -0.13446 | 7.8342 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0338
0xF06D | pH | 2019-03-12 00:52 | 1552351891 | -0.1 | 17.6169 | 33.125 | 62 | 2 | 0 | 21.1 | 48.25 | 11.932 | 0.046 | 7.1234 | 3.662 | -0.93769 | 1.9e-05 | -0.97888 | 2.3e-05 | 7.889e-09 | 1.243e-09 | 290.7669 | 0.0034392 | 33.125 | 0.68271 | 0.53505 | 0.026723 | #DIV/0! | 0.12755 | #DIV/0! | 0.50468 | -0.13447 | 7.8399 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0359
0x8900 | pH | 2019-03-12 00:55 | 1552352071 | -0.09 | 17.6563 | 33.1211 | 64 | 2 | 0 | 21.11 | 48.52 | 11.936 | 0.046 | 7.1234 | 3.662 | -0.93766 | 2.8e-05 | -0.97916 | 2.3e-05 | 7.476e-09 | 1.095e-09 | 290.8063 | 0.0034387 | 33.1211 | 0.68263 | 0.53498 | 0.02672 | #DIV/0! | 0.12735 | #DIV/0! | 0.50471 | -0.13448 | 7.8524 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.038
0xA12C | pH | 2019-03-12 00:58 | 1552352251 | -0.08 | 17.673 | 33.1234 | 66 | 2 | 0 | 21.11 | 48.33 | 11.936 | 0.0462 | 7.1234 | 3.662 | -0.93769 | 1.9e-05 | -0.97947 | 3.7e-05 | 7.887e-09 | 1.12e-09 | 290.823 | 0.0034385 | 33.1234 | 0.68268 | 0.53502 | 0.026722 | #DIV/0! | 0.12727 | #DIV/0! | 0.50473 | -0.13449 | 7.8572 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0401
0x62F5 | pH | 2019-03-12 01:01 | 1552352431 | -0.09 | 17.7024 | 33.1269 | 68 | 2 | 0 | 21.11 | 48.43 | 11.936 | 0.046 | 7.1234 | 3.662 | -0.93773 | 1.8e-05 | -0.97979 | 3.1e-05 | 8.006e-09 | 1.393e-09 | 290.8524 | 0.0034382 | 33.1269 | 0.68275 | 0.53508 | 0.026724 | #DIV/0! | 0.12714 | #DIV/0! | 0.50475 | -0.1345 | 7.8656 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0421
0x3814 | pH | 2019-03-12 01:04 | 1552352610 | -0.1 | 17.7259 | 33.1211 | 70 | 2 | 0 | 21.11 | 48.45 | 11.936 | 0.046 | 7.1234 | 3.663 | -0.93769 | 2.9e-05 | -0.98007 | 3.5e-05 | 7.743e-09 | 7.153e-10 | 290.8759 | 0.0034379 | 33.1211 | 0.68263 | 0.53498 | 0.02672 | #DIV/0! | 0.12701 | #DIV/0! | 0.50477 | -0.13451 | 7.8736 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0442
0xDEA1 | pH | 2019-03-12 01:07 | 1552352791 | -0.09 | 17.7583 | 33.1207 | 72 | 2 | 0 | 21.13 | 48.6 | 11.936 | 0.046 | 7.1234 | 3.663 | -0.93762 | 3.7e-05 | -0.98029 | 2.3e-05 | 8.126e-09 | 1.414e-09 | 290.9083 | 0.0034375 | 33.1207 | 0.68262 | 0.53498 | 0.026719 | #DIV/0! | 0.12685 | #DIV/0! | 0.50479 | -0.13452 | 7.8835 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0463
0x733A | pH | 2019-03-12 01:10 | 1552352971 | -0.09 | 17.792 | 33.1202 | 74 | 2 | 0 | 21.13 | 48.7 | 11.936 | 0.0462 | 7.1234 | 3.663 | -0.93753 | 2.7e-05 | -0.9805 | 2.6e-05 | 7.227e-09 | 1.295e-09 | 290.942 | 0.0034371 | 33.1202 | 0.68261 | 0.53497 | 0.026719 | #DIV/0! | 0.12668 | #DIV/0! | 0.50482 | -0.13454 | 7.8938 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 43536.0484
Formulas:
Adjusted Time = [Date]+AG2/24
Start: = [Start:, row above]
Note on Z19: Hans Jannasch: Use 35 for buffers; CTD underestimates Cl from S.
Note on AA19: Johnson, Ken: From Dickson, Sabine, Christian Best Practices, eqn. 34, Chap. 5. This is in mol/kg H20 and is the desired form.
Note on AC19: Johnson, Ken: From Dickson, Sabine, Christian, Best Practices, Table 2, Chap. 5. Units mol/kg soln - for use with KHSO4
Note on AD19: =pH(TT)-.015*(20-C)
Note on AE19: Johnson, Ken: From Dickson, Sabine, Christian, Best Practices, eqn 33, Chap. 5. Units mol/kg solution
Note on AF19: Johnson, Ken: mol/kg solution
Note on AG19: Johnson, Ken: Empirical fit to data in Table II of Khoo et al Anal Chem 1977
Note on AH19: Johnson, Ken: From Khoo et al. Anal Chem. 1977 For molal conc. Mol/kg H2O
Note on AJ19: Johnson, Ken: Changed -was mol/kg soln, now mol/kg H2O
Note on AL19: =Vrs+R*ln10/F*T*log(aHfreeaCl)

## Pon
SW Calibration
Pump ON
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6
|  | Time | 1970 seconds | P | T | S | Sam ctr | Pwr ctr | Err ctr | Int Temp | Int Hum | Supp (V) | Current | fake pH | Bias V | Vrs | SDev Vrs | Vk | SDev Vk | Ik | Ib | T (K) | 1/TK | S | Ionic Strength | mCl | SO4 (tot) | pHTTSW tot | K(HSO4) | pHtris free | Debye Huckel Const | log activity coef | Beta SO4 | mHfree | aHfree x aCl | k0+k2T | k0 | 0.0003 | -1.3925
0x7648 | pH | 2019-03-11 23:19 | 1552346310 | -0.08 | 16.9425 | 33.1357 | 1 | 2 | 0 | 20.79 | 48.86 | 11.924 | 0.046 | 7.1234 | 3.661 | -0.94066 | 2.7e-05 | -0.97003 | 3e-05 | 7.801e-09 | 7.828e-10 | 290.0925 | 0.0034472 | 33.1357 | 0.68294 | 0.53523 | 0.026731 | #DIV/0! | 0.13096 | #DIV/0! | 0.50415 | -0.1342 | 7.6358 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
0xA136 | pH | 2019-03-11 23:22 | 1552346491 | -0.08 | 16.5555 | 33.164 | 3 | 2 | 0 | 20.87 | 48.09 | 11.928 | 0.0456 | 7.1234 | 3.661 | -0.9405 | 2.2e-05 | -0.96957 | 2.8e-05 | 7.568e-09 | 9.3e-10 | 289.7055 | 0.0034518 | 33.164 | 0.68354 | 0.5357 | 0.026754 | #DIV/0! | 0.13303 | #DIV/0! | 0.50384 | -0.13405 | 7.5172 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #DIV/0!
0x8FBA | pH | 2019-03-11 23:25 | 1552346671 | -0.07 | 16.6014 | 33.1455 | 5 | 2 | 0 | 20.93 | 47.63 | 11.928 | 0.0458 | 7.1234 | 3.661 | -0.94029 | 3.8e-05 | -0.96985 | 3e-05 | 7.784e-09 | 8.148e-10 | 289.7514 | 0.0034512 | 33.1455 | 0.68315 | 0.53539 | 0.026739 | #DIV/0! | 0.13274 | #DIV/0! | 0.50388 | -0.13407 | 7.5333 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0x8ED2 | pH | 2019-03-11 23:28 | 1552346851 | -0.08 | 16.6337 | 33.1385 | 7 | 2 | 0 | 20.97 | 47.27 | 11.928 | 0.0458 | 7.1234 | 3.661 | -0.94007 | 2.2e-05 | -0.97015 | 3.2e-05 | 7.654e-09 | 9.942e-10 | 289.7837 | 0.0034508 | 33.1385 | 0.683 | 0.53527 | 0.026734 | #DIV/0! | 0.13256 | #DIV/0! | 0.5039 | -0.13409 | 7.5439 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0x54BD | pH | 2019-03-11 23:31 | 1552347031 | -0.08 | 16.6583 | 33.1368 | 9 | 2 | 0 | 20.99 | 47.69 | 11.928 | 0.0458 | 7.1234 | 3.661 | -0.93978 | 1.9e-05 | -0.97042 | 3.1e-05 | 7.966e-09 | 6.177e-10 | 289.8083 | 0.0034506 | 33.1368 | 0.68296 | 0.53525 | 0.026732 | #DIV/0! | 0.13243 | #DIV/0! | 0.50392 | -0.1341 | 7.5514 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0xDD0A | pH | 2019-03-11 23:34 | 1552347211 | -0.08 | 16.7119 | 33.1351 | 11 | 2 | 0 | 21.02 | 47.59 | 11.928 | 0.0456 | 7.1234 | 3.662 | -0.93961 | 4.5e-05 | -0.97079 | 4.2e-05 | 7.228e-09 | 1.248e-09 | 289.8619 | 0.0034499 | 33.1351 | 0.68293 | 0.53522 | 0.026731 | #DIV/0! | 0.13214 | #DIV/0! | 0.50396 | -0.13412 | 7.5675 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0xC762 | pH | 2019-03-11 23:37 | 1552347391 | -0.07 | 16.7524 | 33.1277 | 13 | 2 | 0 | 21.03 | 47.57 | 11.928 | 0.0458 | 7.1234 | 3.662 | -0.93947 | 4.4e-05 | -0.97118 | 3e-05 | 7.711e-09 | 1.442e-09 | 289.9024 | 0.0034494 | 33.1277 | 0.68277 | 0.53509 | 0.026725 | #DIV/0! | 0.13192 | #DIV/0! | 0.504 | -0.13414 | 7.5805 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0xA756 | pH | 2019-03-11 23:40 | 1552347571 | -0.08 | 16.804 | 33.1187 | 15 | 2 | 0 | 21.03 | 47.32 | 11.928 | 0.0458 | 7.1234 | 3.662 | -0.93935 | 4e-05 | -0.97156 | 2.4e-05 | 8.035e-09 | 2.34e-09 | 289.954 | 0.0034488 | 33.1187 | 0.68258 | 0.53494 | 0.026718 | #DIV/0! | 0.13163 | #DIV/0! | 0.50404 | -0.13416 | 7.5971 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0xF35B | pH | 2019-03-11 23:43 | 1552347751 | -0.08 | 16.8417 | 33.1231 | 17 | 2 | 0 | 21.03 | 47.12 | 11.928 | 0.0456 | 7.1234 | 3.662 | -0.93922 | 2.5e-05 | -0.97192 | 2.1e-05 | 7.258e-09 | 9.006e-10 | 289.9917 | 0.0034484 | 33.1231 | 0.68267 | 0.53502 | 0.026721 | #DIV/0! | 0.13145 | #DIV/0! | 0.50407 | -0.13417 | 7.6076 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0xE0E5 | pH | 2019-03-11 23:46 | 1552347931 | -0.06 | 16.8738 | 33.1263 | 19 | 2 | 0 | 21.04 | 47.34 | 11.928 | 0.0458 | 7.1234 | 3.662 | -0.93905 | 5e-05 | -0.97222 | 2.5e-05 | 7.616e-09 | 1.51e-09 | 290.0238 | 0.003448 | 33.1263 | 0.68274 | 0.53507 | 0.026724 | #DIV/0! | 0.13129 | #DIV/0! | 0.50409 | -0.13418 | 7.6167 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0x9E7A | pH | 2019-03-11 23:49 | 1552348110 | -0.07 | 16.8954 | 33.1216 | 21 | 2 | 0 | 21.04 | 47.85 | 11.928 | 0.046 | 7.1234 | 3.662 | -0.93895 | 3.3e-05 | -0.97261 | 2.8e-05 | 7.051e-09 | 9.031e-10 | 290.0454 | 0.0034477 | 33.1216 | 0.68264 | 0.53499 | 0.02672 | #DIV/0! | 0.13117 | #DIV/0! | 0.50411 | -0.13419 | 7.6238 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0x95ED | pH | 2019-03-11 23:52 | 1552348291 | -0.08 | 16.9546 | 33.123 | 23 | 2 | 0 | 21.06 | 47.36 | 11.932 | 0.0458 | 7.1234 | 3.662 | -0.93877 | 3.8e-05 | -0.97289 | 2.9e-05 | 8.794e-09 | 3.73e-09 | 290.1046 | 0.003447 | 33.123 | 0.68267 | 0.53502 | 0.026721 | #DIV/0! | 0.13087 | #DIV/0! | 0.50416 | -0.13422 | 7.6412 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
0xBF7E | pH | 2019-03-11 23:55 | 1552348471 | -0.08 | 16.9946 | 33.1248 | 25 | 2 | 0 | 21.07 | 47.49 | 11.932 | 0.0456 | 7.1234 | 3.662 | -0.93874 | 4.4e-05 | -0.97325 | 3.8e-05 | 7.004e-09 | 1.256e-09 | 290.1446 | 0.0034466 | 33.1248 | 0.68271 | 0.53505 | 0.026723 | #DIV/0! | 0.13067 | #DIV/0! | 0.50419 | -0.13423 | 7.6529 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Note on Z19: Hans Jannasch: Use 35 for buffers; CTD underestimates Cl from S.
Note on AA19: Johnson, Ken: From Dickson, Sabine, Christian Best Practices, eqn. 34, Chap. 5. This is in mol/kg H20 and is the desired form.
Note on AC19: Johnson, Ken: From Dickson, Sabine, Christian, Best Practices, Table 2, Chap. 5. Units mol/kg soln - for use with KHSO4
Note on AD19: =pH(TT)-.015*(20-C)
Note on AE19: Johnson, Ken: From Dickson, Sabine, Christian, Best Practices, eqn 33, Chap. 5. Units mol/kg solution
Note on AF19: Johnson, Ken: mol/kg solution
Note on AG19: Johnson, Ken: Empirical fit to data in Table II of Khoo et al Anal Chem 1977
Note on AH19: Johnson, Ken: From Khoo et al. Anal Chem. 1977 For molal conc. Mol/kg H2O
Note on AJ19: Johnson, Ken: Changed -was mol/kg soln, now mol/kg H2O
Note on AL19: =Vrs+R*ln10/F*T*log(aHfreeaCl)

## Summary
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6 |  |  |  | Ver 3.0 | Updated: | By:
|  |  |  | 2019-09-26 | 2021-03-16 | wape
[5 empty] | From | To
|  |  |  | PT Chamber: | 2020-10-26 12:50 | 2020-10-27 16:45
|  |  |  | SW cal: | 0 | 0
Drift |  | 0 | mV/d |  | PT Channel | B
[5 empty] | MSC Logger | xxxxx
k2 calcs |  | 0 | 1600 |  | CTD# | xxxxx
|  | k2 used | -0.00096427 |  | APEX Float# | xxxxx
|  | r2= | #VALUE!
|  | k0 HCl | #VALUE! |  | k0 SW | #N/A
MBARI DuraFET DF#313 + VL#313 PEEK Ver 6
f(P) calcs |  | k6 | k5 | k4 | k3 | k2 | k1 | k0 | R^2 |  | By interpolating Vrs signals to P=0:
| 2 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  |  |  | k0
| 4 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | Vrs-Vc |  |  | #VALUE!
| 6 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | Vrs-Vc-k2T |  |  | #VALUE! | [9 empty] | Engineering Data
| 20 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  | Vrs-Vc-k2T (Up Only) |  |  | #VALUE! | [9 empty] | Time range: | 2020-10-26 12:50 | 2020-10-27 16:45
| Avg All | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | [15 empty] | Average | StDev
[9 empty] | Bromine Offset | [14 empty] | Vrs noise | 0.00011248 | 3.1427e-05
k0 calcs |  | k0 SW | StDev | St Dev pH | k0 comparison (PT and k0 cals) |  | from f(P) calcs | k0 HCl w/Br | 0 |  |  |  |  | k0 w/Br | [9 empty] | Vk noise | 0.00016381 | 4.4221e-05
| SW Pon | #N/A | #N/A | #N/A | Ave PTC/PT | #VALUE! | 2 | #VALUE! |  |  | Vrs-Vc |  |  | #VALUE! | [9 empty] | Ik | -1.8634e-09 | 5.2168e-09
| SW Poff | #N/A | #N/A | #N/A | Ave PTC/f(P) | #VALUE! | 4 | #VALUE! |  |  | Vrs-Vc-k2T |  |  | #VALUE! | [9 empty] | Ib | 1.1468e-09 | 7.9278e-09
| Delta | #N/A | Pump effect | #N/A | Ave SW | #N/A | 6 | #VALUE! |  |  | Vrs-Vc-k2T (Up Only) |  |  | #VALUE! | [9 empty] | f(P) r^2 | #VALUE! | #VALUE!
[5 empty] | Delta k0 | #VALUE! | 20 | #VALUE! |  |  |  | Hysteresis (V) |  | #VALUE! | [9 empty] | k0 (all) | #VALUE! | #VALUE!
|  | k0 Cal run? | no |  | pH offset | #VALUE! | Avg All | #VALUE! |  |  |  | Hysteresis (pH) |  | #VALUE! | [9 empty] | k0 (up only) | #VALUE! | #VALUE!
#VALUE! | #VALUE! | [8 empty] | 10/26-27/2020 | B | 1@SO | 0 | -0.00096427 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | 1/0-0/1900 | #N/A | #N/A | #N/A | #N/A