## Slide 1: Progress in Oceanographic pH Sensors
Gene Massion, Ben Davis, Kenneth S Johnson, Jessica Metter, Josh Plant, Yuichiro Takeshita, and Peter Walz
Monterey Bay Aquarium Research Institute
Moss Landing, United States
Ocean Sciences Meeting 2026, Glasgow Scotland
2026 Feb 23

## Slide 2: Motivation
Defensible pH time series are critical to understanding ocean health challenges
  Ocean acidification: how fast,
  where and why?
  Marine Carbon Dioxide Removal: how effective and what is the impact on the surrounding ecosystem?
  Fishery management
  Coral reef health, …

## Slide 3: Background
MBARI’s chem sensor lab has developed one of the few viable oceanographic pH sensors
  Started in 2008
  ISFET semiconductor technology
  Mostly deployed autonomously: no operator intervention for 3-5+ years
  >670 sensor deployed
    Mostly on Argo floats

## Slide 4: Background
Commercialized in 2010 w/Seabird
We continue to improve our sensor
  Performance
  Manufacturability
  Reliability
  Cost

## Slide 5: 6 Generations of pH Sensors
Original MBARI Deep Sea DuraFET (DSD)
Commercial version of
MBARI DSD
Gasketed DuraFET (GDF)
Pressure balanced GDF
Lionix ISFET SURpH
sensor
Pressure balanced
SURpH sensor
G1 G3 G4 G5 G6

## Slide 6: Details
Increasing Pressure→
←Decreasing Pressure
Hysteresis
~1 mVolt (17 mPH)
Every sensor gets calibrated for pressure, temperature, bulk offset
Historical Issues
  Complicated ∆pH vs. pressure
  Sensor to sensor variability
  Hysteresis, Drift
Decreasing Pressure→
Increasing Presure→
f(p)
~15 mVolts

## Slide 7: Performance Progress
Pressure balancing (oil filling) helps (a lot)
  Same charts
Improved flow
  Still minimizing bubbles
  Flow path aimed directly at sensor face
  Improved response time

## Slide 8: Manufacturing and Cost Progress
Contract Manufacturers are great (mostly)
  We have transferred most of our manufacturing work to Applied Engineering
    10-50 of our pH sensors per year
    Better vendor control and incoming inspection processes
    Well documented assembly procedure process
    Better incoming inspection and rejected material process
    Trained and qualified manufacturing technicians
We’re still responsible for some problems
We’re getting sensors built for about the same price we were building them for (maybe less)

## Slide 9: Reliability Progress
Design and Manufacturability Improvements = Better reliability
XX% of sensors deployed since 20xx still operating
XY% since 20xy

## Slide 10: Reliability Progress
Reliability Engineers use different metrics:
Failures in Time (FIT) is one metric
  Number of failures per billion operating hours
  Calculated or measured and extrapolated
FIT rate for all Argo pH sensors =
FIT rate for GDF sensors = 7092
For reference:
  FIT rate for a Mil-Spec power supply ~ 300-1000
  FIT rate for submarine telecom repeaters <10
    99.912% survive out of 1000 units after 10 years

## Slide 11: Current Status
Hundreds of sensors deployed across the world’s oceans generating defensible 4D pH data
  Accuracy: ~0.01 pH units
  Precision: ~0.005 pH units
All data is publicly available along with other essential ocean variables through Argo Data Centers
  With rigorous on-going QC processing

## Slide 12: Current Work
Move pressure balanced design into production
Reduce the variance in performance metrics
Combined pH and total alkalinty sensor
Understand and mitigate reliability issues
  Shipping mortality, non-electronics failures
Predict seawater pH as an function of sensor mechanical and electrical design

## Slide 13: Acknowledgments
The MBARI pH sensor village includes
The entire Chem Sensor lab
The entire Ocean Biogeochemical Sensing lab,
The MBARI technicians: Chris Beebe, Paul Coenen, Yamir Hernandez, Roman Marin, James McClure, Jim Montgomery, Jose Rosal,
The MBARI machine shop: Jerry Allen, Ray Thompson, Mark Tynes
Generous funding from the David and Lucille Packard Foundation and the US National Science Foundation.

## Slide 14: Abstract
The MBARI Chemical Sensor lab has been developing pH sensor for oceanographic applications since 2008. The original technology was commercialized in 2010. Since that time, the Chem Sensor lab has continued to improve our pH sensor technology. More than 670 sensors have been deployed for a range of science missions, and various failure modes have been identified. Improvements have been made in performance, manufacturability, reliability and cost, which address many of these issues. The current status of and future plans for our continual improvement process will be discussed in this presentation. Performance improvements will be of general interest to the ocean sciences community. The processes we've developed to improve manufacturability, reliability and cost may be of particular interest to the ocean sciences technology development community.