## Slide 1: Quantifying GDF Performance
Goal: Build a data base of relevant GDF and SURpH sensor performance measurements extracted from the cal runs so we can semi-quantitatively call a sensor good or bad relative to selected subsets of all the previous data
  This GDF vs. all delivered GDFs
    Decreasing pressures only
    Increasing pressures only
    Both pressure ramps
  This GDF vs. all AE built GDFs
  This GDF vs. all broached corner GDFs (just an example)

## Slide 2: Current Calibration Run

## Slide 3: Non-ideal Behavior (0.001 volt ~= 0.017 pH)
Burn in or
warmup?
Burn in/warmup/drift (sensor or solution)?
Random Noise?

## Slide 4
Pressure
Vrs
Definitions
Cycle = 0-2000-0 dBar with an increasing pressure ramp and a decreasing pressure ramp
Residuals Increasing = Vrsmeasured – Vrsfit for increasing ramp
Residuals Decreasing = Vrsmeasured – Vrsfit for decreasing ramp
Residuals Both = Vrsmeasured – Vrsfit for the entire cycle (both increasing and decreasing ramps)
Data set grouping:
Individual ramps
Individual cycles
All incr ramps, post warmup incr ramps
All decr ramps, post warmup decr ramps corrected for drift
All cycles, post warmup cycles
Calculations
Fits do not include pressure hold data
Residuals for 1-5 above (vector)
Mean of residuals for 1-5 above (scalar)
Maximum residuals for 1-5 above (scalar)
Decr P residual
Incr P residual

## Slide 5: 4. Mean Decreasing Pressure Residual per Cycle (This is what we use for the deployed fit)
Data set grouping:
Individual ramps
Individual cycles
All incr ramps, post warmup incr ramps
All decr ramps, post warmup decr ramps
All cycles, post warmup cycles
Calculations
Residuals for 1-5 above (vector)
Mean of residuals for 1-5 above (scalar)
Maximum residuals for 1-5 above (scalar)
Decr P residual

## Slide 6: One vs. Many Script

## Slide 8
One f(P) curve from a Deployed GDF (or AE) sensor
All f(P) curves from All Deployed GDF (or AE) sensors

## Slide 14: Bad Sensor

## Slide 15: Bad Sensor
Bad
Good
(reworked)

## Slide 20: What do Mike Trout and the GDF pH sensor have in common??

## Slide 21: OPS+

## Slide 22
When complete, develop a ‘red-line’ number?
Notice GDF32 has highest OPS+ (Davis-Trout) number, failed in situ

## Slide 23: First SURpH Ver2 Sensor with correct pre-load pTC run
First cycle “sets” the piston?
NOT corrected for P or T

## Slide 24: First and Second 15 degC cycles
red = first set, increasing pressure
green = first set, decreasing pressure
blue = second set, increasing pressure
black = second set, decreasing pressure

## Slide 25: How can we evaluate “goodness of fit” Graphical:
Generally speaking, graphical measures are more beneficial than numerical measures because they allow you to view the entire data set at once, and they can easily display a wide range of relationships between the model and the data. Ref: Mathworks
Biggest errors located at the extremes

## Slide 26: Numerical:
The numerical measures are more narrowly focused on a particular aspect of the data and often try to compress that information into a single number. Ref: Mathworks

## Slide 27: Preliminary recommendations for each new sensor
Plot this sensor final decreasing f(P) vs all delivered sensors final decreasing f(P) (1X std dev)
Plot this sensor residuals from final decreasing f(P) fit vs all delivered sensors final decreasing f(P) fit residuals (1X std dev)
Keep track of the histogram of “Davis-Trout” values in the hopes of coming up with a pass/no-pass value or at least a definitely pass value, a definitely no-pass value and a band of questionable values

## Slide 28
https://www.mathworks.com/help/curvefit/evaluating-goodness-of-fit.html

## Slide 29
https://www.mathworks.com/help/curvefit/evaluating-goodness-of-fit.html