## Slide 1: Quantifying GDF Performance
Goal: Build a data base of relevant GDF (Honeywell and Lionix) performance measurements so we can semi-quantitatively call a sensor good or bad relative to selected subsets of all the previous data i.e, this GDF vs all other deployed GDFs

## Slide 2: Quantifying GDF performance
Performance Measurements:
  Vrs (volts measured between reference and ISFET source) vs Pressure “goodness of fit” i.e., how large are the residuals between the measured Vrs values and the polynomial fit
  Vrs vs Temperature “goodness of fit”
  Overall range of Vrs as a function of pressure
  Overall range of Vrs as a function of temperature
  Vrs drift vs time

## Slide 3: Definitions

## Slide 4
Pressure
Vrs
Residuals from decreasing pressure fit
Residuals from increasing pressure fit
Residuals from fit to both increasing and decreasing pressure

## Slide 5: Current Calibration Run
Plot of Vrs with pressure and temperature

## Slide 6: One vs Many
Blur plot of “good” sensor
Blur plot of “bad” sensor

## Slide 7: Ben’s section

## Slide 13: First SURpH Ver2 Sensor with correct pre-load pTC run
First cycle “sets” the piston?

## Slide 14: Non-ideal Behavior (0.001 volt ~= 0.017 pH)
Burn in or
warmup?
Burn in/warmup/drift (sensor or solution)?
Random Noise?

## Slide 15: First and Second 15 degC cycles
red = first set, increasing pressure
green = first set, decreasing pressure
blue = second set, increasing pressure
black = second set, decreasing pressure

## Slide 16: How can we evaluate “goodness of fit” Graphical:
Generally speaking, graphical measures are more beneficial than numerical measures because they allow you to view the entire data set at once, and they can easily display a wide range of relationships between the model and the data.

## Slide 17: Numerical:
https://www.mathworks.com/help/curvefit/evaluating-goodness-of-fit.html
The numerical measures are more narrowly focused on a particular aspect of the data and often try to compress that information into a single number

## Slide 20: One measure of goodness of fit
For all deployed sensors:
  Fit a polynomial to all the decreasing pressure Vrs fit data across all depths and temperatures
  Calculate the standard deviation for every depth bin
Plot the sensor under test’s decreasing Vrs fit data against the N*std dev Vrs data for all sensors (one vs many) plot

## Slide 21: Another measure of goodness of fit
For all deployed sensors:
  Calculate the square error (Vrsmeasured – Vrsfit)2 for every sensor’s Vrs values across all depths and temperatures
  Calculate the standard deviation for every depth bin
Plot the sensor under test’s Vrs residuals against the N*std dev residuals for all sensors (one vs many) plot