## Slide 1: Quantifying GDF Performance
Goal: Build a data base of relevant GDF (Honeywell and Lionix) performance measurements so we can semi-quantitatively call a sensor good or bad relative to selected subsets of all the previous data i.e, this GDF vs all other deployed GDFs

## Slide 2: Quantifying GDF performance
Performance Measurements:
  Vrs (volts measured between reference and ISFET source) vs Pressure “goodness of fit” i.e., how large are the residuals between the measured Vrs values and the polynomial fit
  Vrs vs Temperature “goodness of fit”
  Overall range of Vrs as a function of pressure
  Overall range of Vrs as a function of temperature
  Vrs drift vs time

## Slide 3: Current Calibration Run

## Slide 4
Pressure
Vrs
Definitions
Cycle = 0-2000-0 dBar with an increasing pressure ramp and a decreasing pressure ramp
Residuals Increasing = Vrsmeasured – Vrsfit for increasing ramp
Residuals Decreasing = Vrsmeasured – Vrsfit for decreasing ramp
Residuals Both = Vrsmeasured – Vrsfit for the entire cycle (both increasing and decreasing ramps)
Data set grouping:
Individual ramps
Individual cycles
All incr ramps, post warmup incr ramps
All decr ramps, post warmup decr ramps
All cycles, post warmup cycles
Calculations
Residuals for 1-5 above (vector)
Mean of residuals for 1-5 above (scalar)
Maximum residuals for 1-5 above (scalar)
Decr P residual
Incr P residual

## Slide 5: Ben’s section

## Slide 11: First SURpH Ver2 Sensor with correct pre-load pTC run
First cycle “sets” the piston?

## Slide 12: Non-ideal Behavior (0.001 volt ~= 0.017 pH)
Burn in or
warmup?
Burn in/warmup/drift (sensor or solution)?
Random Noise?

## Slide 13: First and Second 15 degC cycles
red = first set, increasing pressure
green = first set, decreasing pressure
blue = second set, increasing pressure
black = second set, decreasing pressure

## Slide 14: How can we evaluate “goodness of fit” Graphical:
Generally speaking, graphical measures are more beneficial than numerical measures because they allow you to view the entire data set at once, and they can easily display a wide range of relationships between the model and the data. Ref: Mathworks

## Slide 15: Numerical:
The numerical measures are more narrowly focused on a particular aspect of the data and often try to compress that information into a single number. Ref: Mathworks

## Slide 16
https://www.mathworks.com/help/curvefit/evaluating-goodness-of-fit.html

## Slide 17
https://www.mathworks.com/help/curvefit/evaluating-goodness-of-fit.html

## Slide 18: One measure of goodness of fit
For all deployed sensors:
  Fit a polynomial to all the decreasing pressure Vrs fit data across all depths and temperatures
  Calculate the standard deviation for every depth bin
Plot the sensor under test’s decreasing Vrs fit data against the N*std dev Vrs data for all sensors (one vs many) plot

## Slide 19: Another measure of goodness of fit
For all deployed sensors:
  Calculate the square error (Vrsmeasured – Vrsfit)2 for every sensor’s Vrs values across all depths and temperatures
  Calculate the standard deviation for every depth bin
Plot the sensor under test’s Vrs residuals against the N*std dev residuals for all sensors (one vs many) plot