GDF pH Sensor Build Traveler

| Serial Number1: | 2020-001 (also named GDF#2) |  |  |
| --- | --- | --- | --- |
| Version2: | IR | Gasket thickness: | 0.020” silicone |
| Installed on Float WMO Number: |  | Coverplate used: | -020 x 0625 |

| Build Steps |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Step No. | Description |  | Date | By | Notes |  |  |
| 1 | Kit parts |  | 8/10/20 | PW | Same build as GDF#1 on float 18340 |  |  |
| 2 | Visual Inspection |  | 8/10/20 | PW |  |  |  |
| 3 | Dimension checks |  | 8/10/20 | PW |  |  |  |
| Dim A | Dim B | Dim C | Dim D |  | Dim E | Dim F | Dim G |
| 0.057 | 0.045 | 0.098 | 0.297 |  | 0.061 | 0.0187 |  |
| 4 | Gasket thickness and material |  |  |  | Silicone sheet 0.020” |  |  |
| 5 | Build sensor |  | 8/10/20 | PW | JM wired pcb – half moon board |  |  |
| 6 | Photo front and back, no covers3 |  |  |  |  |  |  |
| 7 | Other photos3 |  |  |  |  |  |  |
| 8 | VL pellet resistance |  | 8/10/20 | PW | < 100 kOhm |  |  |
| 9 | ISFET diode measurement |  | 8/10/20 | PW | 0.607 white / 0.608 black |  |  |
| 10 | Counter electrode resistance |  | 8/10/20 | PW | 0.3 Ohm |  |  |
| 11 | Seawater Conditioning |  |  |  |  |  |  |
| 12 | Test / Calibration k2 and f(P) |  |  |  | Yui to test in lab / beaker expts Aug 2020 |  |  |
| 13 | Install on endcap |  |  |  |  |  |  |
| 14 | Test / Calibration K0 |  |  |  |  |  |  |
| 15 | Shipped to |  |  |  |  |  |  |
| 16 | Deployed |  |  |  |  |  |  |

Notes:

1 YYYY-xxx: YYYY = year, xxx = number of GDF build that year e.g., 2019-001, 2019-002, 2020-001, 2020-002, …

2 IR = Initial release (soldered pellet), A = rev A (fuzzy pellet)

3 Photos go in \\atlas.shore.mbari.org\ProjectLibrary\pH2018\Production\YYYY-xxx\photos  YYYY-xxx is the serial number.

.