## Slide 1
pH sensors
What a long way we have come
(in a very short amount of time)
Josh Plant Tanya Maurer Peter Walz Yui Takeshita Luke Coletti
& Ken Johnson
9th BGC Argo Meeting, ADMT 21
November 30 – December 1, 2020
Virtual

## Slide 2
Today!
900,000 QC’ed pH measurements
Largely due to SOCCOM project in the Southern Ocean
Enabling Southern Ocean estimates of CO2 flux, Aragonite Saturation
Just 5 years ago
35,000 QC’ed pH measurements
Largely due to SOCCOM project in the Southern Ocean

## Slide 3
pH pump offset
5906000, Cycle 9
Pump OFF
(spot samples)
Pump ON
(CP mode)
pH profiles exhibit a step change for some MBARI APEX floats
Occurs when CTD shifts to constant profiling mode (CTD pump continuously on) at 985m
Oxygen & Temperature profiles smooth & linear
pH above 985m drops by ~0.007
Magnitude can vary from float to float
Magnitude of offset can vary from profile to profile
Rarely seen in NAVIS floats

## Slide 4
pH pump offset estimates for SOCCOM array
Affects a small % of total profiles
Overall bias small & within sensor precision (-0.005)
Some floats affected more than others
-1 milli pH offset => +1 µatm pCO2
Find root cause & fix sensor
Correct existing data at GDAC affected by problem

## Slide 5
Find root cause & fix sensor
Partially Related to flow over reference electrode or FET (preliminary lab work)
Good news – Mostly isolated to 2018-2019 deployment season
Interesting behavior over time (2018-2019 deployments)
New sensor design may eliminate the problem completely
2019-2020
&
2020 -2021
Deployment seasons
2018-2019
Deployment season
% profiles with
abs(offset) > 0.01
Cycle #
Cycle #
Total profiles

## Slide 6
Correct existing data stored in the GDAC
Single, static offset per float
B) Perform adjustments
in SAGE at 980 m vs
1500m
C) Auto-calculation of
offset per float, per
profile
  Correction methods that have been considered:

## Slide 7
Correct existing data stored in the GDAC
Single, static offset per float
B) Perform adjustments
in SAGE at 980 m vs
1500m
C) Auto-calculation of
offset per float, per
profile
  Correction methods that have been considered:
  Original method when few floats were affected and observed magnitudes were on the order of millipH
  Does not address time-varying nature of offset

## Slide 8
Correct existing data stored in the GDAC
Single, static offset per float
B) Perform adjustments
in SAGE at 980m vs
1500m
C) Auto-calculation of
offset per float, per
profile
  Correction methods that have been considered:
  Maintains accuracy of adjusted surface data
  Deep data would remain unadjusted
  Assumes offset in SAGE is independent of ref depth (980 vs 1500m)

## Slide 9
Correct existing data stored in the GDAC
Single, static offset per float
B) Perform adjustments
in SAGE at 980m vs
1500m
C) Auto-calculation of
offset per float, per
profile
  Correction methods that have been considered:
  No algorithm is perfect
  Thresholds must be set
  How to record magnitude of correction and impact on error/uncertainty??
  Delayed mode or real time corrections?
  Metric for success?

## Slide 10
Correct existing data stored in the GDAC
C) Auto calculation of pump offset – one possible approach………
Assume pH is a smooth function of pressure over a narrow data window centered at 985m
Break data window above & below 985m
Add offset to shallower part of data window
Iterate offset value to optimize polynomial fit over data window (minimize SSR)
Calculate a signal to noise ratio = offset / std(fit residuals)
Set offset & signal to noise thresholds (+/- 0.005 & +/- 10 ??)
Add offsets which exceed thresholds to all data shallower than 985m
Raw
data
7th
iteration
final
iteration

## Slide 11
5906000 Cycle 9
Negative pH offset (-0.03)
High signal to noise ratio (-21)
Apply correction
Example

## Slide 12
5905998 Cycle 2 - No offset
Offset small
Signal to noise ratio low (2.4)
No correction
applied
Example