## GDF031
Serial Number | GDF 031 | Probable Diagnosis of Failure
Date of Failure | 2023-06-21 | Ik was railed from the start of profile two, indicating a bad connection to counter electrode. Combined with a decent first profile, this indicates failre due to environemntal stress at park depth for a long duration (~10 days) at very low temperatuers (~0.5 C).
Deployment Duration | 2 Profiles
Platform | APEX Float
Average Max Depth (m) | 2000
Deployment Temp Range (C) | 0-10
Data Location | \\atlas\Chem\DuraFET\GDF Production\Performances\GDF031\GDF31_081822_ChC
Autopsy Permorfmed? (Y/N) |  | Autopsy Observations
Lab Performance Metrics
f(P) at 2000 dBar (mV) |  |  | ∆k2(P) (uV/C)
∆f(P) at 2000 dBar (mV) |  |  | Avg Hysteresis Magnitude per Cycle (mV)
GDF Production Traveler
Serial Number: |  | GDF 032
Version2: |  | rev A with machined corners
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  |  | MB
2 | Visual Inspection |  |  | MB
3 | Coverplate |  | tall shoulder | MB
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0265 | 0.0355 | 0.298 | 0.097 | 0.0606 | 0.0184 | 0.301
4 | Gasket thickness |  | 0.0313
| Gasket material |  | Viton
| Gasket compression |  | -0.0068
5 | Build sensor |  | 2022-08-11
6 | Photos front and back3 |  | 2022-08-11
7 | Other photos3 |  | 2022-08-11
8 | VL pellet resistance |  | 2022-08-11
9 | ISFET diode measurement |  | 2022-08-11
10 | Counter electrode resistance |  | 2022-08-11
11 | Seawater Cond. (num. days) |  | 2022-08-11
12 | K2 cal (file name) |  | 2022-08-25
13 | f(p) cal (file name) |  | 2022-08-25
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed

## GDF032
Serial Number | GDF 032 | Probable Diagnosis of Failure
Date of Failure | 2023-04-20 | Ib was railed from the start of profile two, along with Vk and Ik beginning to show odd behavior before fully failing on the second profile. Combined with a decent first profile, this indicates failre due to environemntal stress at park depth for a long duration (~10 days) at very low temperatuers (~0.5 C).
Deployment Duration | 2 Profiles
Platform | APEX Float
Average Max Depth (m) | 2000
Deployment Temp Range (C) | 0-10
Data Location | \\atlas\Chem\DuraFET\GDF Production\Performances\GDF032\GDF32_081822_ChD
Autopsy Permorfmed? (Y/N) |  | Autopsy Observations
Lab Performance Metrics
f(P) at 2000 dBar (mV) |  |  | ∆k2(P) (uV/C)
∆f(P) at 2000 dBar (mV) |  |  | Avg Hysteresis Magnitude per Cycle (mV)
GDF Production Traveler
Serial Number: |  | GDF 032
Version2: |  | rev A with machined corners
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  |  | TW
2 | Visual Inspection |  |  | TW
3 | Coverplate |  | tall shoulder | TW
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.0267 | 0.0358 | 0.298 | 0.0975 | 0.0606 | 0.0186 | 0.301
4 | Gasket thickness |  | 0.0313
| Gasket material |  | Viton
| Gasket compression |  | -0.0069
5 | Build sensor |  | 2022-08-11
6 | Photos front and back3 |  | 2022-08-11
7 | Other photos3 |  | 2022-08-11
8 | VL pellet resistance |  | 2022-08-11
9 | ISFET diode measurement |  | 2022-08-11
10 | Counter electrode resistance |  | 2022-08-11
11 | Seawater Cond. (num. days) |  | 2022-08-11
12 | K2 cal (file name) |  | 2022-08-25
13 | f(p) cal (file name) |  | 2022-08-25
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed

## GDF006
Serial Number | GDF 006 | Probable Diagnosis of Failure
Date of Failure | 2023-07-24 | Vrs is in the right range, but it seems to have developed a big lag in response. We got the sensor back in the lab and tested it this morning, and Vk_std is very high (mV), and Ik is erratic. Ib looks fine. Wire walker has fast ascent/decent rate and prone to many vibrations that can harm sensor. Will autopsy when returned. Sensor evaluated at MBARI, initial checkout and test thru deployed electronics were consistent with earlier issues. Sensor surfaces (pellet and isfet) cleaned by Peter W and tested in beaker of seawater. After these cleaning steps the sensor behavior and repsonse was normal. Additional tests in Takeshita lab by Joe W, also showed sensor to be working as expected. Sensor returned to Humboldt for another deployment on wire walker (Aug 2023).
Deployment Duration | 4 months
Platform | Wire Walker
Average Max Depth (m) | 80
Deployment Temp Range (C)
Data Location | \\atlas\Chem\DuraFET\GDF Production\Performances\GDF006\GDF6_secondCal
Autopsy Permorfmed? (Y/N) |  | Autopsy Observations
yes |  | see comments above. Cleaning of isfet surface with methanol, and polishing of ref pellet seems to have caused the sensor to recover and now it works normally
Lab Performance Metrics
f(P) at 2000 dBar (mV) |  |  | ∆k2(P) (uV/C)
∆f(P) at 2000 dBar (mV) |  |  | Avg Hysteresis Magnitude per Cycle (mV)
GDF Production Traveler
Serial Number: |  | GDF 006
Version2: |  | rev A broached corner
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts |  | 2021-03-01 | PW
2 | Visual Inspection
3 | Coverplate |  | tall shoulder
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
0.028 | 0.036 | 0.295 | 0.096 | 0.061 | 0.019 | 0.301
4 | Gasket thickness |  | 0.03125
| Gasket material |  | 1/32" viton sheet
| Gasket compression |  | -0.00625
5 | Build sensor |  | 2021-03-10 | JM
6 | Photos front and back3
7 | Other photos3
8 | VL pellet resistance |  | 3/22/21 | PW | 100 kOhm
9 | ISFET diode measurement |  | 3/22/21 | PW | 0.611/0.611
10 | Counter electrode resistance |  | 3/22/21 | PW | 0.2 Ohm
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed
18 | ∆f(P) (mV) |  | 9 | BD | fn: _alltemp

## Template FailureLog
Serial Number |  | Probable Diagnosis of Failure
Date of Failure
Deployment Duration
Platform
Average Max Depth (m)
Deployment Temp Range (C)
Data Location | \\atlas\Chem\DuraFET\GDF Production\Performances\GDF032\GDF32_081822_ChD
Autopsy Permorfmed? (Y/N) |  | Autopsy Observations
Lab Performance Metrics
f(P) at 2000 dBar (mV) |  |  | ∆k2(P) (uV/C)
∆f(P) at 2000 dBar (mV) |  |  | Avg Hysteresis Magnitude per Cycle (mV)
GDF Production Traveler
Serial Number:
Version2:
Installed on Float WMO:
Step No. | Description |  | Date | By | Notes
1 | Kit parts
2 | Visual Inspection
3 | Coverplate
Dim A | Dim B | Dim C | Dim D | Dim E | Dim F | Dim G
4 | Gasket thickness
Gasket material
| Gasket compression |  | -0.019
5 | Build sensor
6 | Photos front and back3
7 | Other photos3
8 | VL pellet resistance
9 | ISFET diode measurement
10 | Counter electrode resistance
11 | Seawater Cond. (num. days)
12 | K2 cal (file name)
13 | f(p) cal (file name)
14 | Install on endcap
15 | Test / Calibration K0
16 | Shipped to
17 | Deployed