ZEISS LIGHTSOURCE ASS’Y FOR SPECTROPHOTOMETERS	

2018-JIM SCHOLFIELD		X:\ChemSensors\JimS\coletti\Spectr. mount-APEX
This Procedure describes how to properly secure the light sources in the mount plate.
-Assemble the necessary parts and tools	-
	Mount Plates-with finish machining
	Zeiss light source-be aware of gentle bends in fiber optics
	SS M2 x 2.5 set screws,  2 per ass’y
	0.035” Allen key  for above
	Multiple substrate epoxy with 15-30 minute work time
	Eye magnifier recommended, screws are really small
	Special made alignment tool is helpful
	Storage Rack is helpful as well
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-Install a set screw in each threaded hole, just enough to hold in place.
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-Place a drop of epoxy in 2 places on land in mount plate
[image: X:\ChemSensors\JimS\coletti\Spectr. mount-APEX\sp2.jpg]
												p.3
[bookmark: _GoBack]-Slide the light source into place, minding the fiber optics,  Place Alignment tool into position against the flat on the source.   Move/rotate  the source slightly to align the flat surface. Gently tighten both set screws.  The screws hold the source in position and allow positive grounding to the mount plate.
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-Check for proper alignment of ass’y, Move ass’y to curing/storage rack.  This is a good place to store most supplies.
[image: X:\ChemSensors\JimS\coletti\Spectr. mount-APEX\sp6.jpg]
	
image6.jpeg




image1.jpeg




image2.jpeg




image3.jpeg




image4.jpeg




image5.jpeg




