ZEISS LIGHTSOURCE ASS’Y FOR SPECTROPHOTOMETERS

2018-JIM SCHOLFIELD		X:\ChemSensors\JimS\coletti\Spectr. mount-APEX

This Procedure describes how to properly secure the light sources in the mount plate.

-Assemble the necessary parts and tools	-

Mount Plates-with finish machining

Zeiss light source-be aware of gentle bends in fiber optics

SS M2 x 2.5 set screws,  2 per ass’y

0.035” Allen key  for above

Multiple substrate epoxy with 15-30 minute work time

Eye magnifier recommended, screws are really small

Special made alignment tool is helpful

Storage Rack is helpful as well

p.2

-Install a set screw in each threaded hole, just enough to hold in place.

-Place a drop of epoxy in 2 places on land in mount plate

p.3

-Slide the light source into place, minding the fiber optics,  Place Alignment tool into position against the flat on the source.   Move/rotate  the source slightly to align the flat surface. Gently tighten both set screws.  The screws hold the source in position and allow positive grounding to the mount plate.

-Check for proper alignment of ass’y, Move ass’y to curing/storage rack.  This is a good place to store most supplies.