Ver. 1, 6/10/2013

An open source CO2 sensor for education/research

Goal:   develop a sensor for dissolved CO2 using either Senseair K-33 or GE Telaire T6615 sensor modules that can be deployed for time periods of at least one month with hourly measurements of dissolved CO2 that uses widely available components and which can be easily assembled by a technically competent student (but not necessarily an engineer).

Requirements

- Instrument should be submersible to several meters depth.  Water tight housing should be easily manufactured – no complex machining required.

- Data logger interfaced to CO2 sensor module.  Needs real time clock, persistent data storage.

- Data logger to be “low power”.  An Arduino Uno draws about 15 mA when in sleep mode (power supply and USB interface are always on) and that is probably too much if the whole system must run 1 month (1 D cell has 17,000 mamp-hr which gives only about 1000 hours = 40 days just to power system in sleep mode).  The Jeenode is much better, but an Arduino might be a fallback option.

- Wireless data transfer to eliminate underwater connectors.  Note radio signals don’t penetrate seawater, so instrument has to be at surface to offload data.

- Humidity and temperature sensors.  Note the K-33 has these (it also has some data memory).  They would have to be provided for T6615 (we a Honeywell HIH 4030 for humidity and TMP-37 for temperature – both need a 5 V A/D channel which we don’t have?).

- 5 V power – start with 4 D cell batteries.  Might need a 5 V regulator.

- CO2 equilibrator.   Smaller volume will give faster response rate.  Might need a pump to circulate gas phase.  Might need a trap to dry air – both the K-33 and T6615 will not work if humidity is 100%.

- Response time – want <1 hour to reach 99% of response to a step change in CO2 concentration.

www.co2meter.com is a good resource for K-33 knowledge, including calibration procedure and interface to Arduino:

http://www.co2meters.com/Documentation/AppNotes/AN126-K3x-sensor-arduino-uart.pdf

First Steps

- Start talking to the K-33 and/or T6615 with the Jeenode (or Arduino if you want to start with the basics).

- Construct basic, battery powered CO2 data logger (date, time, CO2, temperature, humidity, battery voltage (easy on Arduino but do we have enough ports on a Jeenode to include the Analog Plug)?)

- Experiment with CO2 equilibrator to assess response time, stability, configuration, etc.  Does it need an air pump, water trap….

- Mechanical design – housing and an anti-fouling system (maybe using Cu/Ni tubing).