### A Low Cost, Underwater  CO2 Sensor For Education And Research

### Miguel Uzcategui, University Of The District Of Columbia

##### Mentor: Ken Johnson

## Summer 2013

INDEX

- INTRODUCTION......................................................................................... 4

- COMPONENTS............................................................................................. 6

Microcontroller ................................................................................ 6

RTC plug (clock)............................................................................... 6

Memory plug...................................................................................... 7

UART plug......................................................................................... 8

Relays.................................................................................................. 8

Air pump............................................................................................ 8

CO2  sensor (Telaire 6615F).............................................................. 9

Temperature and humidity Sensor (Sensirion SHT21).................. 10

Nafion tubing...................................................................................... 10

Teflon AF tube.................................................................................... 11

- EXPERIMENTS............................................................................................. 12

Determining the equilibration rate of  the Teflon AF..................... 12

Power consumption............................................................................ 14

Pricing.................................................................................................. 16

- RESULTS........................................................................................................ 17

Test tank.............................................................................................. 17

Harbor................................................................................................. 17

- TUTORIAL.................................................................................................... 20

Assembling the Jeenode..................................................................... 20

Assembling the circuitry together..................................................... 20

Battery pack....................................................................................... 23

Housing............................................................................................... 24

Tubing layout...................................................................................... 26

Code..................................................................................................... 27

User Interface...................................................................................... 30

- CONCLUSION............................................................................................... 31

- REFERENCES................................................................................................ 32

- APPENDIX........................................................................................................33

Appendix A............................................................................................ 34

Appendix B............................................................................................. 37

Appendix C............................................................................................. 51

Abstract: Due to the increasing concern on how CO2  is affecting the seas, and lakes; an underwater CO2  sensing module was developed to fulfill this task . It works by using Teflon AF tubing to equilibrate the pCO2  that is underwater. The sensor used was a Telaire 6615 which is a flow through dual channel sensor that communicates to a microcontroller using UART protocol. By putting the device on the Moss Landing harbor, it was determined that the sensor system works in an accurate way, where there was a constant fluctuation of pCO2  between 470 ppm and 530 ppm which is correlated to the tide.  The conclusion of this project, is that a reliable, low cost underwater CO2  sensing module can be built using Teflon  AF as an equilibrator, in the future it is hoped that the sensing module could serve for future small researches or simply for education.

INTRODUCTION

In the present, humanity have started to realize how important is the chemistry of the water for all the living species on earth. Carbon dioxide plays an important role in the chemistry of our seas, lakes, and rivers (NASA). Many studies are being done to produce better sensors or better algorithms that will detect underwater carbon dioxide levels more efficiently and with more precision (ATABEK) (Ning, Li and Li).   The reason of all this studies is simple, to have a better understanding of how carbon dioxide affects life underwater.  Some species are more suited to live in an environment with more carbon dioxide than others. Different concentrations of  carbon dioxide could trigger different answers to a common stimulus among different species (Damage to Marine Ecosystems as CO2 Emissions Rise). Another reason to study CO2 , could be on how CO2  changes affect the metabolism of respiration and photosynthesis in different underwater plants.

The objective of this project is to develop  a low cost, easy to build, underwater, and low power consumption carbon dioxide sensor. It is not intended to be a high precision sensing device, but is good enough so  that some undergraduate or masters researches could use it to get good data on carbon dioxide concentration that could help them on their respective  researches. It could also be used by anybody with low technical skills who just want to conduct a small research for self-benefit.

In order to measure CO2  concentration underwater, the sensor will be introduced on a water tight housing, where a small Teflon AF tubing is going to be in touch with the water and  going into the sensor. The Teflon AF tubing serves as an equilibrator of Carbon dioxide, meaning that the same partial pressure of CO2  that it is underwater will diffuse through the Teflon AF tubing due to its special capabilities (Dasgupta).

The need of developing radio communications to the sensing module was necessary. This was because, as said before, the sensor will be put in a water tight housing; therefore, the less it is taken out of that housing the better. This is to avoid any human error that will damage the housing while opening and closing it, thus getting water in the sensor and the microcontroller when put into the water.

A complete explanation about the CO2  sensing mote  and the Teflon AF tubing will be addressed on the following sections.

COMPONENTS

Microcontroller (Jeenode):  The Jeenode is a microcontroller based on the Arduino concept that is designed for low power consumption. It has 4 ports in which you can connect any device using the  I2C protocol (Wikipedia, I2C), or also have analog and Digital pins for devices that do not use I2C protocol. The voltage supply is from 3.3 to 17V.  The key about the Jeenode is that is designed for low power consumption. The Jeenode consumes  9mA when is in normal operation and 0.8mA (Jeelabs, Jeenode SMD) while it is on Sleepy state; while the Arduino consumes  50 to 500mA while awake, depending on how is running, and  10 mA while it is on sleeping mode (Mitchell).

Figure 1 (pin out and port description of the Jeenode)

- RTC plug(Clock):  The clock plug is a Jeenode plug that works with the real time chip DS1340. It has a little battery backup so it keeps working after disconnected from the power source. It runs using an I2C bus running at 3.3V. So it can be connected in any Jeenode port. (Jeelabs, RTC plug)

Figure 2 ( picture of the RTC plug along with the Pinout description.)

- Memory plug:  The memory plug is a small board with 1 to 4 EPROM memory chips, each of them have a capacity of 128 Kbytes. It could have up to 512Kbytes of memory; however for this project  2 memory chips are used, which means there is up to 256 Kbytes in memory. The  memory plug uses an I2C bus with 3.3V to interface with any device or any of the four Jeenode ports. It is designed on a way that each memory chip is composed of pages and directories. Each memory chip has 1000 pages with a capacity of 256 bytes per page. (Jeelabs, Memory plug). There is no file system for reading or writing; therefore all the readings and writings have to be done manually by the user. The best way to do it, is by setting up a directory on page 0 byte 0 and save the number of page and the number of byte where you are at. So the next time an operation needs to be done,  the memory can know the location where it has to go.

Figure 3 ( picture of the memory plug along with the Pinout description.)

- UART plug: The UART plug just adds a serial port to the system. It is based on the SC16IS740 chip. It can be interconnected with any devices that needs an UART protocol to communicate up to a baud rate of  230400.  It runs using an I2C bus running at 3.3V. So it can be connected in any Jeenode port (Jeelabs, UART plug).

Figure 4 ( picture of the UART plug along with the Pinout description)

- Relay plug:  The relay plug is as the name says two simple relays that can be switched on and off  when desired. It connects to the Jeenode by using one of the Jeenode ports (Jeelabs, Relay plug).

Figure 5 ( picture of the relay plug along with the Pinout description)

- Air pump:  It is a diaphragm pump created by "KNF" which works with a supply voltage of 6V. It possesses two tubing exits and it does not allow any flow of air from the outside to enter in the main chamber of the pump. For more information, look at the datasheet that is on the website  listed in the "Work Cited" part of the paper. (KNF)

Figure 6 ( picture of the air pump)

- CO2  sensor: For this project, the sensor that is being used is the GE Telaire T6615F which is a dual channel sensor. The sensor works by shining a beam of infrared light, one of the channels measures the reference light that was shined while the other one measures the amount of infrared light that gets back (GE). Then, using Beer-Lambert Law (Wikipedia, Beer_Lambert Law) the concentration of CO2  can be known.  The Telaire 6615F comes with an tight air housing, with two little tubing exits, which does not allow any air to diffuse into the sensor, except by the air that is coming through the tubing.  It communicates using UART protocol (Wikipedia, Universal asynchronous receiver/transmitter) with a supply voltage (Vin) of 5V.

Figure 7 ( picture of the Telaire T6615F along with the Pinout description)

- Temperature and humidity Sensor (Sensirion SHT21):  It is a sensor developed by Sensirion, as the name says it measures the temperature and the humidity on the air.  It works using I2C protocol with a SDA and SCL to do the communications with the microcontroller.  It has a supply voltage (Vin) of 5V. (Sensirion)

Figure 8 (top-down view of SHT21 sensor)

- Nafion Tubing: In simpler terms Nafion is a Teflon backbone with occasional side chains of another fluorocarbon. The side chain terminates in a sulfonic acid (-SO3H). It works as  a water  absorbing trap. It can absorb water in its liquid and vapor phase. Each sulfonic  acid group can absorb up to 13 molecules of water, which are transported through the membranes of the tubing to the outside of the tubing. (Pure)

Figure 9 (Picture of Nafion tubing)

- Teflon AF tubing:  The Teflon AF tubing is a type of tubing made out of Teflon AF-2400. This type of material is permeable to many materials, specially to CO2  (Dasgupta).  It is made out of small membranes which allows the partial pressure of CO2  to equilibrate; therefore, it could act as an equilibrator in an aquatic medium. For this project a 1.6mm OD and 1.0 mm ID was used  (Biogeneral).

Figure 10 (Teflon AF tubing)

| Teflon® AF Gas Permeability, cB * |  |  |
| --- | --- | --- |
| Gas | Teflon® AF 2400 | PTFE |
| CO2 | 280,000 | 1200 |
| O2 | 99,000 | 420 |
| H2 | 220,000 | 980 |
| N2 | 49,000 | 140 |
| * cB = centiBarrer = 10-8 x (cm2 - cm) / (cm - Hg - sec - cm2) |  |  |

Figure 11 (Permeability of Teflon AF 2400 tubing to different materials)

EXPERIMENTS

- Determining the equilibration rate of the Teflon AF

The whole Idea of the project is to create a sensor that can measure CO2  underwater. In order to do that, it has to be tested if the Teflon AF tubing equilibrates underwater. A series of experiments were done with different concentrations of CO2 ; first in air and finally underwater, to see if the tubing truly equilibrates, and if it does, find out the rate of equilibration.

The experiment setup was a glass flask with three tubing ports and one plastic cap that can seal completely the inside of the flask. The Teflon AF equilibration tube is placed in the flask  and connected to the sensor, and the air pump through two of  the ports, the gas supply goes into the last port. The idea was to create a closed loop  between the pump, the sensor and the testing environment. The supply was on for  a certain amount of time. When turned off the flask was closed so no air could go in or out. Air circulated through the air pump, next to the sensor and then to the flask again. The pump was left running the whole time through the experiment.

Figure 12( Experiment setup)

- Equilibration time on air

FIgure 13 (Equilibration in air of the Teflon AF tubing, first graph time constant 9.81 min,  second graph  time constant is 4.3 min)

- Equilibration time on water

Figure 14( Equilibration in Water of Teflon AF, first graph has a time constant of 4.73 min, Second graph has a time constant of 3.73min, Third graph has a time constant of 8.17 ppm/min)

- Power consumption: In order to measure how much time the circuit can run with batteries, first the mAh consumption of the circuit has to be found. A power meter was used in order to look at the amount of power and current that the circuit consumes at each  state. The circuit setup used for this experiment is as follows.

Figure 15 (Circuit to connect the power meter to the sensing module)

After measuring all the devices, the results were as follows:

    - Pump consumption:  111mA

    - CO2 Sensor while awake (lights on):  0.145 mA

    - Microcontroller +SHT21+ 7805 Voltage regulator+ CO2 Sensor  (sleep    	mode)=38.3 mA

Knowing  how much time each of the stages consume, it would be possible to calculate the  total consumption  of the sensing module for each possible time scenario. Nevertheless, the total power consumption calculated for this experiment was done for  a time cycle of an hour. Where  the pump will be turning on for 5 seconds every 3 minutes, the sensor will be in  sleep mode for 50 minutes, and finally the sensor will wake up to warm up for 10 minutes and then do a reading.

In order to know the power consumption in mAh of each state knowing the amount of seconds the state will be effective, all what is needed to do is multiply by 1 hour, multiply by the number of seconds the stage will be on, and finally divide by 3600 s.

    - Pump consumption:  111mA*1h*100s3600s=3.08mAh

    - Sensor while awake (lights on):  145mA*1h*150s3600s=13050mAs=6.04mAh

    - Microcontroller +SHT21+ 7805 Voltage regulator+ Sensor  (sleep    	mode):38.3mA*1h*3600s3600s=38.3mAh

Total power consumption= 47.42mA.h

If the module is operated with 6 D cell batteries which have an average energy of 17000 mA.h. It would be possible to run it for 14.5 days . Nevertheless, it is recommended  to replace the batteries every  13.5 days in order  to be sure that the batteries have not started to discharge, therefore causing the circuit to not run properly.

- Pricing:

| Jeenode | $27 |
| --- | --- |
| RTC plug | $15 |
| Relay plug | $14 |
| UART plug | $11 |
| Memory plug | $9.50 |
| Telaire 6615 | $115 |
| SHT21 | $40 |
| Air pump | $290 |
| 7805 voltage regulator | $2 |
| Switch | $4 |
| Teflon AF | $300 |
| Nafion tubing | $90 |
| PVC endcap | $1 |
| PVC tubing | $1 |
| PVC pipe locker | $3 |
| Cooper nickel tube | $5 |
| Total | $927 |

RESULTS

- Test tank: The module was put in a big tank of around 10 meters of deepness. After 12 hours in the test tank, this are the results that were collected

Figure 16 (Results from the test tank )

As visible,  the CO2  seems to be correlated with the humidity; therefore, a comparison between CO2 and humidity was done. After comparing them, the plot looks as follows.

Figure 17 (CO2)

It seems these two parameters have a linear correlation; therefore a humidity correction needs to be done. The CO2 values will be readjusted to zero humidity using the slope of 2.6 ppm/%humidity (Figure17).

Figure 18 (Results from the test tank (humidity corrected))

- Harbor: The module was put in the  Moss Landing, CA harbor for 24 hours. The data  collected was first humidity corrected as it was done on the test tank results, and afterwards, it  was analyzed in the three following ways.

Figure 19 (Results from the Moss Landing harbor)

TUTORIAL

In this tutorial, it will be demonstrated how to build your own CO2  module and how to make it work underwater.  All the steps are shown, from building the module, how to calibrate the sensor if it goes out of calibration, how to build the water tight housing, and how to hookup all the tubing necessary to measure CO2 .

- Assembling the Jeenode: As described before, the Jeenode has different ports for connecting different devices. Solder the following devices in the  following ports: RTC plug in Port 1, Relay plug in Port 2, Uart plug in Port 3, and Memory plug in Port 4 (remember  to connect the right pins with the right holes of the devices,  just make sure the names of the holes are the same as the names on the Jeenode).  Remember to connect the communication antenna (simply a wire) on the hole that says "A" besides the RF12 board

- Assembling the circuitry together:  The final circuit is  divided into two sub circuits.

- Power lines: first, make sure  that all the devices in the circuit are getting  the voltage that they need to operate. A 7805 voltage regulator will be used, in order to provide  the 5 volts  that the Jeenode, the Telaire CO2  sensor and the SHT21 need to properly function. In addition, add an electrolytic capacitor of 100 uF from the power of the Jeenode to ground, this will help the circuit to properly function when the batteries get a little low.  The power line circuit looks as follows.  Before assembling this circuit please read the section about "battery pack" where more details will be explained about how to use the 7805 5V regulator

Figure 20 (Block diagram of the circuit for the power lines)

- Communication lines:  The communication lines are the lines that are made so that the microcontroller (the Jeenode) can communicate with the Telaire sensor and the SHT21. The Telaire sensor uses  UART  protocol in order to communicate with the microcontroller; therefore, it has to have an TX and RX connection with the microcontroller. The  SHT21 uses I2C protocol in order to communicate with the microcontroller; therefore, it has to have an SDA and SCL connection with the microcontroller. The communication line is as follows.

Figure 21( Block diagram of the circuit for the communication lines)

When everything is assembled, the end result should be something looking as follows:

Figure 22 (Actual CO2  sensor module)

- Battery pack:  Since  the module is using a 7805 5 Volts regulator, a voltage greater than 7.5V and less than 20V has to be used in order to have a good regulation.  The supply voltage will be 9V to allow the circuit to keep working even when the batteries start getting discharged.  This will give us a 1.5V offset of discharge until the circuit stops working.  The 7805 requires special circuitry in order to work, it just needs two capacitors. There  are different  circuits that can be used to do a 5V regulation with the 7805 regulator; nevertheless, this was the circuit used for this project:

Figure 23( 7805 5V regulation circuit)

- Housing:  The CO2  sensing module and battery pack are designed to go inside A PVC tube. The sensing module is attached to a PVC endcap through screw  holes so it could be stable, and  it does not bend too much  (If bended too much , the sensing module board could break). There are also two barb fittings that are used as entrance and exit of the tubing system for the CO2  module.

The outside of the endcap has a copper nickel tube in which Teflon AF tubing can go inside. The cooper nickel tube is going to go attached with the endcap with two swage locks . The last component on the outside are two compression fittings through which  the Teflon AF  is going to go in and out of the system.

Figure 24 (Endcap design)

Figure 25 (Assembled endcap)

Figure 26( PVC tubing  and PVC pipe union)

- Tubing layout: From Figure 21, it is visible that the Teflon AF tubing is going to  go inside  of the Cooper nickel tubing, get out of it and go into the compression fittings.  It is recommendable that at least two loops are made inside of the cooper nickel tubing, the more surface area there is of tubing, then the faster the equilibration process.  For the inside part of the endcap, is no longer necessary to use more Teflon AF since no equilibration is going to take place inside the housing.  After the barb Fittings, it is recommendable to use Tygon tubing. Do not forget about using the Nafion  tubing before going into the sensor, as water vapor could interfere with CO2 measurements. Silica gel dryers packs must be placed inside the pressure housing to reduce humidity on the outside of the Nafion to dry the gas inside.

9

Figure 27 (Tubing layout)

- Code:   The code is made so it can be as user friendly as possible. It is subdivided into three main codes, one for the transmitter, one for the CO2  sensing module, and one for the calibration and sensor functionality. The complete code of the transmitter and the CO2  sensing module can be found in appendix A and B respectively. For  looking at the code of the sensor calibration and functionality, please visit appendix C.

- CO2  sensing module: This code is composed of mainly three functions: read_Save(), load_Memory (char buf) , erase_Memory, and rf12_Interrupt(). The read_Save() part of the code, as the name says, it just reads data and saves it into  the memory. It works using  cycles predefined by the user, where the user can specified the pumping rate and the amount of minutes between each measurement (notice that before a measurement  can be done, the sensor has to warm up at least 10 minutes, that timing system is already included in the code).  If the user wants to control the timing of the pump and sensor the following lines have to be changed

Figure 28 (Code example of cycle timing)

- Transmitter: The transmitter could be just another Jeenode or a Jeelink (same as a Jeenode with the only difference that is more compact and no plugs can be attached to it).  The transmitter is set up so when the user enters a "1" through the serial port, the Sensing module will start to send the data loaded into memory to the transmitter, and then resume normal operation. If the user enters a "2", the sensing module will erase the data inside the memory.

- Sensor calibration and functionality: The sensor calibration and functionality code is just used when the sensor wants to be calibrated, turn the ABC logic off, or test the functionality.  In order to tell the sensor to do either calibration or turn off the ABC logic, some lines have to be commented or uncommented; this lines are shown in Figure 25. If the sensor has to be calibrated, the gas concentration has to be introduced. In the code the setup is to do it at 200 ppm. nevertheless, if the concentration has to be changed, the last two lines of the command shown in Figure 26 have to be changed into the number that was to be input in hexadecimal notation.

Figure 29

Figure 30

Example: The setup for the calibration is 200 ppm which is represented as 00C8 in hexadecimal; therefore the last two lines of Figure 26 are

uart.wrtie(0x00);

uart.write(0xC8);

If the number wants to be changed two 400, the hexadecimal representation would be: 0190; therefore, the last two lines of code in Figure 26 would be as follows:

uart.wrtie(0x01);

uart.write(0x90);

Note: After the calibration process is done, weird data will start appearing on the serial monitor, please wait one minute after the calibration process, and simply turn off the sensor, next time it is turned on, it should be calibrated.

- User interface: A LabView user interface has been created in order to send the radio transmission commands as easy as possible. The LabView interface comes with real time charts that will allow the visualization of data in real time. The data is also  imported to an Excel file if wanted. If the user interface wants to be used, all what it has to be done is open the executable file named CO2 _sensing_module.exe. Nevertheless, sending the serial commands through the Arduino environment is also a good solution.

Conclusion

In this paper  it was proved that a reliable, low power, low cost  underwater  CO2  sensing module could be built in an easy way.  All the parts can be easily obtained online at  not  with an elevated price, if compared to the professional CO2  sensing modules that are already on the sea.  In order to assemble the module some technical skills are needed. Nevertheless, the tutorial presented on this paper is made to be as explicative as possible.

The results from this project showed that Teflon AF is an excellent CO2  equilibrator on air and underwater. How fast equilibrates is going to be proportional to the flow of air that has through it, how long the tubing is, and how big the diameter is. For underwater research, it has to be evaluated and tested the amount of pressure it can withstand , in order to know how deep it could be taken.  In addition a better algorithm and hardware could be developed in order to make the sensing module as low power consuming as possible.

To conclude, it has to be said that this project  serves as the base for a better platform that will allow better  readings of CO2  with a longer battery life period. More research has to be done in order to find a better sensor that  can detect CO2  with more precision and less consumption of power. However, the project presented on this paper fulfills the objective of being easy to build so it can be used for education and little research projects.

REFERENCES

ATABEK, SINAN. CO2 sensors. n.d. document. 1 August 20123.

Biogeneral. Teflon® AF Tubing and Membrane. n.d. BIOGENERAL. Web page. 2 August 13. <http://www.biogeneral.com/teflon.html (picture and table of Teflon AF tubing)>.

Damage to Marine Ecosystems as CO2 Emissions Rise. n.d. Save the seas Foundation. web page. 1 August 2013. <http://saveourseas.com/articles/damage_to_marine_ecosystems_as_co2_emissions_rise>.

Dasgupta, Purnendu. High-Sensitivity GAs Sensor Based on Gas-Permeable Liquid Core Waveguides And Long-Path Absorbance Detection. Lubbock, 1998. Document.

GE. Telaire 6615. n.d. GE, Measurment & Control. 1 August 2013. <http://www.ge-mcs.com/download/co2-flow/920-474D-LR.pdf>.

Jeelabs. Jeenode SMD. n.d. Jeelbas. Web page. 1 August 2013. <http://jeelabs.net/projects/hardware/wiki/JeeNode_SMD>.

—. Memory plug. n.d. Web page. 1 August 2013. <http://jeelabs.net/projects/hardware/wiki/Memory_Plug>.

—. Relay plug. n.d. Web page. 1 August 2013. <http://jeelabs.net/projects/hardware/wiki/Relay_Plug>.

—. RTC plug. n.d. Web page. 1 August 2013. <http://jeelabs.net/projects/hardware/wiki/RTC_Plug>.

—. UART plug. n.d. Web page. 1 August 2013. <http://jeelabs.net/projects/hardware/wiki/UART_Plug>.

KNF. Micro Diaphragm Gas Sampling Pump. 2013. THOMAS. Web Page. 5 August 2013. <http://www.gd-thomas.com/product.aspx?id=13052&tp=p>.

Mitchell, Alan. Alan Mitchell Building Energy Use Monitoring and Analysis. n.d. 1 August 2013. <http://alanbmitchell.wordpress.com/2011/10/02/operate-arduino-for-year-from-batteries/>.

NASA. Effects of changing the Carbon Cycle. n.d. Web. 1 August 2013. <http://earthobservatory.nasa.gov/Features/CarbonCycle/page5.php>.

Ning, Liu, et al. Development of a CO2 chemical sensor and sensor strategy for CO2 monitoring in carbon sequestration. n.d. Web document. 1 August 2013. <http://baervan.nmt.edu/research_groups/carbon_sequestration_membrane_technology/pages/Publications/B-1.pdf>.

Pure, Perma. Perma Pure FAQs. n.d. Perma Pure LLC. Web page. 2 August 2013. <http://www.permapure.com/support/faqs/.>.

Sensirion. SHT21 - Digital Humidity Sensor (RH&T). n.d. Sensirion. Web page. 2 August 2013. <http://www.sensirion.com/en/products/humidity-temperature/humidity-sensor-sht21/>.

Wikipedia. Beer_Lambert Law. n.d. Web document. 5 8 2013. <http://en.wikipedia.org/wiki/Beer%E2%80%93Lambert_law>.

—. I2C. n.d. Web document. 05 August 2013. <http://en.wikipedia.org/wiki/I%C2%B2C>.

—. Universal asynchronous receiver/transmitter. n.d. Web document. 5 8 2013. <http://en.wikipedia.org/wiki/Universal_asynchronous_receiver/transmitter>.

APPENDIX

- Appendix A

#include <JeeLib.h>

// set up the timer for the radio communications

MilliTimer sendTimer;

//global variables

byte needToSend;

void setup()

{

Serial.begin(9600);

rf12_initialize(1,RF12_915MHZ,1);

}

void loop()

{

send_Command();

}

/******************************************************************

* Function name: send_Command()

* type: void

* input parameters: NULL

* output parameter: NULL

* Description: Sends a command, wait for response of the node, and then

* make an action according to the command receicved. If a 1

* is sent, loading of the memory will take place. If a 2 is

* sent, deletion of the memory will take place.

* Creators: MBARI

********************************************************************/

void send_Command()

{

boolean out=true;

boolean receiveCommand=true;

uint8_t buf[1];

char reading[1];

uint8_t data[12];

while(receiveCommand)

{

if(Serial.available())

{

reading[0]=Serial.read();

if(reading[0]=='1'||reading[0]=='2')

{

reading[0]-=48;

receiveCommand=false;

}

}

}

while(out)

{

if (rf12_recvDone() && rf12_crc == 0) // received good data if true

{

buf[0]=rf12_data[0];

if(buf[0]==1)

{

while(1)

{

if (rf12_recvDone() && rf12_crc == 0) // received good data if true

{

for (byte i = 0; i < rf12_len; ++i)

{

data[i]=rf12_data[i];

}

if(data[0]==50)

{

Serial.println("memory empty");

}

else

{

transform(data);

}

}

}

}

else if(buf[0]==2)

{

Serial.println("memory erased");

}

out=false;

}

else

{

if (sendTimer.poll(10))

{

needToSend = 1;

}

if (needToSend && rf12_canSend())

{

rf12_sendStart(2,&reading,1);

needToSend = 0;

}

}

}

}

/******************************************************************

* Function name: transform(char [])

* type: void

* input parameters: uint8_t a[]: array of n bytes which will be loaded

* output parameter: NULL

* Description: Transform an arry of 1 byte integers into something that looks

* like mm//dd//yy h:m:s CO2 ppm= xxx temperature= xx.xx

* humidity= xx.x

* Creators: MBARI

********************************************************************/

void transform( uint8_t a[])

{

int buffer[11];

char string[80];

for(int i=0;i<6;i++) //get the time form the clock

{

buffer[i]= (int)a[i];

}

buffer[6]=((a[6] &0xff)<<8) | (a[7] &0xff); //merge the two bytes of co2

for(int i=8;i<12;i++) // get the values for temperature and humidity

{

buffer[i-1]=int(a[i]);

}

sprintf(string,"%02d/%02d/%02d %02d:%02d:%02d  CO2 ppm= %02d temperature= %02d.%02d humidity= %02d.%02d",

buffer[0],buffer[1],buffer[2],buffer[3],buffer[4],buffer[5],buffer[6],buffer[7],buffer[8],buffer[9],buffer[10]);

Serial.print(string);

Serial.println();

delay(1);

}

- Appendix B

#include <Wire.h>

#include <JeeLib.h>

#include <LibHumidity.h>

// sets up the clock port and the clock module

PortI2C myport (1 /*, PortI2C::KHZ400 */);

DeviceI2C rtc (myport, 0x68);

//sets up the port of the relay

Port relays (2);

//sets up the port of the uart and the uart module

PortI2C i2cBusUart (3);

UartPlug uart (i2cBusUart, 0x4D);

//sets up the port of the memory and the memory module

PortI2C i2cBusMem (4);

MemoryPlug mem (i2cBusMem);

MemoryStream stream (mem);

//sets up the Sensirion sensor for temperature and humidity

LibHumidity humidity_Temp = LibHumidity(0);

MilliTimer sendTimer; //setup the time for the radio communications

// must be defined in case we're using the watchdog for low-power waiting

ISR(WDT_vect) {

Sleepy::watchdogEvent();

}

//global variables

char bufData[12];//buffer for the data is going to go into the memory

void setup()

{

Serial.begin(9600);

uart.begin(19200);

// initialize the radio for this node

rf12_initialize(2, RF12_915MHZ, 1);

//initialize the relays

relays.digiWrite(0);

relays.mode(OUTPUT);

relays.digiWrite2(0);

relays.mode2(OUTPUT);

//uncomment the next two lines when you wat to use the memory for the first time

//this will start the directory to the page 0 adress 0

//after compiling and running it for 2 seconds comment them again and compile the code again

//int buf[2]={0,0};

//mem.save(0,0,buf,4);

//uncomment he following line to set up the date and time for the clock for the first time

//run the code for the 1 seconds comment the following line again and compile

//setDate (13,8,9,10,03,0);

/*while(1)

{

if(Serial.read()=='1')

{

char buf[12];

load_Memory(buf);

}

}*/

delay(10000);

Serial.end();

}

void loop()

{

read_Save();

}

/******************************************************************

* Function name: read_Save(void)

* type: void

* input parameters: void

* output parameter: NULL

* Description: function that records a Co2 value from the telaire sensor

* and stores it into memory

* Creators: MBARI

********************************************************************/

void read_Save()

{

boolean cycle=true;

int co2Value=0;

int pumpingCounter=3; // number of minutes until the pump turn on

int pumpRunTime=5000; //number of miliseconds for which the pump will run

int timeForMeasurements=20;

boolean idle=true; //number of minutes it will pass before the reading happens

// the sensor will turn on 10 minutes before that in order to warm up

int counter=0;

byte time[6];

int i=0;

//send the Telair to sleep

relays.digiWrite2(1);//turn on the sensor

delay(10000);

Telaire_IdleMode(idle);

//clearing the memory data array

for(int j=0;j<12;j++)

{

bufData[j]=NULL;

j++;

}

while(cycle)

{

//check for communication

rf12_sleep(-1);

rf12_recvDone();

receive_Command();

rf12_sleep(1);

rf12_sleep(0);

//sleep for n number of seconds if the pump does not turn on during the iteration

if(pumpingCounter!=3)

{

Sleepy::loseSomeTime(pumpRunTime);

}

// turn the pump on for n number of seconds

else if(pumpingCounter==3)

{

relays.digiWrite(1);

Sleepy::loseSomeTime(pumpRunTime);

relays.digiWrite(0);

pumpingCounter=0;

}

Sleepy::loseSomeTime((word)(60000-pumpRunTime));

counter++;

pumpingCounter++;

if(counter==timeForMeasurements-10)

{

Telaire_IdleMode(!idle);

}

if(counter==timeForMeasurements)

{

Serial.begin(9600);

delay(20);

co2Value=read_Co2();

while((co2Value>2000||co2Value<=0)&&i<=10)

{

co2Value=read_Co2();

if(i==10)

{

co2Value=1;

}

i++;

delay(10000);

}

float temp=get_Temperature();

delay(20);

float humidity=get_Humidity();

//getting the date and time in the buffer for the memory

getDate(time);

bufData[0]=time[1];

bufData[1]=time[2];

bufData[2]=time[0];

bufData[3]=time[3];

bufData[4]=time[4];

bufData[5]=time[5];

delay(20);

Serial.print(time[1]);

Serial.print("/");

Serial.print(time[2]);

Serial.print("/");

Serial.print(time[0]);

Serial.print(" ");

Serial.print(time[3]);

Serial.print(":");

Serial.print(time[4]);

Serial.print(":");

Serial.print(time[5]);

Serial.print(" ");

Serial.print(co2Value);

Serial.print(" ");

Serial.print(temp);

Serial.print(" ");

Serial.print(humidity);

Serial.println();

saved(bufData);

delay(10);

cycle=false;

counter=0;

Serial.end();

}

}

}

/******************************************************************

* Function name: load_Memory(char [])

* type: void

* input parameters: char a[]: array of n bytes which will be loaded

* output parameter: NULL

* Description: Loads the the information that is on memory and transmit

* it thorugh radio communications to a base station.

* Creators: MBARI

********************************************************************/

void load_Memory(char a[]) //function for loading characters in a fast way

{

int finish[2];

int remember[2]={

1,0    };

if(mem.isPresent())

{

mem.load(0,0,finish,4);

delay(20);

if(finish[0]!=0 && finish[1]!=0)

{

while((remember[0]!=finish[0])||(remember[1]!=finish[1]))

{

mem.load(remember[0],remember[1],a,12);

delay(50);

send_Radio(a,12);

delay(1);

transform(a);

if(remember[1]<243)

{

remember[1]+=12;

}

if(remember[1]>243)

{

remember[0]+=1;

remember[1]=0;

}

}

}

else

{

Serial.println("memory empty");

uint8_t checkDeletion[1]={50};

rf12_sendStart(1,checkDeletion,1);

}

}

}

/******************************************************************

* Function name: erase_Memory()

* type: void

* input parameters: void

* output parameter: NULL

* Description: Erases the memory.resets the directory making the memory

* to start overriding previous values every time it saves

* new data

* Creators: MBARI

********************************************************************/

void erase_Memory()

{

int erase[2]={0,0};

if(mem.isPresent())

{

mem.save(0,0,erase,4);

Serial.println("memory erased");

}

}

/******************************************************************

* Function name: receive_Command()

* type: void

* input parameters: NIULL

* output parameter: NULL

* Description: Checks if there is any radio communication command

* from the basen station and takes and action according

* to the command

* Creators: MBARI

********************************************************************/

void receive_Command()

{

uint8_t buf[1];

char buffer[12];

byte InfoReceived=0;

byte needToSend=0;

int i=0;

boolean out=true;

Serial.begin(9600);

delay(20);

while(i<50 && out)//i<2000 && out

{

i++;

if (rf12_recvDone() && rf12_crc == 0) // received good data if true

{

buf[0]=rf12_data[0];

InfoReceived=1;

needToSend=1;

}

if(needToSend && rf12_canSend() && InfoReceived)

{

rf12_sendStart(1, buf,1);

InfoReceived=0;

needToSend = 0;

out=false;

if(buf[0]==1)

{

load_Memory(buffer);

}

if(buf[0]==2)

{

erase_Memory();

}

}

}

Serial.end();

delay(20);

}

/******************************************************************

* Function name: read_Co2(void)

* type: int

* input parameters: void

* output parameter: int

* Description: tells tehe sensor to start doing measurments and store

* the two bytes in a varible called co2value

* Creators: MBARI

********************************************************************/

int read_Co2()

{

byte buf[56];

byte i=0;

int co2Value;

uart.write(0xFF);

uart.write(0xFE);

uart.write(0x02);

uart.write(0x02);

uart.write(0x03);

delay(20);

// put he two bit answer from the sensor into a buffer

while(uart.available())

{

buf[i]=uart.read();

i++;

}

//save the two bits of Co2 into memory buffer

bufData[6]=buf[3];

bufData[7]=buf[4];

// do a bitwise operation to acomodate the two bits

// and make them an integer

co2Value=0;

co2Value|=buf[3] & 0xFF;

co2Value=co2Value<<8;

co2Value|=buf[4] & 0xFF;

return co2Value;

}

/******************************************************************

* Function name: get_Temp(void)

* type: float

* input parameters: void

* output parameter: float

* Description: tells the Sensirion sensor to start doing measurements

* of temperature, after that it divides the integer and

* float part of the number in two bytes so they can be

* stored into memory. The function uses the library

* LibHumidity from www.moderndevices.com in order to control

* the sensor

* Creators: MBARI

********************************************************************/

float get_Temperature()

{

float temp;

byte tempInt;

byte tempFloat;

/* get the measurement of temperature and divides them into an

integer part and a float part*/

temp=humidity_Temp.GetTemperatureC();

tempInt=floor(temp);

tempFloat=(temp*100)-(tempInt*100);

//get the date into memroy

bufData[8]=tempInt;

bufData[9]=tempFloat;

return temp;

}

/******************************************************************

* Function name: get_Humidity(void)

* type: float

* input parameters: void

* output parameter: float

* Description: tells the Sensirion sensor to start doing measurements

* of humidity, after that it divides the integer and

* float part of the number in two bytes so they can be

* stored into memory. The function uses the library

* LibHumidity from www.moderndevices.com in order to control

* the sensor

* Creators: MBARI

********************************************************************/

float get_Humidity()

{

float humidity;

byte humidityInt;

byte humidityFloat;

/* get the measurement of humidity and divides them into an

integer part and a float part*/

humidity=humidity_Temp.GetHumidity();

humidityInt=floor(humidity);

humidityFloat=(humidity*100)-(humidityInt*100);

//get the date into memroy

bufData[10]=humidityInt;

bufData[11]=humidityFloat;

return humidity;

}

/****************************************

* Function name: bind2bcd(byte)

* type: byte

* input parameters: byte val

* output parameter:  val + 6 * (val / 10);

* Description:

* Creators: Jeelabs.

*****************************************

*/

static byte bin2bcd (byte val)

{

return val + 6 * (val / 10);

delay(1);

}

/****************************************

* Function name: bcd2bcd(byte)

* type: byte

* input parameters: byte val

* output parameter: val - 6 * (val >> 4);

* Description:

* Creators: Jeelabs

*****************************************

*/

static byte bcd2bin (byte val)

{

return val - 6 * (val >> 4);

delay(1);

}

/***************************************************************************************************

* // Function name: setDate(byte,  byte, byte, byte, byte)

* type: static void

* input parameters: byte yy: year

* byte mm:month

* byte dd: day

* byte h: hour

* byte m: minutes

* byte s: seconds

* output parameter: NULL

* Description:  This function is used in order to setup a specific day in the format yy/mm/dd h:m:s

* Compile it once in order to set up the clock. Erase the call of this function from the code.

* Finally compile again in order to make the clock to start running indepedently.

* Creators: Jeelabs

********************************************************************************************************

*/

static void setDate (byte yy, byte mm, byte dd, byte h, byte m, byte s)

{

rtc.send();

rtc.write(0);

rtc.write(bin2bcd(s));

rtc.write(bin2bcd(m));

rtc.write(bin2bcd(h));

rtc.write(bin2bcd(0));

rtc.write(bin2bcd(dd));

rtc.write(bin2bcd(mm));

rtc.write(bin2bcd(yy));

rtc.write(0);

rtc.stop();

}

/******************************************************************

* Function name: getDate (void)

* type: static void

* input parameters: bye* buf: byte array in which the date and time would be introduced

* output parameter: NULL

* Description: get the date and time from the clock and puts it into

* the byte array specified

* Creators: MBARI

********************************************************************/

static void getDate (byte* buf)

{

rtc.send();

rtc.write(0);

rtc.stop();

rtc.receive();

buf[5] = bcd2bin(rtc.read(0));

buf[4] = bcd2bin(rtc.read(0));

buf[3] = bcd2bin(rtc.read(0));

rtc.read(0);

buf[2] = bcd2bin(rtc.read(0));

buf[1] = bcd2bin(rtc.read(0));

buf[0] = bcd2bin(rtc.read(1));

rtc.stop();

}

/******************************************************************

* Function name:saved(char[])

* type: void

* input parameters: char a[]= array whcih will be saved into memory

* output parameter: NULL

* Description: saves the input array into the memory. The way it works is

* by checking the directory of the memory which is at location

* [0,0] if there is nothing in memory start saving on adress [1,0]

* if there is something on memory start saving where you left off

* Creators: MBARI

********************************************************************/

void saved(char a[])

{

int remember[2];

if(mem.isPresent())

{

mem.load(0,0,remember,4);

delay(10);

if(remember[0]==0 && remember[1]==0)

{

mem.save(1,0, a,12);

delay(10);

remember[0]=1;

remember[1]+=12;

mem.save(0,0,remember,4);

delay(10);

}

else

{

if(remember[1]<243)

{

mem.save(remember[0],remember[1],a,12);

delay(10);

remember[1]+=12;

}

if (remember[1]>243)

{

remember[0]+=1;

remember[1]=0;

mem.save(remember[0],remember[1],a,12);

remember[1]+=12;

}

mem.save(0,0,remember,4);

delay(10);

}

}

}

/******************************************************************

* Function name: send_Radio( char a[], int length)

* type: void

* input parameters: char a[]: array which will be transmitted

* int length: length of the array to be transmitted

* output parameter: NULL

* Description: sends an array of n length through radio communications

* to a base station

* Creators: MBARI

********************************************************************/

void send_Radio(char a[], int length)

{

delay(10);

while (!rf12_canSend()) // wait until sending is allowed

{

rf12_recvDone();    // processes any incoming data, it all happens "in the background"

}

rf12_sendStart(1,a, length);  // sends the data out

}

/******************************************************************

* Function name: transform(char [])

* type: void

* input parameters: char a[]: array of n bytes which will be loaded

* output parameter: NULL

* Description: Transform an arry of characters into something that looks

* like mm//dd//yy h:m:s CO2 ppm= xxx temperature= xx.xx

* humidity= xx.x

* Creators: MBARI

********************************************************************/

void transform( char a[])

{

int buffer[11];

char string[80];

for(int i=0;i<6;i++) //get the time form the clock

{

buffer[i]= (int)a[i];

}

buffer[6]=((a[6] &0xff)<<8) | (a[7] &0xff); //merge the two bytes of co2

for(int i=8;i<12;i++) // get the values for temperature and humidity

{

buffer[i-1]=int(a[i]);

}

sprintf(string,"%02d/%02d/%02d %02d:%02d:%02d  CO2 ppm= %02d temperature= %02d.%02d humidity= %02d.%02d",

buffer[0],buffer[1],buffer[2],buffer[3],buffer[4],buffer[5],buffer[6],buffer[7],buffer[8],buffer[9],buffer[10]);

Serial.print(string);

Serial.println();

delay(1);

}

/******************************************************************

* Function name: Telaire-IdleMode(boolean mode)

* type: void

* input parameters: boolean mode,

* output parameter: NULL

* Description: Indicates wheter the sensor is

going to go in Idle mode or not. True=Idle mode,

False=wake up

* Creators: MBARI

********************************************************************/

void Telaire_IdleMode(boolean mode)

{

boolean out=true;

byte buf[3];

//if true the Telaire will go into sleep mode

if(mode==true)

{

while(out)

{

int i=0;

uart.write(0xFF);

uart.write(0xFE);

uart.write(0x02);

uart.write (0xB9);

uart.write (0X01);

while(uart.available())

{

buf[i]=uart.read();

i++;

}

if((buf[0]==255)&&(buf[1]==250)&&(buf[2]==0))

{

out=false;

}

}

}

//if not true, the Telair will wake up

if(mode==false)

{

while(out)

{

int i=0;

uart.write(0xFF);

uart.write(0xFE);

uart.write(0x02);

uart.write (0xB9);

uart.write (0X02);

while(uart.available())

{

buf[i]=uart.read();

i++;

}

if((buf[0]==255)&&(buf[1]==250)&&(buf[2]==0))

{

out=false;

}

}

}

}

- Appendix C

//uncomment the following line to calibrate the sensor

//Calibrate_Telair();

}

void loop()

{

//read CO2 every 10 seconds

int co2=read_Co2();

Serial.print("CO2 ppm=");

Serial.println(co2);

delay(10000);

}

/******************************************************************

Function name: read_Co2(void)

type: int

input parameters: void

output parameter: int

Description: tells tehe sensor to start doing measurments and store

the two bytes in a varible called co2value

Creators: MBARI

********************************************************************/

int read_Co2()

{

byte buf[56];

byte i=0;

int co2Value;

uart.write(0xFF);

uart.write(0xFE);

uart.write(0x02);

uart.write(0x02);

uart.write(0x03);

delay(20);

// put he two bit answer from the sensor into a buffer

while(uart.available())

{

buf[i]=uart.read();

i++;

}

// do a bitwise operation to acomodate the two bits

// and make them an integer

co2Value=0;

co2Value|=buf[3] & 0xFF;

co2Value=co2Value<<8;

co2Value|=buf[4] & 0xFF;

return co2Value;

}

/******************************************************************

Function name: Abc_Logic_Off()

type: void

input parameters: NULL

output parameter: NULL

Description:Tells the sensor to turn off the ABC logic (self calibration

process)

Creators: MBARI

********************************************************************/

void Abc_Logic_Off()

{

byte buf[10]={};

boolean out =true;

while(out)

{

int i=0;

uart.write(0xFF);

uart.write(0xFE);

uart.write(0x02);

uart.write(0xB7);

uart.write(0x02);

delay(20);

// put he two bit answer from the sensor into a buffer

while(uart.available())

{

uart.read();

}

delay(20);

//check if the abc logic was turned of

uart.write(0xFF);

uart.write(0xFE);

uart.write(0x02);

uart.write(0xB7);

uart.write(0x00);

delay(20);

// put he two bit answer from the sensor into a buffer

while(uart.available())

{

buf[i]=uart.read();

i++;

}

Serial.println(buf[3]);

if(buf[3]==0x01)

{

Serial.println("ABC logic on");

}

if(buf[3]==0x2)

{

Serial.println("ABC loigc turned off");

out=false;

}

else if((buf[3]<1||buf[3]>2))

{

Serial.println("error reading ABC logic status");

}

}

}

/******************************************************************

Function name: Check_Abc_Logic()

type: void

input parameters: NULL

output parameter: NULL

Description:Tells the sensor chek for the ABC logic status

Creators: MBARI

****************