## Sheet1
Calibration Checks | 2023may 22-25
For calibration see:
Apendixies of 500-4435 Short Puck Conversion Process_Modified_05.21.2023
2GESP_Calibration_CH_PHv1-espXXX-date.xlsx
Calibration_SE.doc
| espRUSH | espNiagara | espNessie
Sleepy |  | ESP2 v3.42 I2C Gateway v3.20 for core bus | ESP2 v3.42 I2C Gateway v3.20 for core bus
Dwarfs |  | All dwarf controllers running firmware version 3.85 | All dwarf controllers running firmware version 3.85
Clamp o-rings changed (PC, CC) | Yes | Yes | Yes
Syringes (SS, CS, PS) | unchanged, not checked | unchanged, not checked | unchanged, not checked
Manipulator: GARAGE | Updated calibration | Updated calibration
Manipulator: CC | Updated calibration | Updated calibration
Manipulator: PC | Updated calibration | Updated calibration
Manipulator: Storage Carousel (SC) - positions 1..7 | Updated calibration | Updated calibration
Elevator | verified, no change | Updated (sePuck, seBottom, SEbeamHeight, seTop, place threshold
Clamp
CC | updated calibration | verified, no change | verified, no change
PC | verified no change | verified, no change | verified, no change
Temperature Sensors
CH | New calibration (room temperature) | New calibration (room temperature)
PH | New calibration (room temperature) | New calibration (room temperature)
Pressure Sensors
CS | New calibration | verified, no change
SS | New calibration | verified, no change
Hold Pressure - puck in clamp?
CC | Yes | Yes
PC | Yes | Yes
Light Puck Test (USED LIGHT PUCK 0) | Yes -Light puck 0
Run: require 'shuffle'; shuffle 2 #two pucks in SC1 | Yes | Yes
Current state: | Running Fluidics: shortDA, some fluid retention | PSR reagent burkert valve stuck open at sub1. -- needs replacement | Manipulator motor is bad. Being sent to McLane for rebuild
| Array images improved after rinsing out PBV.waste1 line with lots of warm water | PSR sub2 reagent burkert valve clogged -- needs replacement
Reagents on CSR need priming
Wash is fouled, replace
TO DO before June 19th | 1. 2x shortDA | 1. Replace bad burkert valves & update configuration | 1. Replace manipulator
| 2. CC.loadFlush {purgeUpper, purgeLower; CSR.purge}; PS.flush | 2. Replace wash | 2. Manipulator Calibrations - Garage, CC, PC
| 3. Run 3 peat phase | 3. Light puck test (can do now) | 3. Storage carousel positions
|  | 4. CC.loadFlush {purgeUpper, purgeLower; CSR.purge}; PS.flush | 4. Manipulator to Carousel tubes
|  | 5. 2x shortDA | 5. Heater calibration (CH and PH)
|  | 6. Run 3 peat phase | 6. Pressure calibration check (CS, SS)
7. Leak Check
8. Light puck
9. shuffle pucks
10. Rinse out PBV.waste1 with hot water
11. CC.loadFlush {purgeUpper, purgeLower; CSR.purge}; PS.flush
12. 3x shortDA
13. Run 3 peat phase