## NIAGARA - Manipulator
ESP NIAGARA - Manipulator Calibrations
Date: 5/23/23
MANIPULATOR TO GARAGE
Forearm | Elbow
-12512 | 11742
-12511 | 11743
-12510 | 11744
-12510.5 | 11743.5
MANIPULATOR TO CC
Forearm | Elbow
-12731 | 54236
-12733 | 54244
-12748 | 54242
-12725 | 54241
-- | 54240.75

## RUSH - Heater Cal
ESP Rush
Date: 5/24/23
NOTE: The thermal pads on ESP Rush's CC and PC are compressed, likely from the overextended forearm to clamp moves before calibration of the forearm. Because thermal pad on temp sensor finger was compressed on both clamps, forearm has to slightly deflect temp sensor finger for good thermal contact. This was done without deflecting the finger too far back and risking breaking the temp sensor finger.
** Values before forearm to clamp calibrations, post Preston's cal update | [5 empty] | ** Values obtained after Danna's heater calibrations and forearm to clamp calibrations
CH - Heater Check |  |  |  | PH- Heater Check |  | CH - Heater Check |  |  | PH - Heater Check
seconds | Thermometer | ESP | minutes | Thermometer | ESP | seconds | Thermometer | ESP | minutes | Thermometer | ESP
ambient | 21.4 | 21.4 | ambient | 21.8 | 22 | ambient | 22.1 | 21.9 | ambient | 21.6 | 22
0 | 56.1 | 68.1 |  |  | 26 | 0 | 66.3 | 68 | 0 | 23.9 | 26
30 | 66.3 | 70 |  |  | 32.2 | 30 | 70.1 | 70 | 1 | 30 | 31.8
60 | 66.9 | 70.6 |  | 32.4 | 32.4 | 60 | 71.5 | 70.1 | 2 | 32.1 | 32.6
90 | 68.2 | 70.6 |  |  | 32.5 | 90 | 71.6 | 69.5 | 3 | 32.2 | 32.2
120 | 68.8 | 70.2 |  | 31.5 | 31.6 | 120 | 70.9 | 68.9 | 4 | 31.8 | 31.4
150 | -- | 69.8 |  | 30.7 | 30.7 | 150 | 70.4 | 68.7 | 5 | 31.1 | 30.6
180 | -- | 69.7 |  | 30.3 | 29.9 | 180 | 70.4 | 69.1 | 6 | 30.3 | 29.8
210 | -- | 69.6 |  | 29.4 | 29.1 | 210 | 70.7 | 69.7 | 7 | 29.5 | 29.1
240 | -- | 69.7 |  | 28.8 | 28.4 | 240 | 71 | 70.1 | 8 | 28.8 | 28.4
270 | 68.8 | 70 |  |  |  | 270 | 71.3 | 70.2 | 9 | 28.2 | 27.9
300 | 68.8 | 70 |  |  |  | 300 | 71.3 | 70.2 | 10 | 28.1 | 28
330 | 68.9 | 70.1 |  |  |  | 330 | 71.4 | 70.1 | 11 | 28.2 | 28.2
360 | 68.9 | 70.1 |  |  |  | 360 | 71.2 | 70 | 12 | 28.2 | 28
390 | 68.9 | 70.1 |  |  |  | 390 | 71.1 | 70 | 13 | 28 | 27.9
420 | 68.9 | 70 |  |  |  | 420 | 71.1 | 70 | 14 | 28.1 | 28.1
450 | 68.9 | 70 |  |  |  | 450 | 71 | 69.9 | 15 | 28.1 | 28
480 | 68.9 | 70 |  |  |  | 480 | 71 | 70 | 16 | 28.1 | 28
510 | 68.9 | 70 |  |  |  | 510 | 71.1 | 70.1 | 17 | 28.1 | 28
540 | 68.9 | 70 |  |  |  | 540 | 71.1 | 70 | 18 | 28.1 | 27.9
570 | 68.5 | 68.8 |  |  |  | 570 | 71.1 | 70.1 | 19 | 28.1 | 28
600 | 67.6 | 66.6 |  |  |  | 600 | 71.1 | 70 | 20 | 28.1 | 28.1
630 | 66.3 | 64.3
660 | 64.7 | 62.1
690 | 62.7 | 60.1
720 | 61 | 58.2
750 | 59.1 | 56.3
780 | 57.4 | 54.6
810 | 52.8 | 52.8
**Values after forearm to clamp calibrations, possibly before updated config with Danna's calibration was pushed
CH - Heater Check
scconds | Thermometer | ESP
ambient | 21.6 | 21.9
0 | 66.4 | 68.1
30 | 70.2 | 70
60 | 71.2 | 70
90 | 70.5 | 69.2
120 | 70.3 | 68.7
150 | 70.4 | 68.9
180 | 70.9 | 69.4
210 | 70.8 | 69.8
240 | 71.2 | 70.1
270 | 71.5 | 70.2
300 | 71.8 | 70.1
330 | 71.4 | 70.1
360 | 71.3 | 69.9
390 | 71.3 | 69.9
420 | 71.3 | 69.9
450 | 71.3 | 70
480 | -- | --
510 | -- | --
540 | -- | --
570 | -- | --
600 | -- | --

## RUSH - Clamp
ESP RUSH - CC Calibration
Date: 5/24/23
NOTE: We could not visually see the Collection Clamp heater pad deflecting below the top plate. Top clamp was backed off until torque ratio and ON/OFF spring values were okay. CC will need to be removed, cleaned and recalibrated after deployment. Visually inspect wave spring and bottom clamp components.
**PC torque ratio and ON/OFF spring values checked by Preston and looked okay.
CC CALIBRATION | CC.config.maxCurrent = 350
closed = 24265
Torque Ratio | ON/OFF | max Current
2.22 | 11/28 | --
2.32 | 10/27 | --
2.39 | 10/27 | --
2.39 | 11/28 | 200
2.31 | 11/28 | 212
2.39 | 11/28 | 214
onguides = 24779
closed = 23765
Torque Ratio | ON/OFF | max Current
2.21 | 10/28 | 213
**Top clamp position moved up 500 counts, but heater pad was not compressing enough at closed position =23765

## RUSH - LeakCheck
CC & PC Leak Check
PC.loadPuck :SAF,2 | *Load SAF puck from tube 2
PSV.PSR | *Fill line with flush
PSR.flush
PS.fill
PSV.waste1
PS.empty
*Connect pressure gauge to the PBV
PSV.PSR
PSR.flush
PS.fill
PSV.PTV
PTV.puck
PBV.air
PS.push 1
PS.push 2.5 | *Incremented 1, 0.5, 0.5, 0.5 mL (2.5mL is max)
delay 10:00 | *check pressure drop, start = 9psi (pressure held)
CC.loadPuck :SH1,1 | *Load SH1 puck from tube 1
CSV.flush
CS.fill
CSV.CTV
CBV.DBV
CTV.CSV
CS.push 5
CS.push 0.5 | *release air, pressure gauge in-line
CS.push 1
CS.push 3
delay 10:00 | *check pressure drop, start = 56.2psi, end, 48psi (14.6% drop after 17 min)
Third Leak Test checked by Tina and Preston.

## RUSH - Pressure Cal
ESP RUSH - Pressure Sensor Calibration
s/n: |  | date: | 2023-05-25
[7 empty] | SS (outlet) | pre-cal |  |  |  |  | CS (intake) | pre-cal
| Absolute PSI | Offset | bias |  |  |  | -50.1931 | [5 empty] | -105.5616
| Absolute PSI | Slope | skew |  |  |  | 0.0561 | [5 empty] | 0.058388
| cnts@-15psi | Vacuum |  |  |  |  | 915.0743 | [5 empty] | 1837.8417
| mV/psi*75 | Full Scale |  |  |  |  | 437.863 | [5 empty] | 406.3128
| mV | Offset |  |  |  |  | 352.115 | [5 empty] | 646.185
| mV/psi | Slope |  |  |  |  | 5.8382 | [5 empty] | 5.4175
| (10V/3V) | Excitation |  |  |  | 3.3333 | -16.9922 |  |  |  |  | 3.3333 | -16.9922
| 100/1000(mV/V) | Gain |  |  | 0.1 |  | -5.0977 |  |  |  | 0.1 |  | -5.0977
| 10K | Offset |  | 1 |  |  | -0.50977 |  |  | 1 |  |  | -0.50977
| (4096/16/2.5V) | cnts/V | 102.4 |  |  |  | 2.4902 |  | 102.4 |  |  |  | 2.4902
Abs PSI | Gauge Target | Gauge PSI / inHg | SS (outlet, outPress) | (V) | (V) | (mV) |  |  | CS (intake, inPress) | (V) | (V) | (mV)
14.7 | 0 | 0 | 1172 | 11.4453 | 10.4453 | 104.4531 | 348.1771 |  | 2084 | 20.3516 | 19.3516 | 193.5156 | 645.0521
17.5 | 3 | 2.8 | 1246 | 12.168 | 11.168 | 111.6797 | 372.2656 |  | 2137 | 20.8691 | 19.8691 | 198.6914 | 662.3047
21.4 | 7 | 6.7 | 1309 | 12.7832 | 11.7832 | 117.832 | 392.7734 |  | 2201 | 21.4941 | 20.4941 | 204.9414 | 683.138
29.6 | 15 | 14.9 | 1447 | 14.1309 | 13.1309 | 131.3086 | 437.6953 |  | 2334 | 22.793 | 21.793 | 217.9297 | 726.4323
45.5 | 30 | 30.8 | 1714 | 16.7383 | 15.7383 | 157.3828 | 524.6094 |  | 2592 | 25.3125 | 24.3125 | 243.125 | 810.4167
68.2 | 50 | 53.5 | 2099 | 20.498 | 19.498 | 194.9805 | 649.9349 |  | 2965 | 28.9551 | 27.9551 | 279.5508 | 931.8359
92.7 | 75 | 78 | 2517 | 24.5801 | 23.5801 | 235.8008 | 786.0026 |  | 3370 | 32.9102 | 31.9102 | 319.1016 | 1063.6719
64.6 | 50 | 49.9 | 2035 | 19.873 | 18.873 | 188.7305 | 629.1016 |  | 2907 | 28.3887 | 27.3887 | 273.8867 | 912.9557
45.4 | 30 | 30.7 | 1710 | 16.6992 | 15.6992 | 156.9922 | 523.3073 |  | 2590 | 25.293 | 24.293 | 242.9297 | 809.7656
32.2 | 17 | 17.5 | 1486 | 14.5117 | 13.5117 | 135.1172 | 450.3906 |  | 2372 | 23.1641 | 22.1641 | 221.6406 | 738.8021
31.7 | 15 | 15.6 | 1451 | 14.1699 | 13.1699 | 131.6992 | 438.9974 |  | 2345 | 22.9004 | 21.9004 | 219.0039 | 730.013
22.1 | 7 | 7.4 | 1313 | 12.8223 | 11.8223 | 118.2227 | 394.0755 |  | 2209 | 21.5723 | 20.5723 | 205.7227 | 685.7422
18 | 3 | 3.3 | 1243 | 12.1387 | 11.1387 | 111.3867 | 371.2891 |  | 2143 | 20.9277 | 19.9277 | 199.2773 | 664.2578
14.7 | 0 | 0 | 1183 | 11.5527 | 10.5527 | 105.5273 | 351.7578 |  | 2087 | 20.3809 | 19.3809 | 193.8086 | 646.0286
12.2931 | -5 | -4.9 | 1100 | 10.7422 | 9.7422 | 97.4219 | 324.7396 |  | 2006 | 19.5898 | 18.5898 | 185.8984 | 619.6615
9.788 | -10 | -10 | 1012 | 9.8828 | 8.8828 | 88.8281 | 296.0938 |  | 1923 | 18.7793 | 17.7793 | 177.793 | 592.6432
7.8723 | -15 | -13.9 | 946 | 9.2383 | 8.2383 | 82.3828 | 274.6094 |  | 1860 | 18.1641 | 17.1641 | 171.6406 | 572.1354
14.7 | -22 |  |  | 0 | -1 | -10 | -33.3333 |  |  | 0 | -1 | -10 | -33.3333
14.7 | -27 |  |  | 0 | -1 | -10 | -33.3333 |  |  | 0 | -1 | -10 | -33.3333
Sampler.intakePSI |  | Pressure on the Collection side of the membrane, has calibration data applied
Sampler.outletPSI |  | Pressure on the Sample side of the membrane
Sampler.intakePress |  | Uncalibrated Press based on A/D values
Sampler.outletPress |  | Uncalibrated Press based on A/D values
Sampler.calibrate |  | Tells the machine to apply calibration to Press values
SS.status (inPress / outPress) |  | Raw A/D counts of outlet
CS.status (inPress / outPress) |  | Raw A/D counts of intake