Syringe Calibrations
· Sent all syringes to .home.to :home
· DO NOT do an ESP.ready! before calibrating syringes especially for PS. Motors are strong and can break parts.
· Full 25mL: Jog syringe up until bottom of syringe plunger lip is at the middle of top line. (PS.status.current - Holding current was at 20)
· Empty 0mL: Jog syringe down until bottom of syringe plunger lip is at bottom line. It is hard to see the bottom so check holding current. When current jumps up, back up 20 counts until holding current is back. (Eddie’s holding current was 4)
· Jogged in increments of 1000, 500, 200, 20
· SS.status  Take position counts and input in configure file
· Save old configuration to new file  cp configure.rb configurebackup.rb
· Continue for all syringes 25mL, 10mL, 2.5mL

