Dwarf rValve Board Setup and Testing

**Programming**

- Set power supply to 3.3V

- Connect XILINX DLC7 to computer’s parallel port and PS/2 port or USB-PS/2 adaptor.  //Light on XILINX should be yellow

- Connect central connector of 3-way split cable to rvalve board J2.

  - Connect BLK/WHT/RED twisted wires to XILINX

  - Connect banana plugs to 3.3V power and turn on power supply output.  //XILINX light should turn green

- Start iMPACT

  - Select “load most recent project” which should be rvalve.ipf

(if rvalve.ipf does not appear in window, navigate to: ProjectLibrary\ESP\2G\Project.Drawings\Project.EE.Dwgs\Programmable Logic\Rvalve\Rvalve\U4\implement\ver0\rev2 and select rvalve.ipf)

    - Make sure “xrc3384xl u4_2ea1.jed” appears.

  - Right-click icon and select Program

    - Check Erase Before Programming

    - Check Verify

    - Check XPLA and type u4_2ea1 in text box.

  - Click OK

  - Right-click and select “Get Device Checksum” //should see 2ea1 below

  - Right-click and select “Get XPLA Device UES”  //u4_2ea1 below

**Power up system with core board and rValve board**

Set dip switch on tested core board to “processing” address (on/off/on/off/off/off)

With power off, install rvalve board on top of core board.

Set power supply voltage to 13V +/- 1V

Connect power supply to Sleepy J2  //Board should draw nominal 0.2A – 0.3A

Attach DwarfBus ribbon cable from Sleepy J6/J8/J10 to Core J3

Attach power cable from J5 on Sleepy to J2 on core board

**Login**

telnet esptech

esp

genosensor

esp //make sure that there are no errors (30sec load time)

**setup**

require ‘testDwarf’ //True

testDwarf

**Testing**

(This testing done with Dwarf firmware v3.68, ESP2 v3.00 I2C Gateway 3.01)

- Solenoid test

  - Connect 20 pin connector from LED protoboard to J5 of rValve board

  - ProcessingValves[0].open  //1st (top) row LED should flash green on left

  - ProcessingValves[0].bypass  //1st (top) row LED should flash red on right

  - ProcessingValves[1].open  //2nd row LED should flash green on left

  - ProcessingValves[1].bypass  //2nd row LED should flash red on right

  - ProcessingValves[2].open  //3rd row LED should flash green on left

  - ProcessingValves[2].bypass  //3rd row LED should flash red on right

  - ProcessingValves[3].open  //4th row LED should flash green on left

  - ProcessingValves[3].bypass  //4th row LED should flash red on right

  - ProcessingValves[4].open  //5th row LED should flash green on left

  - ProcessingValves[4].bypass  //5th row LED should flash red on right

  - ProcessingValves[5].open  //6th row LED should flash green on left

  - ProcessingValves[5].bypass  //6th row LED should flash red on right

  - ProcessingValves[6].open  //7th row LED should flash green on left

  - ProcessingValves[6].bypass  //7th row LED should flash red on right

  - ProcessingValves[7].open  //8th (bot) row LED should flash green on left

  - ProcessingValves[7].bypass  //8th (bot) row LED should flash red on right

- Heater Control Testing

  - Connect 20 pin connector from protoboard’s bank of LEDs to J5 of rValve board

  - Measure voltage on red and black test points on protoboard  //red should be +5V relative to ground (black)

  - Connect DMM to yellow and black test points on protoboard (heater control signal).  Meter should read 0V.

  - PH.hold 700,100  //voltage should increase dramatically (3.29V)

  - PH.stop  //voltage should return to the previous value (0V)

- Rotary Valve Testing

  - Connect ‘bottom row’ connectors of motor/encoder test cable to motor and encoder

  - Connect other end of motor/encoder test cable to J1 of rValve board

  - PSV.advance 1; PSV.retard 1  //motor should spin one way, then the other

  - Remove connector from J1 of rValve board and connect to J3

  - PTV.advance 1; PTV.retard 1  //motor should spin one way, then the other

  - Connect ‘top row’ connectors of motor/encoder test cable to motor and encoder (leave other end connected to J3)

  - PRV.advance 1; PRV.retard 1  //motor should spin one way, then the other

  - Remove connector from J3 of rValve board and again connect it to J1

  - PBV.advance 1; PBV.retard 1  //motor should spin one way, then the other