Dwarf rValve Board Setup and Testing

**Programming**

- Set power supply to 3.3V

- Connect XILINX DLC7 to computer’s parallel port and PS/2 port or USB-PS/2 adaptor.  //Light on XILINX should be yellow

- Connect central connector of 3-way split cable to rvalve board J2.

  - Connect BLK/WHT/RED twisted wires to XILINX

  - Connect banana plugs to 3.3V power and turn on power supply output.  //XILINX light should turn green

- Start iMPACT

  - Select “load most recent project” which should be rvalve.ipf

(if rvalve.ipf does not appear in window, navigate to: ProjectLibrary\ESP\2G\Project.Drawings\Project.EE.Dwgs\Programmable Logic\Rvalve\Rvalve\U4\implement\ver0\rev2 and select rvalve.ipf)

    - Make sure “xrc3384xl u4_2ea1.jed” appears.

  - Right-click icon and select Program

    - Check Erase Before Programming

    - Check Verify

    - Check XPLA and type u4_2ea1 in text box.

  - Click OK

  - Right-click and select “Get Device Checksum” //should see 2ea1 below

  - Right-click and select “Get XPLA Device UES”  //u4_2ea1 below

**Power up system with core board and rValve board**

Set dip switch on tested core board to “processing” address (on/off/on/off/off/off)

With power off, install rvalve board on top of core board.

Set power supply voltage to 13V +/- 1V

Connect power supply to Sleepy J2  //Board should draw nominal 0.2A – 0.3A

Attach DwarfBus ribbon cable from Sleepy J6/J8/J10 to Core J3

Attach power cable from J5 on Sleepy to J2 on core board

**Login**

telnet esptech

everlove

techtransfer

esp //make sure that there are no errors (30sec load time)

**setup**

require ‘testDwarf’ //True

testDwarf

**Testing**

(This testing done with Dwarf firmware v3.65, ESP2 v2.92 I2C Gateway 2.00)

- Solenoid test

  - Connect 20 pin connector from protoboard’s bank of LEDs to J5 of rValve board

  - PSR.to :sig1  //top row’s LED should flash green on left (row 1L green)

  - PSR.to :sig2  //next row’s LED should flash green on left;  row above will flash red on right (row 2L green; row 1R red)

  - PSR.to :sig3  //row 3L green; row 2R red

  - PSR.to :conjugate  //row 4L green; row 3R red

  - PSR.to :sub1  //row 5L green; row 4R red

  - PSR.to :sub2  //row 6L green; row 5R red

  - PSR.to :wash  //row 7L green; row 6R red

  - PSR.to :flush  //row 8L green; row 7R red

  - PSR.to :air  //row 8R red (none green)

- Rotary Valve Testing

  - Connect ‘bottom row’ connectors of motor/encoder test cable to motor and encoder

  - Connect other end of motor/encoder test cable to J1 of rValve board

  - PSV.advance 1; PSV.retard 1  //motor should spin one way, then the other

  - Remove connector from J1 of rValve board and connect to J3

  - PTV.advance 1; PTV.retard 1  //motor should spin one way, then the other

  - Connect ‘top row’ connectors of motor/encoder test cable to motor and encoder (leave other end connected to J3)

  - PRV.advance 1; PRV.retard 1  //motor should spin one way, then the other

  - Remove connector from J3 of rValve board and again connect it to J1

  - PBV.advance 1; PBV.retard 1  //motor should spin one way, then the other

- Heater Control Testing

  - Connect 20 pin connector from protoboard’s bank of LEDs to J5 of rValve board

  - Measure voltage on exposed red and black wires bundled with the wires to the two BNC connectors  //red should be +5V relative to ground (black)

  - Yellow/Black BNC is heater control signal: connect to DMM

  - PH.hold 700,100  //voltage should increase dramatically (???)

  - PH.stop  //voltage should return to the previous value (???)