Dwarf Sensor Board Setup and Testing

**Power up system with core, rValve, and sensor boards**

Install jumpers on sensor board JP1: jumper 1-2 and 3-4

Set dip switch on tested core board to “processing” address (on/off/on/off/off/off)

With power off, install sensor board on top of tested rvalve and core board

Set power supply voltage to 13V +/- 1V

Connect power supply to Sleepy J2  //Board should draw nominal 0.2A – 0.3A

Attach DwarfBus ribbon cable from Sleepy J6/J8/J10 to Core J3

Attach power cable from J5 on Sleepy to J2 on core board

**Login**

telnet esptech

everlove

techtransfer

esp //make sure that there are no errors (30sec load time)

**setup**

require ‘testDwarf’ //True

testDwarf

**Testing**

(This testing done with Dwarf firmware v3.65, ESP2 v2.92 I2C Gateway 2.00)

- Temperature Testing

  - Connect temperature/pressure test cable to J1 on sensor board

  - Connect temperature branch of test cable to puck temp sensor board

  - PH.status  //note the current values for “temp” and “ambient” on screen

  - Place finger over sensor on end of puck temp sensor board

  - PH.status  //value for “temp” should increase while “ambient” does not

  - Place finger over sensor in middle of puck temp sensor board

  - PH.status  //value for “ambient” should increase while “temp” does not

  - Spray freeze spray lightly over both sensors on to puck temp sensor board

  - PH.status  //value for “ambient” and “temp” should drop dramatically

- Pressure Testing

  - Connect temperature/pressure test cable to J1 on sensor board (should still be connected from temperature test)

  - Connect pressure branch of test cable to pressure sensor and syringe

  - PS.status  //note (inPress? press1?) value

  - Push plunger on syringe

  - PS.status  // (inPress? press1?) value should increase

  - Pull plunger on syringe past starting point

  - PS.status  // (inPress? press1?) value should decrease below original value