Building a Heater assembly (1002131)

Parts

· Main parts

	Item
	Description

	1002131-001
	Sample Heater Base

	1002131-003
	Watlow Heater Mount

	1002131-004
	3/8” – 40 Nut

	1002131-005
	Teflon Gasket

	P-200181-006
	Titanium Conductive Mount

	1002115
	Watlow Thin Film Heater


· Small parts

	Qty
	Size
	Item

	1
	BUNA-N 70 Shore A 2-008
	O-Ring 

	1
	CR 16003
	Oil Seal

	1
	C150-H2-S17
	Century Spring - Wave Spring

	
	Dow Corning 340
	Conductive Heat Sink Compound


· Fasteners

	Qty
	Size
	Item

	1
	Truarc 5100-112 - S1.125
	External Snap Ring

	4
	#8-32 x 5/8”
	Socket Head Cap Screw

	4
	#8
	Spring-lock washer


Tools

· Snap Ring Pliers
· Vise

· 9/64” Hex key

Assembly

1. Trim the Heat Conductive Adhesive on the Watlow Heater to fit the Teflon gasket.
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2. Slide the Teflon gasket down the Heater wires.
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3. Place the Wave Spring onto the Watlow Heater Mount.
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4. Place the Watlow Heater Mount into the Heater Base.
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5. Temporarily place the Conductive Mount into the Watlow Heater Mount.
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6. Place the Assembly into a vise and compress the spring until the Snap ring groove is exposed.
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7. Using the Snap Ring Pliers, slip the Ring into the groove. 
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8. Release the Assembly from the Vise and remove the Conductive Mount.

9.  Spread the Heat sink compound onto the bottom of the Watlow Heater. Slide the Heater wires through the Teflon seal and the Heater mount. 
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10.  Replace the Conductive Mount and secure with the 3/8”- 40 Nut.  
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11.  Place the O-Ring into the groove on the Conductive Cap.
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