(DRAFT)

To revive the Deep ESP in 2013:

- Ambient reservoir is needed. (The blue bag tore at a fatigue point during the last recovery.) There are spare bags, identical to the one that tore. We could also look to purchase or design a rolling diaphragm compensator. Needs to be 30L.

- A sea pressure measuring transducer can be installed in the Sample can, directly measuring sea pressure. This will avoid the CTD/DWSM pressure calibration step, and the no-CTD issue during shore pressure testing.

- Testing of the bag switch assemblies shows that almost every barb fitting slowly leaks silicone oil, even with a clamp. A patch may be to glue every connection. But the switch assembly and bag cap should be redesigned.

- A tubing modification in the resample can will make air purging possible after the cans are sealed.

- Addition of 375V chargers to 48V battery and bus. (Mostly done, need to test in the cold for heating issues.)

- Replace Ethernet hub with one that auto senses.

- Inspect sphere oring stock for aging. Order new ones if needed.

- Thermal management of ESP and sphere, with new chargers a bypass of the insulation will likely be needed.

- Make new drop weight(s). We no longer have one (was left at Juan de Fuca)

- DWSM cleaning process and equipment is needed for longer term operation. (Ambient toxicant, flush, etc.)

- Battery replacement in DSPL boxes

Other Platform devices:

- 4k ISUS was destroyed. Is a new one needed, was there ever a determination of how or when it flooded? Will there be changes to the probe assembly?

- ISMS… is it going, are there changes?

- CTD, are we adding another sensor? Oxygen?

- DWSM Toxicant storage and selection valves, and protocols for cleaning DWSM.

- Are we planning for two sampling points (like the Juan de Fuca wand?)

Rebuild prior to deployment.

- The bags, tubing, and plastic fittings need to be replaced. They were flushed out but left in so assembly is intact.

- Bag switches need to be checked and set point adjusted and verified.

- Bulkhead valve assemblies cleaned and tested, over pressure set points adjusted if needed.

- Full flush out of fittings and tubing in DWSM and bulkhead valves.

- Clean up of High Pressure Interface, flush out rust

- Full cycle pressure test at or above planned deployment pressure.