Deep ESP 2011 Juan deFuca Deployment July 2011

TO DO LIST

Deep ESP Elevator Equipment

Sphere:  	Gordon

MFB #4?

On elevator:	DWSM

CTD

ISUS

ISMS

ISMS:		Send back to Harvard?

Get dimensions and mass of other housing, recalc weight balance

Make new mounts

Modify cabling?

Receive new pressure housing

Assemble, and Test

Mount on elevator

ISUS:		Receive fittings and probes

MinK’s in

P Caps in

Relief Valves ordered

P Test covers in shop

Orings/seals in

Order new probes

Assemble 4000m housings

Rosal group mounting MinK’s

Leak and hydro test

Mod cables

Wire ISUS

Calibrate ISUS for sulfides

Assemble ISUS for deployment

Leak test, check ground isolation

Mount on elevator

WAND:	Have functional reqmts? Comments?

Tested optional pumping system, review options.

Finalize schematic

Order components

Design parts

Assemble system

Test in tank

Measure flowrate, adjust.

HPI:		Order parts

Replace pump, valve, hoses, etc

Proof test HPI

Set relief valves

DWSM: 	New jacket bolt torque tool/method

Evaluate open rings for new pressure…

Compliance bags issues

Proof test DWSM for JdeF depth operation

Inspect IEP Pump

IEP accumulator?

Designed

Parts to shop

Assemble

Test

Install

System test, Low pressure

High Pressure test.

Adjust IEP and dwarf driver

Test IEP more

Test out HPP encoder

Check with mfg what damage may be from

Review HPP drive methodology

SPHERE:	New isolation valves

New valves ordered

Test and validate new valve tree (including overpressure valves)

Add heat exchangers to heater ring

Heat exchangers ordered

Dogleg to machine shop

Modify hanger ring

Replace purge plugs

ELEVATOR: Rewire/cast Battery Combiner box

Acquire back up modems?

Wet mate release for recovery, ROV/sled pack

SHIP:		Check on WF Hull transducer

Review weight ring for dunking transducer

Review tow point/recover lines with WF crew

CHECKLIST:

In trailer work

Proof test system to JdeF depth (2000m)

Remove Gordon

Modify sphere components

Modify DWSM?

Replace water bags?

Test bag switch points

Flush/clean HPI

Core process, develop array, PCR?

Full cycle test Under HP

CTD/ISUS in a bucket

Chill Down to 4C

Full cycle cold test.

Test ISMS in lab

Can ISMS

Canning Gordon

Add hobo’s

Load reagents

Load Pucks

Empty wastes

Install sphere insulation

Close Sphere

Torque bolts, cross pattern  (xx in-lbs)

Go-no go gauge test

Leak test and N2 purge.

Clean and reset sphere relief (5 turns out, 4 psi)

Can DWSM

Leak test DWSM bells

Return to trailer

Chill Down to 4C

Post CAN runs

Elevator and Foam pack to High Bay

Inspect sphere vent ports

Install Drop Weight Release

Mount Dummy Jack (Drill for thru bolts.)

Rebuild battery boxes (8)

Mount Battery Combiner box

Mount 8 Batteries on elevator

Install ISMS

Change CTD batts

Install CTD

Install ISUS

Add foam to pack

Change Ac modem batteries

Install Ac. Modem

Install Pinger, Strobe, etc.

Attach bumper rings

Test Tank

Weigh, trim and balance

Test burn wire release (3 minutes)

Test Fluidic Stream (Sample point/DWSM/ISMS/ISUS/CTD…)

Test manual release

Modem Comm

Post Can QC

Prep for deploy

Make lists of equipment, tools, spare, etc.

Pack it all

Clean/Mop trailer

Load Trailer

Anchor equip in trailer

Verify generator and trailer compatibility

Set trailer temp to 11C for shipping

Load ship