## Slide 1: 4km DWSM Operation
Notes: Add seawater sensors at bottom of each oil chamber
Sample tube in bag to ¾ depth
Upper/lower purge may need to wait until, or be moved to after sampling.
May need to add an RO water rinse fluid. (non-toxicant) Verify that air compliance is at right volume.
Also look at cast 17-4 for bell housings.
brent: One could use a serial port or ethernet for communications to the ESP
brent: Core must be work from 2C to 35C
gene: Test urathane bags for oil compatibility
gene: Pack toxicant bag in PVC or similar hard container
gene: Don't servo on pressure, just monitor it during decompression
brent: measure amplitude of sample pump's pressure impulses
brent: Modify ESP scripts to purge into DWSM before starting sample processing
gene: Add an external bag for DI Water to mix with toxicant
brent: Add filter on raw water inhaled into bags when rinsing them
brent: Measure compliance when filling bells with oil
jon, gene, et al.: Invert bells when filling with oil
gene: Dish out top plate to control air bubble size
gene: consider stainless for casting bell housings
jon: 3 threads on HP oil pump motor housing is bare minimum
gene: O-ring on HP oil pump motor housing may be damaged on assembly
mike: put O-ring on inside wall to prevent this

## Slide 2: Mechanical Systems Simple Operating Theory
Intake seawater
Decompress
Transfer and release gas
Sample (thru core)
Recompress
Exhaust seawater
Decon System

## Slide 3
Oil
Res.
Deep Water Sampler
Process Diagram
Flush
TSV
SRV
FOV
POV
HP Pump
IE Pump
ReSample Bag
ESP Intake
ESP Exhaust
Sample
Pump
Sample Bag
SOV
RPT
SPT
SSS
SES
SBS
RSS
RES
RBS
RS
IS
REV
SBV
TOX
REX
RST
RST
SST
Exhaust
Exhaust
Intake
SAX
SAX
HPP
IEP
Free
SMP
Free
RO
ISUS
CTD
Sample Point
WAND
Sea Vent
SP
TF
FL
TX
(50’)
+
-
+
-

## Slide 4
Oil
Res.
Deep Water Sampler
Process Diagram
SRV
FOV
POV
HP Pump
IE Pump
ReSample Bag
ESP
Sample
Pump
Sample Bag
SOV
RPT
SPT
SSS
SES
SBS
RSS
RES
RBS
RS
IS
SBV
REX
RST
RST
SST
SAX
SAX
HPP
IEP
Free
SMP
Free
RO
+
-
+
-

## Slide 5
High Pressure Interface
Process Diagram
OIL
HPO
Fill Pump
BYP
+
-
Sea
Water
SW
OIL
BYP
HPW
+
-
+
-
DWSM
High
Pressure
Pump
Reg.
Orifice
Valve
Relief
Valve
Check
Valve
Oil Fill Valve
HP Oil Valve
HP Water
Valve
Water
Fill
Valve
Accumulator
Air Solenoid
Valve
Air Reg

## Slide 6
Deep Water Sampler
Aqua Lung
ESP In
ESP Out
Sample Bag
Pressure Plate
Oil Chamber
Resample Bag
Pressure Plate
Oil Chamber
Resample
HP pump
Valves
IE pump
Sample
Valves
Gateway
Power
Sample
pump
Pressure
Sensor
Pressure
Sensor
SSS
SES
SBS
ESS
EES
EBS
Stir
Pump
Stir
Pump
Wand
Pump, Cap and Manifold
System Power
Battery