ESP LLNL CDR Follow-Up
07May07
Attending:

Doug Pargett, Scott Jensen, Chris Scholin, Brent Roman, Addie Harris, Jim Birch





Agenda:

1. Review LLNL CDR status

2. Action items

3. Legal status of IP

New Items:
1. From the science side, we are happy with the development.
2. Limitation is single channel reaction at a time (4 channel detection, but 1 reaction at a time)

3. Not how science would run in the lab.

4. Technically it is good; heater is new, untested; software and firmware, temp sensitivity.

5. May need to ask LLNL for calibration process that could be done remotely, or a temp probe that could be detected. (Molecular beacon.)
6. We want them to characterize the stability of the calibration.

7. Handshaking issue, they have not committed to fixing it.

Action items:

1. Formal responses to CDR… Brent/Doug/etc
2. Need progress on mFB, (Pargett)
3. Need to work on SPE, (Scholin)
4. Software driver for PCR (originally requested for Bob Helien.) Need to find out who will be able to work on this.
5. Intellectual Property status and allowable, usage, production, publishing, etc with regards to the PCR module. All legal issues from now on are Chris Scholin and Jim Birch to address.
