ESP Project Meeting 

13 Sept 2007
Attending:

Doug Pargett, Scott Jensen, Addie Harris, Brent Roman, Jim Birch, Cheri Everlove





Next milestones: 
Agenda:
Old Items:

1. 1k DWSM to get cleaned, disassembled, inspected. Mothball and store for future use. Still needs to happen.
2. After network deployment- general discussion workshop in December, C-MORE focused workshop in January.
3. Brent to document how to calibrate an instrument (Cheri or Doug) and look into developing self calibration. 1 wk to document, 3 weeks to code an automatic cal procedure.

New Items:

4. GG/NEO bug list:

a. NEO has a heating problem; work around is functioning, need to inspect and test when back in lab.

b. Bulk head valves; rust issue.

c. Camera seal spring too stiff

d. Elevator rod sleeve; need one that works and doesn’t clunk.

e. Firmware issues; should have a way of specifying the current limit

f. Forearm slide; possible dust, possible additional runs
g. NEO samples slow

h. Swirly images; still no conclusive reason why
i. New code; a bunch of changes that need to be tested in the instrument.

5. Core build, next 3 instruments

a. All the COTS parts for replacing the MFB borrowed parts (needed for unit 3)

b. Heater assemblies are going together this week. Some interference, being modified. 
c. Sub assemblies are completed for the first 2 units. 

d. Schedule is to complete build by Dec 07, then turn over to QC.

e. Cheri to work with Brent to find him mort play parts to test firmware.

6. 4k 

a. We’ve got some new ideas, we’ll need to get science input and expect that they’ll have a number of concerns and ideas. Science needs to have time for design reviews. Need to have design review Monday or Tuesday next week.

b. 4k housing changing from 40 to 34 inch ID as primary design. There is no value in additional space in the deep housing that doesn’t match space in the shallow housing.

c. Brown to make up a new design and get preliminary quotes. If it’s close, we’ll go with the larger

7. MFB

a. Jim and Addie are ready to clean the MFB and run real reagents.

b. Getting reagent plumbed today.

c. Addie starting re-design of the MFB to accept two PCR modules, and make room for reagents and parts.

d. Also need to start getting parts for mounting and structure of the MFB.

e. Brent thinks the counts to ml is off, and needs to be checked. Addie will check that.

f. Scott’s not sure shat we’re getting by having 2 PCR modules. For now we just want to have the potential.
