ESP Project Meeting 

23 Aug 2007
Attending:

Doug Pargett, Scott Jensen, Chris Scholin, Addie Harris, 





Next milestones: 
· Networked Surface ESP, Aug 30 – Sept 27 (GG, A1, 2 Moorings, ISUS, CTD, 2 turns?)

Agenda:
Old Items:

1. Science side working on PCR-reagent stability over time.
2. Science also to work on the internal positive control (IPC) two options.
3. 1k DWSM to get cleaned, disassembled, inspected. Mothball and store for future use. Still needs to happen.
4. After network deployment- general discussion workshop in December, C-MORE focused workshop in January.
5. Next 3 units are to be completed and tested this year.

6. Need to account for parts that were “borrowed” for MFB, or repairs, and get ordered.
7. Addie to learn the PCR process from Preston once science has time (after deployment)

8. We probably have enough info to design up the heater.


New Items
9. Brent found an anomaly in GG, one of the clamps spins on the motor shaft. Scott to check it out. Doug checked (and tightened and is seems good now.)

10. Jensen has 2 ISUS cables; Luke will have two ISUS ready on Friday.

11. Need to make screens, Jensen to do.
12. We would have 2 version of software (NEO code and GG code.) Brent to deal with CVS issues, etc.

13. Roman to prep the CTD’s for the cruise.

14. Right now NEO is not qual’ed for deployment. (Not behaving as well as GG.)
15. Syringes leaked in the 10C (8C) test, syringes have been replaced with ones tested to 2C (25mL) and 4C (10ml).
16. Brent fixed a retry issue on NEO, code needs to get put on GG. Can test on mock-run mode.

17. Auto exposure change is already on both machines. 
18. End of this week we can probably lock down the software. Still adjusting scripts, and ESP code.

19. Addie made a new MFB diagram.

20. MFB, start testing with the speeds as slow as they can go. Brent modified the firmware on MFB and got close to 1 micro-L per sec flow rates.

21. Addie ran the reaction slug at 12, and the air plugs would break up with 5 micro-L, but not with 10 micro-L.

22. Addie to work on the decon scripts in coordination with a slug process. Run it a bunch of times then it’ll be ready for the real thing. Brent recommends that we home the syringe often.
23. Need to start talking to LLNL more often. Bubble detection, PCR unit delivery, kit prices, kit assembly, etc. Chris to talk to John D.
24. Addie buying prototyping parts (backfilling ones borrowed form Cheri.)

