## Slide 1
Sample processing system (based on heritage ESP) [temp]
Optical Microscope images (minimally) the 0.2 micron substrate; also the 10 micron pre-filter (and larger one if used)
Sample deposited on membrane (e.g. 0.2 micron SimPore microporous membrane that has surface area of total of ~ 5 mm2)
Motorized stage capable of x, y dimensional movement holds membrane cartridge; contains calibration substrates (?).
Magnification | 5.75
CCD | 256 x 512 pixel
Imaged area | 1 mm x 2 mm
Resolution | 4 microns/px
Light source | 3 red, 3 green, 3 blue and 3 UV LEDs
Heritage MECA
stage
1.02 mm
2.48 mm
AFM low resolution scanning : Small squares (100 x 100 mm) are scaled to the large rectangle
Piezoresistive cantilever
Sample Imaging Scenario:
2 low resolution OM images required under this configuration with 4 microns per pixel resolution to image 5 mm2 substrate.
Analyisis: JPEG compression identifies high information regions; alternatively use autonomous, AI-driven software for image analysis and decision-making to direct AFM to areas of interest in OM image.
AFM images at low resolution ## 50 to 100 mm sectors (representing 12-25 pixels from OM image). {optional ]
Analysis: JPEG/Machine Learning/AI decision making on AFM subsector selection.
Targeted subsectors at 10 x 10 mm or 5 x 5 mm are scanned for topography and force spectroscopy.
Analysis: process mechanical force data and volatilize soft/adhesive particles.
Soft or adhesive particles are volatilized and ingested into the Mass Spec (Baseline)
AFM high resolution subsector imaging 5-10 mm squares
Motorized AFM scanner moves in x, y, z dimensions to approach sample
To be revised: ANTONIE Schematic
Imaged area – low res | 50-100 micron square {optional}
Pixel density | 256 x 256 or 512 x 512 pixel?
Resolution | 390 or 195 nm/pixel
Imaged area –high res | 5 x 5 or 10 x 10 um
Pixel density | 256 x 256 pixel [lowest]
Resolution | 40 nm/pixel [req: image 100-200 nm cells/virus particles]
Cantilever | Piezoresistive ~ 20 nm tip
Scanner | x,y,z dimensions; precision req?
Calibration targets | Size and force detection
1.
2.
3.
4.
5.
Requirements | Sample melted
Sample hydrated | Insolubles removed
Sample characterized (e.g. conductivity, transmissivity) | Sample size fractionated (Pressure monitored)
FOV 1 image
Notes: To be edited and updated with relevant images; and specs

## Slide 2
G3: Characterize Surface properties:
Putative abiotic particle characteristics
Morphology supporting hypotheses for crystaline vs. amorphous particles
Size distribution
Mechanical properties
G2:Assess Habitability:
Particle aggregates suggest oceanic origin
abiotic evidences biogenic
G1: Search for biosignatures:
Cell-like or organic particle morphology
Size distribution of cell-like & organic particles
Mechanical properties suggest soft/adhesive objects
Particle or bulk sample is carbon-containing
Sample processing system (based on heritage ESP) [temp]
Optical Microscope images (minimally) the 0.2 micron substrate; also the 10 micron pre-filter (and larger one if used)
Sample deposited on membrane (e.g. 0.2 micron SimPore microporous membrane that has surface area of total of ~ 5 mm2)
Motorized stage capable of x, y dimensional movement holds membrane cartridge; contains calibration substrates (?).
Magnification | 5.75
CCD | 256 x 512 pixel
Imaged area | 1 mm x 2 mm
Resolution | 4 microns/px
Light source | 3 red, 3 green, 3 blue and 3 UV LEDs
Heritage MECA
stage
1.02 mm
2.48 mm
AFM low resolution scanning : Small squares (100 x 100 mm) are scaled to the large rectangle
Piezoresistive cantilever
AFM high resolution subsector imaging 5-10 mm squares
Motorized AFM scanner moves in x, y, z dimensions to approach sample
1.
2.
3.
4.
5.
Requirements | Sample melted
Sample hydrated | Insolubles removed
Sample characterized (e.g. conductivity, transmissivity) | Sample size fractionated (Pressure monitored)
FOV 1 image

## Slide 3: Antonie system components
Sample fractionation
AFM
Optical microscopy
Particle identification and size distribution. Fluorescence, other low res optical ideas?
Extends size distribution to below 3um. Fluorescence, other low res optical ideas?
Heritage scale
High risk/Low TRL
Low risk/High TRL
Ice chip melting, filtration and distribution on to substrates for Optical and Atomic Force Microscopy