Fido Control Board

Version 1.0

September 4, 2024

Inputs/Outputs

- Pressure Sensors: 2 4-20mA pressure sensors input circuits

  - Currently unused

  - Could be used to monitor incoming sample pressure

  - Power switched

- Force Resistor (Flexiforce): 1 analog input circuit with relay control

  - Used for sensing force of clamp

  - Clamp mechanics preloads sensor to 15 pounds

  - Unloads when in full open position

  - “relay” allows sensor to be disconnected and the circuits bias level measured

  - Always powered

- Environmental Sensor: 1 onboard humidity and pressure sensors, with SPI output

  - Used as “metadata”

  - Always powered

- Flow Sensor:

  - Input: 1 pulse counter input with voltage level shifting to 3.3V

  - Output: 24-10V fused at 375 mA output

  - Used to measure sample volume on the down stream side of puck

  - Power switched through a linear regulator

  - Regulator set to 28V to prevent over voltage if system power is above 28V

- Vacuum Pump:

  - Input: 1 pulse counter input with level shifting to 3.3V

  - Outputs:

    - 1 PWM speed control 0-5V

    - Switched battery voltage: 10-35V, fused at 6.3A, 4A max expected

  - Used to:

    - Manipulate pucks

    - Vacuum puck to empty air before sampling

    - Clean fluid from clamp

  - Due to its high start current (4A) it is switched on its own high current switch

  - Speed control is possible with PWM

  - Speed monitor is possible with returned frequency

- LED indicators:

  - 3 LED driving circuits:

    - one connected to 3.3V power

    - Two switched with GPIO

- NFC Reader:  three pin output for antenna

  - Tag transceiver to write and read meta data of tag

- Optical interrupters:

  - Inputs: 8 phototransistor collector inputs

  - Outputs: 4 LED pulse circuits (3.3V Pulse)

  - Used to locate Arm and Cylinder

  - Pulsed LED to create AC signal for ambient light insensitivity

- Puck Counter:

  - Input: 1 pulse counter input

  - Output: Switched 3.3V, 100mA max

    - Used to measure the number of pucks in a tube

    - Since it looks down the hole it should be able to see if the cylinder is installed

- Solenoids:

  - Outputs: 8 1A solenoid drivers

  - SPI control of solenoid drivers

- Sample Pump:

  - Comms: RS232

  - Output: 1 GPIO

- Stepper Motors:

  - Outputs: 5 2A stepper motor drives

  - GPIO controlled (Step, Dir, Sleep)

- Vacuum Pressure Sensor:

  - Input: analog voltage 0-3.3V proportional to pressure differential

  - Output: switched 3.3V

- Bucket and Cylinder Detect

  - Input: 2 GPIO

  - Output: 3.3V

Communications

- Programming: SWD 6pin programming header

- RS232 comms to sample pump

- USB Python Console

- USB FIDO Command Console

Power

- Input battery voltage range: 10-35V

  - 3.6 V Switching converter, designed for 1A output

  - 3.3V, 1A linear regulator

  - 5V, 2A max (Rock Pi)

  - 10-24V, 375mA linear regulator (flow sensor)

  - 5V, 200mA linear regulator (sample pump)