## Sheet1
Rev0
LL = Logic Level
Pin Name/Function | Format | Interface | Periperhal | Description | U Port | U Pin | CN | CN pin | CN port | U Port X | U Port # | Block | Pin
[5 empty] | PA0 | 34 | CN11 | 28 | PA0 | PA | 0
[5 empty] | PA1 | 35 | CN11 | 30 | PA1 | PA | 1
Force | Voltage | Analog (0-3.3) | ADC12_IN7 | Analog input that sense the clamping pressure of the clamp on the puck | PA2 | 36 | CN12 | 35 | PA2 | PA | 2 | Flexi | J8-1
Pressure1 | 4-20mA | Analog (0-3.3V) | ADC12_IN8 | 4-20mA pressure sensor using a 165Ohm sense resistor | PA3 | 37 | CN12 | 37 | PA3 | PA | 3 | 4-20mA | 1
Pressure2 | 4-20mA | Analog (0-3.3V) | ADC12_IN9 | 4-20mA pressure sensor using a 165Ohm sense resistor | PA4 | 40 | CN11 | 32 | PA4 | PA | 4 | 4-20mA | 3
[5 empty] | PA5 | 41 | CN12 | 11 | PA5 | PA | 5
[5 empty] | PA6 | 42 | CN12 | 13 | PA6 | PA | 6
[5 empty] | PA7 | 43 | CN12 | 15 | PA7 | PA | 7
|  | USB | OTG_FS_SOF | Reserved for USB for Python | PA8 | 100 | CN12 | 23 | PA8 | PA | 8
|  | USB | OTG_FS_VBUS | Reserved for USB for Python | PA9 | 101 | CN12 | 21 | PA9 | PA | 9
|  | USB | OTG_FS_ID | Reserved for USB for Python | PA10 | 102 | CN12 | 33 | PA10 | PA | 10
|  | USB | OTG_FS_DM | Reserved for USB for Python | PA11 | 103 | CN12 | 14 | PA11 | PA | 11
|  | USB | OTG_FS_DP | Reserved for USB for Python | PA12 | 104 | CN12 | 12 | PA12 | PA | 12
|  | USB | OTG_FS_NOE | Reserved for USB for Python | PA13 | 105 | CN11 | 13 | PA13 | PA | 13
[5 empty] | PA14 | 109 | CN11 | 15 | PA14 | PA | 14
[5 empty] | PA15 | 110 | CN11 | 17 | PA15 | PA | 15
Motor STEP 1 | LL output | Output Capture | TIM3_CH3 | Step control of motor. The step lines uses Output Capture to time the step pulses | PB0 | 46 | CN11 | 34 | PB0 | PB | 0 | motor1 | 15
[5 empty] | PB1 | 47 | CN12 | 24 | PB1 | PB | 1
|  | USB | OTG_FS_CRS_SYNC | Reserved for USB for Python | PB2 | 48 | CN12 | 22 | PB2 | PB | 2
[5 empty] | PB3 | 133 | CN12 | 31 | PB3 | PB | 3
[5 empty] | PB4 | 134 | CN12 | 27 | PB4 | PB | 4
[5 empty] | PB5 | 135 | CN12 | 29 | PB5 | PB | 5
[5 empty] | PB6 | 136 | CN12 | 17 | PB6 | PB | 6
[5 empty] | PB7 | 137 | CN11 | 21 | PB7 | PB | 7
Solenoid Fault 1 | LL input |  | GPIO | MAX22200 Fault signal #1 | PB8 | 139 | CN12 | 3 | PB8 | PB | 8 | solenoid1 | 16
Solenoid Fault 2 | LL input |  | GPIO | MAX22200 Fault signal #2 | PB9 | 140 | CN12 | 5 | PB9 | PB | 9 | solenoid2 | 16
Flow | LL input | Input Capture | TIM2_CH3 | Paddle wheel output from flow sensor. Input capture for flow rate calculation along with totalization | PB10 | 69 | CN12 | 25 | PB10 | PB | 10 | flow | 2
Motor STEP 3 | LL output | Output Capture | TIM2_CH4 | Step control of motor. The step lines uses Output Capture to time the step pulses | PB11 | 70 | CN12 | 18 | PB11 | PB | 11 | motor3 | 15
Solenoid 1 SS | LL output |  | GPIO | Chip select of MAX22200 #1 | PB12 | 73 | CN12 | 16 | PB12 | PB | 12 | solenoid1 | 2
Solenoid 2 SS | LL output |  | GPIO | Chip select of MAX22200 #2 | PB13 | 74 | CN12 | 30 | PB13 | PB | 13 | solenoid2 | 2
Load Cell 1 CLK | LL output |  | GPIO | Load Cell Clock. HX711 defines up to 26 clock pulses per sample as well as programming of next sample | PB14 | 75 | CN12 | 28 | PB14 | PB | 14 | scale1 | 4
Load Cell 2 CLK | LL output |  | GPIO | Load Cell Clock. HX711 defines up to 26 clock pulses per sample as well as programming of next sample | PB15 | 76 | CN12 | 26 | PB15 | PB | 15 | scale2 | 4
[5 empty] | PC0 | 26 | CN11 | 38 | PC0 | PC | 0
[5 empty] | PC1 | 27 | CN11 | 36 | PC1 | PC | 1
[5 empty] | PC2 | 28 | CN11 | 35 | PC2 | PC | 2
[5 empty] | PC3 | 29 | CN11 | 37 | PC3 | PC | 3
Sample Pump TX | RS-232 | UART | UART3 |  | PC4 | 44 | CN12 | 34 | PC4 | PC | 4
Sample Pump RX | RS-232 | UART | UART3 |  | PC5 | 45 | CN12 | 6 | PC5 | PC | 5
Motor Sleep 1 | LL output |  | GPIO | Logic Level where sleep is low | PC6 | 96 | CN12 | 4 | PC6 | PC | 6 | motor1 | 14
Motor Sleep 2 | LL output |  | GPIO | Logic Level where sleep is low | PC7 | 97 | CN12 | 19 | PC7 | PC | 7 | motor2 | 14
Motor Sleep 3 | LL output |  | GPIO | Logic Level where sleep is low | PC8 | 98 | CN12 | 2 | PC8 | PC | 8 | motor3 | 14
Motor Sleep 4 | LL output |  | GPIO | Logic Level where sleep is low | PC9 | 99 | CN12 | 1 | PC9 | PC | 9 | motor4 | 14
Motor Sleep 5 | LL output |  | GPIO | Logic Level where sleep is low | PC10 | 111 | CN11 | 1 | PC10 | PC | 10 | motor5 | 14
Motor DIR 1 | LL output |  | GPIO | Logic level to define forward or reverse rotation of motor | PC11 | 112 | CN11 | 2 | PC11 | PC | 11 | motor1 | 16
Motor DIR 2 | LL output |  | GPIO | Logic level to define forward or reverse rotation of motor | PC12 | 113 | CN11 | 3 | PC12 | PC | 12 | motor2 | 16
Motor DIR 3 | LL output |  | GPIO | Logic level to define forward or reverse rotation of motor | PC13 | 7 | CN11 | 23 | PC13 | PC | 13 | motor3 | 16
[5 empty] | PC14 | 8 | CN11 | 25 | PC14 | PC | 14
[5 empty] | PC15 | 9 | CN11 | 27 | PC15 | PC | 15
|  |  |  | Reserve for SPI2 | PD0 | 114 | CN11 | 57 | PD0 | PD | 0
|  |  |  | Reserve for SPI2 | PD1 | 115 | CN11 | 55 | PD1 | PD | 1
BME280 SS | LL output |  | GPIO | Chip select of BME280 | PD2 | 116 | CN11 | 4 | PD2 | PD | 2 | BME280 | 5
|  |  |  | Reserve for SPI2 | PD3 | 117 | CN11 | 40 | PD3 | PD | 3
|  |  |  | Reserve for SPI2 | PD4 | 118 | CN11 | 39 | PD4 | PD | 4
ST25R3916B SS | LL output |  | GPIO | Chip select of STR25 | PD5 | 119 | CN11 | 41 | PD5 | PD | 5 | NFC
Solenoid Enable 1 | LL output |  | GPIO | MAX22200 Enable signal #1 | PD6 | 122 | CN11 | 43 | PD6 | PD | 6 | solenoid1 | 18
Solenoid Enable 2 | LL output |  | GPIO | MAX22200 Enable signal #2 | PD7 | 123 | CN11 | 45 | PD7 | PD | 7 | solenoid2 | 18
Indicator 1 | LL output |  | GPIO | Output for a TBD LED status indicator | PD8 | 77 | CN12 | 10 | PD8 | PD | 8
Indicator 2 | LL output |  | GPIO | Output for a TBD LED status indicator | PD9 | 78 | CN11 | 69 | PD9 | PD | 9
Motor DIR 4 | LL output |  | GPIO | Logic level to define forward or reverse rotation of motor | PD10 | 79 | CN12 | 65 | PD10 | PD | 10 | motor4 | 16
Motor DIR 5 | LL output |  | GPIO | Logic level to define forward or reverse rotation of motor | PD11 | 80 | CN12 | 45 | PD11 | PD | 11 | motor5 | 16
[5 empty] | PD12 | 81 | CN12 | 43 | PD12 | PD | 12
[5 empty] | PD13 | 82 | CN12 | 41 | PD13 | PD | 13
[5 empty] | PD14 | 85 | CN12 | 46 | PD14 | PD | 14
Motor STEP 2 | LL output | Output Capture | TIM4_CH4 | Step control of motor. The step lines uses Output Capture to time the step pulses | PD15 | 86 | CN12 | 48 | PD15 | PD | 15 | motor2 | 15
Cylinder 0 | LL input |  | GPIO | 1 of 6 total inputs that define the position of the turntable. There are 12 valid states which map to the 12 tubes of the cylinder | PE0 | 141 | CN12 | 64 | PE0 | PE | 0 | opto | 2
Cylinder 1 | LL input |  | GPIO | 2 of 6 | PE1 | 142 | CN11 | 61 | PE1 | PE | 1 | opto | 3
Cylinder 2 | LL input |  | GPIO | 3 of 6 | PE2 | 1 | CN11 | 46 | PE2 | PE | 2 | opto | 4
Cylinder 3 | LL input |  | GPIO | 4 of 6 | PE3 | 2 | CN11 | 47 | PE3 | PE | 3 | opto | 5
Cylinder 4 | LL input |  | GPIO | 5 of 6 | PE4 | 3 | CN11 | 48 | PE4 | PE | 4 | opto | 6
Cylinder 5 | LL input |  | GPIO | 6 of 6 | PE5 | 4 | CN11 | 50 | PE5 | PE | 5 | opto | 7
Arm 0 | LL input |  | GPIO | 1 of 4 total inputs that define the position of the arm. The are may pass through a numbre of usused states with 4 definite states | PE6 | 5 | CN11 | 62 | PE6 | PE | 6 | opto | 8
Arm 1 | LL input |  | GPIO | 2 of 4 | PE7 | 58 | CN12 | 44 | PE7 | PE | 7 | opto | 9
Arm 2 | LL input |  | GPIO | 3 of 4 | PE8 | 59 | CN12 | 40 | PE8 | PE | 8 | opto | 10
Arm 3 | LL input |  | GPIO | 4 of 4 | PE9 | 60 | CN12 | 52 | PE9 | PE | 9 | opto | 11
Vacuum detect | LL input |  | GPIO | Digital input that indicates that the presence of adequate vacuum within the suction grabber | PE10 | 63 | CN12 | 47 | PE10 | PE | 10 | opto | 13
Spare | LL input |  | GPIO |  | PE11 | 64 | CN12 | 56 | PE11 | PE | 11 | opto | 14
Load Cell 1 DATA | LL input |  | GPIO | Load Cell Data | PE12 | 65 | CN12 | 49 | PE12 | PE | 12 | scale1 | 3
Load Cell 2 DATA | LL input |  | GPIO | Load Cell Data | PE13 | 66 | CN12 | 55 | PE13 | PE | 13 | scale2 | 3
Cylinder Present | LL input |  | GPIO | Input indicating that a cylinder has been installed | PE14 | 67 | CN12 | 51 | PE14 | PE | 14
Bucket Present | LL input |  | GPIO | Input indicating that a puck bucket is in place | PE15 | 68 | CN12 | 53 | PE15 | PE | 15
[5 empty] | PF0 | 10 | CN11 | 53 | PF0 | PF | 0
[5 empty] | PF1 | 11 | CN11 | 51 | PF1 | PF | 1
[5 empty] | PF2 | 12 | CN11 | 52 | PF2 | PF | 2
[5 empty] | PF3 | 13 | CN12 | 58 | PF3 | PF | 3
[5 empty] | PF4 | 14 | CN12 | 38 | PF4 | PF | 4
[5 empty] | PF5 | 15 | CN12 | 36 | PF5 | PF | 5
Motor STEP 4 | LL output | Output Capture | TIM5_CH1 | Step control of motor. The step lines uses Output Capture to time the step pulses | PF6 | 18 | CN11 | 9 | PF6 | PF | 6 | motor4 | 15
Motor STEP 5 | LL output | Output Capture | TIM5_CH2 | Step control of motor. The step lines uses Output Capture to time the step pulses | PF7 | 19 | CN11 | 11 | PF7 | PF | 7 | motor5 | 15
[5 empty] | PF8 | 20 | CN11 | 54 | PF8 | PF | 8
[5 empty] | PF9 | 21 | CN11 | 56 | PF9 | PF | 9
[5 empty] | PF10 | 22 | CN12 | 42 | PF10 | PF | 10
[5 empty] | PF11 | 49 | CN12 | 62 | PF11 | PF | 11
[5 empty] | PF12 | 50 | CN12 | 59 | PF12 | PF | 12
[5 empty] | PF13 | 53 | CN12 | 57 | PF13 | PF | 13
[5 empty] | PF14 | 54 | CN12 | 50 | PF14 | PF | 14
[5 empty] | PF15 | 55 | CN12 | 60 | PF15 | PF | 15
[5 empty] | PG0 | 56 | CN11 | 59 | PG0 | PG | 0
[5 empty] | PG1 | 57 | CN11 | 58 | PG1 | PG | 1
common SPI SCLK | SPI | SPI-CLK | SPI1 | SPI clock to MAX22200's, BME280 and ST25R3916B (NFC Tag reader) | PG2 | 87 | CN11 | 42 | PG2 | PG | 2 | solenoid1 | 4 | solenoid2 | 4 | BME280 | 8 | NFC | ?
common SPI MISO | SPI | SPI-MISO | SPI1 | SPI SDI Data into STM32 | PG3 | 88 | CN11 | 44 | PG3 | PG | 3 | solenoid1 | 8 | solenoid2 | 8 | BME280 | 7 | NFC | ?
common SPI MOSI | SPI | SPI-MOSI | SPI1 | SPI SDO Data out of STM32 | PG4 | 89 | CN12 | 69 | PG4 | PG | 4 | solenoid1 | 6 | solenoid2 | 6 | BME280 | 6 | NFC | ?
[5 empty] | PG5 | 90 | CN12 | 68 | PG5 | PG | 5
[5 empty] | PG6 | 91 | CN12 | 70 | PG6 | PG | 6
Console TX | RS-232 | UART | LPUART1 |  | PG7 | 92 | CN12 | 67 | PG7 | PG | 7 | console | 7
Console RX | RS-232 | UART | LPUART1 |  | PG8 | 93 | CN12 | 66 | PG8 | PG | 8 | console | 8
[5 empty] | PG9 | 124 | CN11 | 63 | PG9 | PG | 9
[5 empty] | PG10 | 125 | CN11 | 66 | PG10 | PG | 10
[5 empty] | PG11 | 126 | CN11 | 70 | PG11 | PG | 11
[5 empty] | PG12 | 127 | CN11 | 65 | PG12 | PG | 12
[5 empty] | PG13 | 128 | CN11 | 68 | PG13 | PG | 13
[5 empty] | PG14 | 129 | CN12 | 61 | PG14 | PG | 14
[5 empty] | PG15 | 132 | CN11 | 64 | PG15 | PG | 15
[5 empty] | PH0 | 23 | CN11 | 29 | PH0 | PH | 0
[5 empty] | PH1 | 24 | CN11 | 31 | PH1 | PH | 1
[5 empty] | PH3 | 138 | CN11 | 7 | PH3 | PH | 3
[5 empty] | NRST | 25 | CN11 | 14 | RESET
[5 empty] | VBAT | 6 | CN11 | 33 | VBAT
[5 empty] | VDD | 17
[5 empty] | VDD | 39
[5 empty] | VDD | 52
[5 empty] | VDD | 62
[5 empty] | VDD | 72
[5 empty] | VDD | 84
[5 empty] | VDD | 108
[5 empty] | VDD | 121
[5 empty] | VDD | 144
[5 empty] | VDDA | 33
[5 empty] | VDDIO2 | 95
[5 empty] | VDDIO2 | 131
[5 empty] | VDDUSB | 106
[5 empty] | VREF- | 31
[5 empty] | VREF+ | 32 | CN12 | 7 | VREF+
[5 empty] | VSS | 16
[5 empty] | VSS | 38
[5 empty] | VSS | 51
[5 empty] | VSS | 61
[5 empty] | VSS | 71
[5 empty] | VSS | 83
[5 empty] | VSS | 94
[5 empty] | VSS | 107
[5 empty] | VSS | 120
[5 empty] | VSS | 130
[5 empty] | VSS | 143
[5 empty] | VSSA | 30
[7 empty] | CN11 | 10 | -
[7 empty] | CN12 | 8 | U5V
[7 empty] | CN11 | 5 | VDD
[7 empty] | CN11 | 24 | VIN
[7 empty] | CN11 | 26 | -
[7 empty] | CN11 | 67 | -
[7 empty] | CN11 | 16 | 3.3V
[7 empty] | CN11 | 18 | 5v
[7 empty] | CN12 | 32 | AGND
[7 empty] | CN11 | 6 | E5V
[7 empty] | CN11 | 8 | GND
[7 empty] | CN11 | 19 | GND
[7 empty] | CN11 | 20 | GND
[7 empty] | CN11 | 22 | GND
[7 empty] | CN11 | 49 | GND
[7 empty] | CN11 | 60 | GND
[7 empty] | CN12 | 9 | GND
[7 empty] | CN12 | 20 | GND
[7 empty] | CN12 | 39 | GND
[7 empty] | CN12 | 54 | GND
[7 empty] | CN12 | 63 | GND
[7 empty] | CN11 | 12 | IOREF