## Slide 1: 3G ESP General Overview of Cartridges Archive and Lyse-n-Go
MBARI, Monterey Bay Aquarium Research Institute
2019 MBARI, Confidential Information

## Slide 2
3G ESP Overview
Sampling
(filter water)
Archive
Hand-off
Process
(create
homogenate)
(add preservative,
bring home)
Send Results Home
Analyze
Cartridge
Analytical Module
3G ESP Sampler
2019 MBARI, Confidential Information
Notes: Points
1. First part: focus on the cartridge.
2. Each cartridge is specialized. Allows the user to develop own pathway. Process from encompassing filtration, and processing all on board.

## Slide 3: Archive Process
Preserve Sample
Water Sample:
Concentrate Particulates
2019 MBARI, Confidential Information
Requirements
  Pressure range 0-500 psi
  Temperature 3C – 30C
  Multiple 25 mm filter porosities/types
    30 µm – 0.1 µm
  Chemicals: seawater, freshwater, 10% bleach, 10% HCl, silicone oil
  Deliver fixative across/through whole filter (No bubbles)

## Slide 4: Archive Cartridge -- Features
2019 MBARI, Confidential Information
Actuator Rod (magnetic)
Actuator Rod
Actuator Rods
Actuator Rod (magnetic)
Connects to 3G Core Sample Loop
Connection to 3G core Sample Loop
5 Actuator Rods Actuate
  4 Syringes
  1 Valve (sample or reagent)
4 Syringes
  Reagent/Air or waste
4) Not in use
Valve
1) Preservative/Air
2) Waste
3) Waste
Air Refill
Vent
Vent
1/8” EPDM ball w/ spring (10lb/in)
Filter
1/8” EPDM ball w/ spring (1.4lb/in)

## Slide 5: Archive – Water Filtration Concentrate Particulates
2019 MBARI, Confidential Information
Water Sample in
Filtered water out
Sampling (surface to 300m)
  Test psi range 0-500 psi
  Multiple 25 mm filter porosities/types
  Sample path under vacuum in 3G/LRAUV.
  Sample enters top of filter first before full fluid path through cartridge is opened.
When sampling is complete move valve to reagent position.
Filter
Valve Sample Position
Preservative (1.6ml)
Actuator Rod

## Slide 6: Archive – Preserve Sample
Preservative pushed in and displaces the water to waste. Preservative reaches >95% concentration.
Incubate
Refill Syringe 1 with air and displace preservative
2019 MBARI, Confidential Information
Valve Reagent Position
Preservative (1.6ml)
Actuator Rod (magnetic)
Air
waste
waste
Sample
Preserved

## Slide 7: Lyse-n-Go Process
Evacuate Water
Water Sample:
Concentrate Particulates
Add Lysis Buffer and Heat
Recover Extract
Off Cartridge analysis. May involve dilution with reagent(s).
2019 MBARI, Confidential Information
Requirements
  Pressure range 0-500 psi
  Temperature 3C – 30C
  Multiple 25 mm filter porosities/types
    30 µm – 0.1 µm
  Chemicals: seawater, freshwater, 10% bleach, 10% HCl, silicone oil
  Deliver lysis buffer and recover cell lysate.
  Perform dilution
  Deliver a bubble free extract off cartridge for analysis.

## Slide 8: Lyse-n-Go Cartridge -- Features
Connection to Sample Loop
5 Actuator Rods
4 Syringes
1 Valve (sample or reagent)
Make/Break cartridge connection to downstream components (waste, analytical module)
3x 3/32” EPDM Ball with spring
(10 lb/in)
Spacer with DuckBill Valve (FVMQ)
2) Lysis Buffer
/Extract
1) Not in use
3) Diluent
4) Air/Vent
3/32” EPDM
Ball w/ spring
(4lb/in)
Vent
Air Refill Vent
Filter
Valve (Sampling)
2019 MBARI, Confidential Information
Burkert Valve Tree
Make/Break Cartridge Connection w/ pressure sensor
Waste
Analytical Module
Blocked
Connects to 3G Core Sample Loop
p
Bubble trap/
Puck vent
Actuator Rod (magnetic)
Actuator Rod
Actuator Rods
Actuator Rod (magnetic)

## Slide 9: Lyse-n-Go Cartridge – Water Filtration Concentrate Particulates
Filter
Valve (Sampling)
Burkert Valve Tree
Waste
Analytical Module
Blocked
p
Actuator Rod
Sampling (surface to 300m)
  Test psi range 0-500 psi
  Multiple 25 mm filter porosities/types
  Sample path under vacuum in 3G/LRAUV.
  Sample enters top of filter first before full fluid path through cartridge is opened.
When sampling is complete engage handoff (open to waste) & move valve to reagent position.
Water Sample in
Filtered water out
1
2
2019 MBARI, Confidential Information
*Check valve sees full sea pressure

## Slide 10: Lyse-n-Go Cartridge – Evacuate Water
Burkert Valve Tree
Waste
Analytical Module
Blocked
p
Refillable air/vent syringe (4) used to push water to waste
  All other plungers in syringes backed up with actuator rods
To remove remainder of water, must break bubble point of the filter.
  Burkert tree at blocked
  Repeated cycles with air/vent syringe (4) until reach specified psi
Drain Water to waste
Vent top of puck
  Plunger below bottom vent & boss of plunger displaces check valve
Valve (Reagent Position)
3
5
4
2019 MBARI, Confidential Information
Vent
6
X
Vent
Actuator Rod (magnetic)
4) Air/Vent

## Slide 11: Lyse-n-Go Cartridge – Add Lysis Buffer & Heat
Burkert Valve Tree
Waste
Analytical Module
Blocked
p
2019 MBARI, Confidential Information
Vent
X
Vent
Actuator Rod (magnetic)
2) Lysis Buffer (1.4 ml)
Fill off cartridge line with lysis buffer (300 µl)
  Puck top vented, open to waste
Fill majority of filter chamber with lysis buffer (syringe 2)
  Puck top Vented (syringe 4)
  Burkert tree at blocked
Pressurize filter chamber & Heat
  Burkert tree at blocked
  Puck top sealed. Plunger moved above vent with pressure balanced move with Syringe 4 & Syringe 2.
  Drive syringe 2 to empty.
  No bubble over filter
  Heat to 85℃, 14 minutes
9
8
7
Heat

## Slide 12: Lyse-n-Go Cartridge – Recover Extract
Burkert Valve Tree
Waste
Analytical Module
Blocked
p
2019 MBARI, Confidential Information
Vent
X
Vent
Actuator Rod (magnetic)
4) Air/Vent
Recover 900 µl of Extract
  Pressure filter chamber with Air/Vent Syringe
  Draw back on actuator rod (2) when pressure exceeds threshold. Plunger moves due to pressure.
Relieve pressure in system
  Open to Waste
  Vent Puck Top
10
11
2) Lysate

## Slide 13: Lyse-n-Go Cartridge – Off Cartridge Analysis
Burkert Valve Tree
Waste
Analytical Module
Blocked
p
2019 MBARI, Confidential Information
Vent
X
Vent
Actuator Rod (magnetic)
2) Lysate
Syringes 2 & 3 push at different rates to “mix” reagents at appropriate dilution
  Initial mixture pushed to waste before pushing to analytical module
12
3) Diluent