## Slide 1
Cartridge and Filter Project:
Lyse-n-Go Design Overview
Date: 12-17-13
By: Accel Biotech Team
Mission
Our mission is to provide clients throughout the world with product development for medical, diagnostic, biodefense and biotech products. Accel’s skilled team of engineers, designers and technicians will work together to provide services from concept to development to manufacturing support

## Slide 2: Agenda
Requirements
Cartridge overview
Valve design
Clamp design

## Slide 3: MBARI Tech Specifications
ESP Platform in a Nose Cone for deployment
  11.125” x 24” (the total system) but scalable, modular
  30 degree deviations from normal
  14v, ?amp, ?watts ->need soon
  0-300m depths, 15psi-450psi plus pump pressure
  0 -50 C inactive, 2-40 C active
  3g acceleration inactive, 0.5g active (plus 1g rolls)
Interface Specifications -TBD
  Port Size – Screw on ports for testing phases
  Output Port Design/Specs
  Etc.
  PCB Could be the Cartridge Mount – Later Phases

## Slide 4: MBARI Tech Specifications
Genetic Marker Classification (algae, bacteria, viruses, invertebrate larvae, biotoxins) – Nucleic Acid Prep
  60 turret slots
  Save sample before and after lysis
  Expression and translation of specific genes
  mRNA, DNA
  2ml-2L of raw seawater @ rates of 1L/hr using 0.2 um filters
  Exclude particles larger than 0.5mm
  Variable filter (two types)
  300-500 ul lysate (30 minute processing)
  6 Reagents

## Slide 5: Lyse-n-Go Cartridge Reqs.
Venting of the Puck chamber
Air evac of puck
1 input Reagent
Up to 2 on board dilutions
Waste off board
Sample off board: straight and diluted

## Slide 6: Lyse-n-Go Layout
CAM Latch
Additional Valve
Extended Manifold Plate
Reworked syringe manifold plates
Evac & Vent
Reagent
Storage/Dilution
Cartridge extended by 10mm
Need lower actuator!
Interface plate?

## Slide 7: Puck Venting
Venting is part of plunger
Controlled with actuator
Possibly need to add spring washer on bottom to provide a soft stop for evac
Indent in wall allows air to pass over o-ring
Air Path
Boss depresses check valve

## Slide 8: Valving Matrix
Per our last meeting:
 | Puck |  |  |  | Cartridge |  |  | Actuators Used | 
Step | Over Filter | Under Filter | Vent | Pressure | Sea In | Sea Out | Off Board |  | Notes
1 | Sea In | Sea Out | Closed | 500 | Open | Open | Closed* | N/A | Filter Seawater
2 | Evac In | Off Board | Closed | 90 | Closed* | Closed* | Open | Air | Clear residual seawater
3 | Evac In | Off Board | Open | 15 | Closed* | Closed* | Closed* | Vent | Vent off pressure above filter
4 | Reagent In (1st half) | Off Board | Open | 15 | Closed* | Closed* | Closed* | Lysis, Vent | Begin adding reagent, vent displaced air
5 | Reagent In(2nd half) | On Board Storage | Closed | 25 | Closed* | Closed* | Closed* | Lysis, Vent, dil | Close vent and add remaining reagent
6 | ~ ~ ~ ~ ~ ~ ~ ~ ~ ~ ~Heat~ ~ ~ ~ ~ ~ ~ ~ ~ ~ ~ ~ ~ ~ ~ |  |  |  |  |  |  |  | Heat Regent in Puck
7 | Evac In | On Board Storage | Closed | 90 | Closed* | Closed* | Closed* | Air | Move sample to on-board storage to mix w/ additional reagents
8 | N/A | No Back Flow | N/A | N/A | Closed* | Closed* | Open | dil | Move sample off-board to analytical unit, etc.
 | = Changed position |  |  |  | *Only matters if fluid flows past port |  |  |  | 

## Slide 9: Valving-Step 1
Sea In
Sea Out
Filter Seawater

## Slide 10: Valving-Step 2
Off Board Waste
Evac Air In
Clear residual seawater

## Slide 11: Valving-Step 3
Pressure Out
Vent off pressure above filter

## Slide 12: Valving-Step 4
Pressure Out
Reagent In
1st half
Add reagent, vent off displaced air

## Slide 13: Valving-Step 5
Reagent In
2nd half
Close vent, add remaining reagent
Tube to puck volume:100ul

## Slide 14: Valving-Step 6
Heat!!!

## Slide 15: Valving-Step 7
Move sample to on board storage to mix with additional reagents
Should there be active mixing (pumping)?
Volume from Puck to syringe≈ 300µl

## Slide 16: Valving-Step 8
Move sample off board to analytical unit, etc.
Volume from syringe to output is≈ 200µl

## Slide 17: Puck Clamp
Cam style clamp
Quick release
One solid structure
Finger detent, etc. TBA
Dual contact points

## Slide 18: Lyse-n-Go
Additional items to address
  O-ring seal for puck
  Custom gaskets for manifold plates
  Decrease puck volume to ideal level
  Heater mounting/wire routing/contact points
  Interface with downstream port
  Evac syringe return mechanism
    Magnets?
    Spring?
    Socket style ball detent?