G3 Technology Shopping List.

|  | G3 Function | Requirements | Technologies | COTS |
| --- | --- | --- | --- | --- |
|  | Isolation | -Isolation of seawater path -Up to 500 psi, 1mm min port dim -Non-corroding -Small scale -Low power Size: less than c-battery, | Direct lift solenoid valves Custom body valves (thread into a SAE 4 port) Size: less than c-battery, Axial flow thru valves Turn or axial pull valves Multi port valves Automated actuators of above valves. Multiple bank valves integrated on a chip |  |
|  | Filter |  | 25mm diameter surface filtering media Similar area, but other geometry filters Filter integration in mount/seal/chip/cartridge |  |
|  | Filter Pumping | -75 psi (5 bar bubble point) -Variable speed -Max flow at least 4L/hr (@22psi) -Min flow 300mL/hr -Dosing accuracy <1% -Positive displacement (or flow measuring) -Unaffected by saltwater -Minimal shear damage to microbes. -Bidirectional (not required, but valuable) Size (with drive mech, but not motor) < soda can | Piston pumps Rotating piston (Similar to DWSM) Syringes (with or without flow valves) Gear pumps (low expectation) Peristaltic (low expectation) Progressive cavity (typically large frame) Diaphragm Others? |  |

|  | Lyse | -Up to 1mL of buffer + beads -Sample particles on filter | Bead Beating Heating Integrated heater/filter fit or holder Grinding Magnetic driven beating/grinding Acoustic/Sonic Automation of the above in closed volumes |  |
| --- | --- | --- | --- | --- |
|  | Sample purify | Purification of 500-1000uL of lysate RNA and DNA in single sample desired. | RNA and DNA extraction from one sample Magnetic bead Glass bead/fiber columns Non-centrifuge wash processes. Automated systems that perform the above. |  |
|  | Reagent storage | Bulk storage to 500mL Single use storage 50uL to 1mL | Blister packs Prefilled syringes Imbedded syringes Lyophilized Dry powder (beads) |  |
|  | Reagent dosing | Deliver 50-750uL doses of liquid reagent Deliver single shot reagent doses Deliver/dose powder (beads) | Compressed gas systems Linear or cam actuators of blisters/syringes Actuation methods for any of the above storages, such as Dissolving/melting pills or isolators. |  |
|  | Transfer to AM | 90+ ports in to 3 ports out. | Mega-multi port valves. (90+ ports) Chip based valve to make multiport system Low pressure, robotic open/seal fittings. |  |