## Sheet1
Span | 75
FS | 37.61 | 37.9861
Zero | 0 | 1.3164
Zero TC | 0.07 | 0.052654
Span TC | 0.02 | 1.01
T | 25
0 | 0 | 1.3164 | 2.625 | 2.625
1 | 0.50147 | 1.8228 | 3.635 | 2.635
2 | 1.0029 | 2.3293 | 4.645 | 2.645
3 | 1.5044 | 2.8358 | 5.655 | 2.655
4 | 2.0059 | 3.3423 | 6.665 | 2.665
5 | 2.5073 | 3.8488 | 7.675 | 2.675
6 | 3.0088 | 4.3552 | 8.685 | 2.685
7 | 3.5103 | 4.8617 | 9.695 | 2.695
8 | 4.0117 | 5.3682 | 10.705 | 2.705
9 | 4.5132 | 5.8747 | 11.715 | 2.715
10 | 5.0147 | 6.3812 | 12.725 | 2.725
11 | 5.5161 | 6.8876 | 13.735 | 2.735
12 | 6.0176 | 7.3941 | 14.745 | 2.745
13 | 6.5191 | 7.9006 | 15.755 | 2.755
14 | 7.0205 | 8.4071 | 16.765 | 2.765
15 | 7.522 | 8.9136 | 17.775 | 2.775
16 | 8.0235 | 9.4201 | 18.785 | 2.785
17 | 8.5249 | 9.9265 | 19.795 | 2.795
18 | 9.0264 | 10.433 | 20.805 | 2.805
19 | 9.5279 | 10.9395 | 21.815 | 2.815
20 | 10.0293 | 11.446 | 22.825 | 2.825
21 | 10.5308 | 11.9525 | 23.835 | 2.835
22 | 11.0323 | 12.4589 | 24.845 | 2.845
23 | 11.5337 | 12.9654 | 25.855 | 2.855
24 | 12.0352 | 13.4719 | 26.865 | 2.865
25 | 12.5367 | 13.9784 | 27.875 | 2.875
26 | 13.0381 | 14.4849 | 28.885 | 2.885
27 | 13.5396 | 14.9913 | 29.895 | 2.895
28 | 14.0411 | 15.4978 | 30.905 | 2.905
29 | 14.5425 | 16.0043 | 31.915 | 2.915
30 | 15.044 | 16.5108 | 32.925 | 2.925
31 | 15.5455 | 17.0173 | 33.935 | 2.935
32 | 16.0469 | 17.5238 | 34.945 | 2.945
33 | 16.5484 | 18.0302 | 35.955 | 2.955
34 | 17.0499 | 18.5367 | 36.965 | 2.965
35 | 17.5513 | 19.0432 | 37.975 | 2.975
36 | 18.0528 | 19.5497 | 38.985 | 2.985
37 | 18.5543 | 20.0562 | 39.995 | 2.995
38 | 19.0557 | 20.5626 | 41.005 | 3.005
39 | 19.5572 | 21.0691 | 42.015 | 3.015
40 | 20.0587 | 21.5756 | 43.025 | 3.025
41 | 20.5601 | 22.0821 | 44.035 | 3.035
42 | 21.0616 | 22.5886 | 45.045 | 3.045
43 | 21.5631 | 23.095 | 46.055 | 3.055
44 | 22.0645 | 23.6015 | 47.065 | 3.065
45 | 22.566 | 24.108 | 48.075 | 3.075
46 | 23.0675 | 24.6145 | 49.085 | 3.085
47 | 23.5689 | 25.121 | 50.095 | 3.095
48 | 24.0704 | 25.6275 | 51.105 | 3.105
49 | 24.5719 | 26.1339 | 52.115 | 3.115
50 | 25.0733 | 26.6404 | 53.125 | 3.125
51 | 25.5748 | 27.1469 | 54.135 | 3.135
52 | 26.0763 | 27.6534 | 55.145 | 3.145
53 | 26.5777 | 28.1599 | 56.155 | 3.155
54 | 27.0792 | 28.6663 | 57.165 | 3.165
55 | 27.5807 | 29.1728 | 58.175 | 3.175
56 | 28.0821 | 29.6793 | 59.185 | 3.185
57 | 28.5836 | 30.1858 | 60.195 | 3.195
58 | 29.0851 | 30.6923 | 61.205 | 3.205
59 | 29.5865 | 31.1987 | 62.215 | 3.215
60 | 30.088 | 31.7052 | 63.225 | 3.225
61 | 30.5895 | 32.2117 | 64.235 | 3.235
62 | 31.0909 | 32.7182 | 65.245 | 3.245
63 | 31.5924 | 33.2247 | 66.255 | 3.255
64 | 32.0939 | 33.7312 | 67.265 | 3.265
65 | 32.5953 | 34.2376 | 68.275 | 3.275
66 | 33.0968 | 34.7441 | 69.285 | 3.285
67 | 33.5983 | 35.2506 | 70.295 | 3.295
68 | 34.0997 | 35.7571 | 71.305 | 3.305
69 | 34.6012 | 36.2636 | 72.315 | 3.315
70 | 35.1027 | 36.77 | 73.325 | 3.325
71 | 35.6041 | 37.2765 | 74.335 | 3.335
72 | 36.1056 | 37.783 | 75.345 | 3.345
73 | 36.6071 | 38.2895 | 76.355 | 3.355
74 | 37.1085 | 38.796 | 77.365 | 3.365
75 | 37.61 | 39.3025 | 78.375 | 3.375

## Sheet2
Warm up
10v excititaion
7:12 | 22.7
7:20 | 22.84
7:34 | 23.23
7:47 | 23.53
7:58 | 23.77
8:21 | 24.06
8:32 | 24.13
8:51 | 24.24
9:06 | 24.28
9:21 | 24.33
10:08 | 24.35
10:24 | 24.41
11:13 | 24.5
11:39 | 24.58
11:54 | 24.55

## Sheet3
Bias | 1.3
Vfs | 0.03761
Voff | -0.03 |  | FS+ | Zero | FS-
vin+ | 0 |  | 1.3188 | 1.3 | 1.2812
vin- | 0.03 |  | 1.2512 | 1.27 | 1.2888
rg | 6206.5267
vr |  |  |  | 0.38328
Gain | 37.2241
Calc off | -0.28328
Vo | -1.1167 |  | 2.9 | 1.5 | 0.1
Vo2 | 0.75834 |  | 1.8208 | 1.9958 | 2.1708
Vo2 | Vin+
Vin- | -0.06 | -0.05 | -0.04 | -0.03 | -0.02 | -0.01 | 0 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06
-0.06 | 0.525 | 0.48472 | 0.44444 | 0.40416 | 0.36388 | 0.3236 | 0.28332 | 0.24304 | 0.20276 | 0.16248 | 0.1222 | 0.081918 | 0.041638
-0.05 | 0.57778 | 0.5375 | 0.49722 | 0.45694 | 0.41666 | 0.37638 | 0.3361 | 0.29582 | 0.25554 | 0.21526 | 0.17498 | 0.1347 | 0.094418
-0.04 | 0.63056 | 0.59028 | 0.55 | 0.50972 | 0.46944 | 0.42916 | 0.38888 | 0.3486 | 0.30832 | 0.26804 | 0.22776 | 0.18748 | 0.1472
-0.03 | 0.68334 | 0.64306 | 0.60278 | 0.5625 | 0.52222 | 0.48194 | 0.44166 | 0.40138 | 0.3611 | 0.32082 | 0.28054 | 0.24026 | 0.19998
-0.02 | 0.73612 | 0.69584 | 0.65556 | 0.61528 | 0.575 | 0.53472 | 0.49444 | 0.45416 | 0.41388 | 0.3736 | 0.33332 | 0.29304 | 0.25276
-0.01 | 0.7889 | 0.74862 | 0.70834 | 0.66806 | 0.62778 | 0.5875 | 0.54722 | 0.50694 | 0.46666 | 0.42638 | 0.3861 | 0.34582 | 0.30554
0 | 0.84168 | 0.8014 | 0.76112 | 0.72084 | 0.68056 | 0.64028 | 0.6 | 0.55972 | 0.51944 | 0.47916 | 0.43888 | 0.3986 | 0.35832
0.01 | 0.89446 | 0.85418 | 0.8139 | 0.77362 | 0.73334 | 0.69306 | 0.65278 | 0.6125 | 0.57222 | 0.53194 | 0.49166 | 0.45138 | 0.4111
0.02 | 0.94724 | 0.90696 | 0.86668 | 0.8264 | 0.78612 | 0.74584 | 0.70556 | 0.66528 | 0.625 | 0.58472 | 0.54444 | 0.50416 | 0.46388
0.03 | 1 | 0.95974 | 0.91946 | 0.87918 | 0.8389 | 0.79862 | 0.75834 | 0.71806 | 0.67778 | 0.6375 | 0.59722 | 0.55694 | 0.51666
0.04 | 1.0528 | 1.0125 | 0.97224 | 0.93196 | 0.89168 | 0.8514 | 0.81112 | 0.77084 | 0.73056 | 0.69028 | 0.65 | 0.60972 | 0.56944
0.05 | 1.1056 | 1.0653 | 1.025 | 0.98474 | 0.94446 | 0.90418 | 0.8639 | 0.82362 | 0.78334 | 0.74306 | 0.70278 | 0.6625 | 0.62222
0.06 | 1.1584 | 1.1181 | 1.0778 | 1.0375 | 0.99724 | 0.95696 | 0.91668 | 0.8764 | 0.83612 | 0.79584 | 0.75556 | 0.71528 | 0.675
Vo2 | Vin+
Vin- | -0.06 | -0.05 | -0.04 | -0.03 | -0.02 | -0.01 | 0 | 0.01 | 0.02 | 0.03 | 0.04 | 0.05 | 0.06
-0.06 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334 | 2.6057 | 2.9779 | 3.3502 | 3.7224 | 4.0947 | 4.4669
-0.05 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334 | 2.6057 | 2.9779 | 3.3502 | 3.7224 | 4.0947
-0.04 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334 | 2.6057 | 2.9779 | 3.3502 | 3.7224
-0.03 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334 | 2.6057 | 2.9779 | 3.3502
-0.02 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334 | 2.6057 | 2.9779
-0.01 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334 | 2.6057
0 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612 | 2.2334
0.01 | -2.6057 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489 | 1.8612
0.02 | -2.9779 | -2.6057 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167 | 1.489
0.03 | -3.3502 | -2.9779 | -2.6057 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448 | 1.1167
0.04 | -3.7224 | -3.3502 | -2.9779 | -2.6057 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224 | 0.74448
0.05 | -4.0947 | -3.7224 | -3.3502 | -2.9779 | -2.6057 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0 | 0.37224
0.06 | -4.4669 | -4.0947 | -3.7224 | -3.3502 | -2.9779 | -2.6057 | -2.2334 | -1.8612 | -1.489 | -1.1167 | -0.74448 | -0.37224 | 0

## Sheet4
Full Scale P | 75
Sensor Common Ratio | 0.5
Sensor Vfs | 0.03
Sensor ZeroP Voff | 0
Excitation V | 10
Actual Excitation | 3.3
Sensor Vfs | 0.0099
Sensor ZeroP Voff | 0
Vout at ZeroP | 0.65
Vout at MaxP | 3.15
Vout at MinP | 0.15
P at MaxP | 75
Pat MinP | -15
DeltaV | 3
DelatP | 90
Scale V/P | 0.033333
Sesnor V/P | 0.000132
Gain | 252.5253
Offset | 0.65
Rg | 807.9984
| Vout | Vout2
At Max 75psi | 3.15
At -15psi | 0.15
At 0psi | 0.65
Sensor CMV | 1.65
Sensor V+ at 75 | 1.6549 | 2.2125 | 3.15
Sensor V- at 75 | 1.645
Sensor V+ at -15 | 1.649 | 2.5875 | 0.15
sensor V- at -15 | 1.651

## Sheet5
PCB | [5 empty] | Sensor
22 |  | 5 |  |  |  | 5
70.8 | 2.88 | 75 | 3.024 | 75.4374 | -0.43738 | 72.1 | 2.99 | 74.3548 | -2.2548 | 75.2 | 3.103 | 77.9527 | -2.7527
17.2 | 1.188 | 28.2 | 1.534 | 27.9961 | 0.20394 | 36.1 | 1.872 | 38.7579 | -2.6579 | 0 | 0.716 | 1.9511 | -1.9511
0 | 0.66 | 0 | 0.664 | 0.29543 | -0.29543 | 0 | 0.728 | 2.3332 | -2.3332
Slope | 31.8398 |  | 31.7396 |  |  |  | 31.8732 |  |  |  | 31.504
Intercept | -20.8462 |  | -20.8481 |  |  |  | -23.3238 |  |  |  | -22.5568

## Sheet6
Vfs | 0.1
Vos | 0.02
Ri | 10000
Gf | 10
Vfs33 | 0.033
Vos33 | 0.0066
Gt | 90.9091 |  | 180 | 75.7576
Gd | 9.0909 |  | 18 | 7.5758
Rg | 48888.8889 |  | 15384.6154 | 77647.0588
Voff | 1.458 | 1.603
Vdoh | 3.15
Vdol | 0.15
V0 |  |  | 0.1 | 0.1
Vref |  |  | 1.513 | 1.513 | 1.518
Vp0 |  |  | 1.288 | 0.6 | 0.65
Vp75 |  |  | 7.228 |  | 3.15
Ga |  |  | 0.42088
Vnewp0 |  |  | 0.6
VnewRef

## Sheet7
Vfs | 0.15 | [8 empty] | Vfsnom | 0.1
Vos | 0.02 | [8 empty] | Pfs (bar) | 20
Ri | 10000 | [8 empty] | Pfs | 290.0754
Gf | 2.5 | [8 empty] | V%+ | 0.5 | 1.3336
Pfs | 507.6 | (35 bar abs) |  | bar | PSI |  |  |  |  | V%- | -0.3
Pmax | 600 |  |  | 20 | 290.0754 |  |  |  |  | Vos | -0.015
[10 empty] | Vref | 3
Vfs33 | 0.058511 |  |  |  |  | Rg | 0.128 | 0.1 | 0.072 | Roff | 10000
Vos33 | 0.0066 | [5 empty] | 10507 | 7763 | 5301 | Rg | 10000
[10 empty] | Vhigh | 3.15
Gt | 51.2727 |  | 180 | 42.7273 | [5 empty] | Vlow | 0.15
Gd | 20.5091 |  | 72 | 17.0909 | [5 empty] | Phigh | 290
[10 empty] | Plow | -290
Rg | 12895.6624 |  | 2985.0746 | 16541.3534 | [5 empty] | Pfs | 507.6
Voff | 0.92664 | 1.2974 | [7 empty] | Gf | 2.5
Vdoh | 3.15
Vdol | 0.15 |  |  |  |  | Voff | 0.128 | 0.1 | 0.072 | X | 0.18853
[6 empty] | 0.02 | 0.903 | 0.859 | 0.78 | Y | 63.6489
V0 |  |  | 0.1 | 0.1 |  | -0.02 | 1.22 | 1.265 | 1.34 | a | -25
Vref |  |  | 1.042 | 1.042 | 1.1492 |  |  |  |  | b | -2545.9561
Vp0 |  |  | 1.288 | 0.382 | 0.65 |  |  |  |  | c | 3439.6938 | Sensor FS error allowable with given Rg
Vp75 |  |  | 11.8199 |  | 3.15 |  |  |  |  | E | 1.3336
Ga |  |  | 0.23737 | [7 empty] | 0.07
0.15
Vnewp0 |  |  | 0.382 | [6 empty] | Vfs+ | 0.058511
VnewRef | [9 empty] | Vfs- | 0.027305
[10 empty] | Vos | -0.00495
From Danny Uehara at TE
The sensitivity tolerance for our non-amplified sensor is -30%/+50%
[10 empty] | Gt+ | 51.2727
[10 empty] | Gt- | 109.8701
[10 empty] | Gd+ | 20.5091
[10 empty] | Gd- | 43.9481
[10 empty] | Rg+ | 12895.6624
[10 empty] | Rg- | 5135.045
[10 empty] | Delta R | 7760.6174
[10 empty] | Vos+ | 1.2795
[10 empty] | Vos- | 0.84446
[10 empty] | Delta Vos | 0.43509
[10 empty] | Ios | 4.3509e-05
[10 empty] | Rlow | 19408.9835
[10 empty] | Rhigh | 39542.9472

## CalibrationOld
Steps
Amplifier Configuration | High | Jumpers | 1 | Selecting amplifier configuration
Low Vout @ Low Vi | 0.7742 | Low Vi 1-2, 5-6, 7-8 | 2 | Disconnect Sensor, Power with 3.3V (1-8), Measure Vo (4-8)
High Vout @ High Vi | 1.3446 | High Vi 1-2, 3-5, 7-8 | 3 | Install jumpers for Low Vi, Measure and record Vout Low
Vo @ P = Pmin | 0.2854 |  | 4 | Install jumpers for High Vi, Measure and record Vout High
Pressure | 462.7 |  | 5 | Remove Jumpers, Install Sensor, Record Vo at Pmin
Vo @ Pressure | 2.4281 |  | 6 | Pressurize to Pressure, Record Vo
|  |  | 7 | While at Pressure Adjust Offset to obtain Vo
@Pressure Adjust O for Vo | 2.4485 | Volts | 8 | While at pressure Adjust Gain to obtain Vo
@Pressure Adjust G for Vo | 2.9328 | Volts
|  | Low Pressure Settings | High Pressure Settings
|  | Low | High
Jumpers
LowVi @ Low Vout | 0.01 | 0.005 | 0.01 | Vi with jumper setting
High Vi @ High Vout | 0.02 | 0.015 | 0.02 | Vi with jumper setting
Gain final stage | 2.5 | 10 | 2.5 | Final gain stage value
Pmin | 14.7 | 0 | 14.7 | pressure to use as low point calibration (ambient)
Pmax | 500 | 75 | 500 | pressure to use as high point claibration
Pmin | 14.7 | 0 | 14.7 | I think this is not used
Vmax | 3.15 | 3.15 | 3.15 | ideal Vout at maximum pressure
Vmin | 0.238 | 0.65 | 0.238 | ideal Vout at ambient pressure
Output stage offest | 1.67 | 1.67 | 1.67
Gain Total | 57.04
Gain Diff | 22.816
Vref | 1.0835
Vi @ Pmin | 0.0014306
Vi @ Pressure | 0.038995
Vsens | 8.385e-05
Vfs | 0.041925
Vos | 0.00019798
Offset | 0.0014306
0.038995
Vi @ Pmax | 0.040593
ideal Vo @Pressure | 2.9328
Gain total final | 69.4572
Vo @ P=Pmin | 0.32361
Vref need to be @ P=14.7 | 1.0917
| 7.0135 | New offset
| 2.5853 | Adjusted V at Pmin
| 39 | psi
| 2.595 | V at psi
| 0.00016999 | delta V
| 4.3586e-06 | V/psi
| 0.0021152 | V full scale
| 550.6715 | Gain
| -4.8346 | final offset
| -27.0253 | Adjust O until Vo at psi
| -24.9131 | Adjust G until Vo at psi

## Sheet8
Full Scale Pabs | 507
Sensor Common Ratio | 0.5
Sensor Vfs | 0.08
Sensor ZeroP Voff | 0.0047
Excitation V | 10
Actual Excitation | 3.3
Sensor Vfs | 0.0264
Sensor ZeroP Voff | 0.001551
Vout at Atm | 0.2235
Vout at MaxP | 3.15
Vout at MinP | 0.15
P at MaxPabs | 600
Pat MinPabs | 0
DeltaV | 3
DelatP | 600
Scale V/P | 0.005
Sesnor V/P | 5.2071e-05
Gain | 96.0227
Offset | 0.0010688
Rg | 2197.2534
| Vout | Vout2
At Max 75psi | 0.525
At -15psi | 0.075
At 0psi | 0.15
Sensor CMV | 1.65
Sensor V+ at Pmax | 1.6664 | 2.2936 | 3.15
Sensor V- at Pmax | 1.6336
Sensor V+ at Pmin | 1.6508 | 2.6686 | 0.15
sensor V- at Pmin | 1.6492

## Calibration
Steps
Amplifier Configuration | Sample | Jumpers | 1 |  |  | Selecting amplifier configuration
Low Vout @ Low Vi | 1.2236 | Low Vi 1-2, 5-6, 7-8 (1,3,4) | 2 |  |  | Disconnect Sensor, Power with 3.3V (1-8), Measure Vo (4-8)
High Vout @ High Vi | 1.6703 | High Vi 1-2, 3-5, 7-8 (1,2,4) | 3 |  |  | Install jumpers for Low Vi, Measure and record Vout Low
Pmin cal | 0 |  | 4 |  |  | Install jumpers for High Vi, Measure and record Vout High
Vo @ Pmin cal | 0.8456 | z | 5 |  |  | Remove Jumpers, Install Sensor, Record Vo at Pmin
Pmax cal | 199 |  | 6 |  |  | Pressurize to Pressure, Record Vo
Vo @ Pmax cal | 1.3536 |  | 7 |  |  | While at Pressure Adjust Offset to obtain Vo
|  |  | 8 |  |  | While at pressure Adjust Gain to obtain Vo
Vfs actual | 82.86 | mV
Vos actual | 6.99 | mV
@Pressure Adjust O for Vo | 2.0875 | Volts
@Pressure Adjust G for Vo | 2.6793 | Volts
|  | Low Pressure Settings | High Pressure Settings | Super High Pressure Settings | Sample Pressure Setting
|  | Low | High | Super High | Sample
Sensor | 20 Bar Gauge | 5 Bar Gauge | 35 Bar Absolute | 35 Bar Absolute | 20 Bar Gauge
LowVi @ Low Vout | 0.015 | 0.005 | 0.01 | 0.015 | 0.015 | Vi with jumper setting
High Vi @ High Vout | 0.03 | 0.015 | 0.02 | 0.03 | 0.03 | Vi with jumper setting
Vext | 3 | 3.3 | 3.3 | 3 | 3 | Sensor Excitation Voltage
Gain final stage | 2.5 | 10 | 2.5 | 2.5 | 2.5 | Final gain stage value
Vfs max | 150 | 45 | 150 | 150 | 150 | Maximum sensitivity (mV at FS)
Vfs min | 70 | 21 | 70 | 70 | 70 | Minimum sensitivity (mV at FS)
Vos Max | 100 | 30 | 100 | 100 | 100 | Vos Maximum (mV)
Vos Min | -100 | -30 | -100 | -100 | -100 | Vos Minimum (mV)
Pmin | 0 | 0 | 14.7 | 15 | 0 | pressure to use as low point calibration (ambient)
Pfs | 290 | 75 | 508 | 508 | 290 | sensor full scale rating
Pmax | 290 | 75 | 600 | 600 | 290 | pressure to use as high point of calibration
Pmin | -290 | 0 | 14.7 | 0 | -290 | pressure at the lowest point of calibration
Vmax | 3.15 | 3.15 | 3.15 | 3 | 3.15 | ideal Vout at maximum pressure
Vmin | 0.15 | 0.65 | 0.2235 | 0.3 | 0.15 | ideal Vout at ambient pressure
Output stage offest | 1.67 | 1.67 | 1.67 | 1.67 | 1.67 | Output stage offestimum
Gain Total | 29.78 | current state
Gain Diff | 11.912 | of |  | 0.2235
Vref | 1.3128 | amplifier
Vi @ Pmin cal | 0.0023069 | input voltage at pressure |  | #DIV/0! |  |  | -0.01 | 0.3
Vi @ Pmax cal | 0.019365 | input voltage at pressure |  |  |  |  | 0.086 | 3
Vsens | 8.5721e-05 | sensor actual sensitivity (V/P)
Vfs | 0.082863 | sensor actual FS @10V |  |  |  |  | 28.125 | 0.58125
Vos | 0.0069907 | sensor actual Vos @10V
Offset | 0.0023069 | input voltage at pressure
| 0.019365 | input voltage at pressure
Ideal M | 0.0051724 |  | Vi @ Pmax | 0.024859 | input V @ Pmax
Ideal B | 1.65
Ideal Vo @ Pmax | 2.6793
Ideal Vo @ Pmin | 1.65
Gain total final | 60.3403
Offset Final | 1.5108
Ideal Vref | 1.6063
Vo @ P=Pmin | 0.15
Vref need to be @ Pmin cal | 1.0063
| 7.1266 | New offset
| 2.8552 | Adjusted V at Pmin
| 39 | psi
| 2.595 | V at psi
| -0.0087367 | delta V
| -0.00022402 | V/psi
| -0.064965 | V full scale
| -18.4715 | Gain
| 6.439 | final offset
| 0.87609 | Adjust O until Vo at psi
| 1.3645 | Adjust G until Vo at psi

## Design Calc
Position | Sample | Leak | Handoff
Sensor range | 20 bar | 35 bar | 5 bar
Part number | XPM4-20BG-/Z02 | XPM4-35BA-/HA/SI/ZI/ET1 | XPM4-5BG-/ET1
Vfsnom | 0.1 | 0.1 | 0.03 | Sensor full scale output V
Vext | 10 | 10 | 10 | Sensor specified excitation V
Pfs (bar) | 20 | 35 | 5 | Pressure sensor full scale pressure
Pfs | 290.0754 | 507.6319 | 72.5189 | converted to psi
V%+ | 0.5 | 0.5 | 0.5 | vfs + tolerence across devices
V%- | -0.3 | -0.3 | -0.3 | vfs - tolerance across devices
Vos | 0.015 | 0.015 | 0.015 | sensor offset
Vaext | 3 | 3 | 3 | Amplifier excitation voltage
Vref | 3 | 3 | 3 | amplifier offset reference
Vhigh | 3.15 | 3.15 | 3.15 | Amplifier output voltage
Vlow | 0.15 | 0.15 | 0.15
Phigh | 290 | 600 | 75 | highest pressure
Plow | -290 | 0 | -15 | lowest pressure
Gf | 2.5 | 2.5 | 11 | Final amplifier gain
Rg mag | 100000 | 100000 | 10000 | magnitude of gain resistors
Roff | 10000 | 10000 | 10000 | offest adjust potentiometer
R7 | 13300 | 13300 | 13300 | Two calibration resistor values
R17 | 16500 | 16500 | 16500
Vcal low | 0.015 | 0.015 | 0.0045 | two calibration voltages
Vcal high | 0.03 | 0.03 | 0.009
R1 (E96) | 8450 | 9530 | 6650 | Minimum Gain Resistor
R2 | 20000 | 10000 | 20000 | Gain Trim pot
R5 | 100000 | 100000 | 10000 | Final stage input resistor
R6 (E96) | 14300 | 2320 | 10000 | Low Vref Resistor
R10 | 150000 | 150000 | 100000 | Final stage feedback resistor
R11 (E96) | 150 | 150 | 45.3
R12 (E96) | 150 | 150 | 45.3
R16 (E96) | 10500 | 8450 | 8660 | High Vref Resistor
R1 | 8482.1753 | 9614.0465 | 6693.0816 | Minimum Gain Resistor
R6 | 14384.9599 | 2322.2342 | 10072.6328 | Low Vref Resistor
R16 | 10605.9425 | 8362.0431 | 8595.7754 | High Vref Resistor
R11 | 150.5051 | 150.5051 | 44.8345
R12 | 150.5051 | 150.5051 | 44.8345
Vref high | 2.0907 | 1.7872 | 2.1005 | Highest Vref
Vref low | 1.2333 | 0.33681 | 1.054 | Lowest Vref
Vfs+ | 0.045 | 0.045 | 0.0135
Vfsnom | 0.03 | 0.03 | 0.009 | Nominal Vfs at amplifier excitation
Vfs- | 0.021 | 0.021 | 0.0063
V/psi + | 0.00015513 | 8.8647e-05 | 0.00018616
V/psi nom | 0.00010342 | 5.9098e-05 | 0.00012411 | Volts/PSI nominal
V/psi - | 7.2395e-05 | 4.1369e-05 | 8.6874e-05
| 0.089977 | 0.053188 | 0.016754
V range nom | 0.059984 | 0.035459 | 0.01117 | V range from sensor across Phigh-Plow
| 0.041989 | 0.024821 | 0.0078187
+ | 33.342 | 56.4035 | 179.0589
Gt nom | 50.013 | 84.6053 | 268.5883 | Total gain to achieve output Vhigh-Vlow
- | 71.4471 | 120.8647 | 383.6976
| 13.3368 | 22.5614 | 16.2781
Gdiff | 20.0052 | 33.8421 | 24.4171 | difference amplifer gain based upon Gf
| 28.5789 | 48.3459 | 34.8816
+ | 23990.0202 | 11388.6007 | 17733.5137
Rg | 13328.7127 | 6934.3006 | 10300.1881
- | 8482.1753 | 4614.0465 | 6693.0816
dRg | 15507.8448 | 6774.5542 | 11040.432 | R2 range
dV low + +off | -0.029988 | 0.015 | 0.012208 | sensor voltage Plow, Vfs+ and +Voff
dV low nom +off | -0.014992 | 0.015 | 0.013138 | sensor voltage Plow, Vfs nom and +Voff
dV low - +off | -0.0059945 | 0.015 | 0.013697 | sensor voltage Plow, Vfs- and +Voff
dV low+ | -0.044988 | 0 | -0.0027924 | sensor voltage Plow, Vfs+ and 0 off
dV low nom | -0.029992 | 0 | -0.0018616 | sensor voltage Plow, Vfs nom and 0 off
dV low- | -0.020995 | 0 | -0.0013031 | sensor voltage Plow, Vfs- and 0 off
dV low + -off | -0.059988 | -0.015 | -0.017792 | sensor voltage Plow, Vfs+ and -Voff
dV low nom -off | -0.044992 | -0.015 | -0.016862 | sensor voltage Plow, Vfs nom and -Voff
dV low - -off | -0.035995 | -0.015 | -0.016303 | sensor voltage Plow, Vfs- and -Voff
Vref low + +off | 1.4619 | 0.72358 | 1.3331 | Vref at Plow, Vfs+ and +Voff
Vref low nom +off | 1.3619 | 0.55437 | 1.211 | Vref at Plow, Vfs nom and +Voff
Vref low - +off | 1.2333 | 0.33681 | 1.054 | Vref at Plow, Vfs- and +Voff
Vref low + | 1.662 | 1.062 | 1.5773 | Vref at Plow, Vfs+ and 0 off
Vref low nom | 1.662 | 1.062 | 1.5773 | Vref at Plow, Vfs nom and 0 off
Vref low - | 1.662 | 1.062 | 1.5773 | Vref at Plow, Vfs- and 0 off
Vref low + -off | 1.8621 | 1.4004 | 1.8214 | Vref at Plow, Vfs+ and -Voff
Vref low nom -off | 1.9621 | 1.5696 | 1.9435 | Vref at Plow, Vfs nom and -Voff
Vref low - -off | 2.0907 | 1.7872 | 2.1005 | Vref at Plow, Vfs- and -Voff
dVref | 0.85737 | 1.4504 | 1.0464
Iref | 8.5737e-05 | 0.00014504 | 0.00010464

## E96 calc
Gf | 2.5 | min error | 0 | Gf | 2.5 | min error | 0 | Gf | 11 | min error | 0
Ratio | 1.5 | R5 | 0 | Ratio | 1.5 | R5 | 0 | Ratio | 10 | R5 | 0
A-B | 17 | R10 | 17 | A-B | 17 | R10 | 17 | A-B | 96 | R10 | 96
| B | A |  |  | B | A |  |  | B | A
0 | 0 | 17 | 1.5 | 0 | 0 | 17 | 1.5 | 0 | 0 | 96 | 10
0.0065359 | 1 | 18 | 1.5098 | 0.0065359 | 1 | 18 | 1.5098 | 0.0039216 | 1 | 97 | 10.0392
0.0031746 | 2 | 19 | 1.5048 | 0.0031746 | 2 | 19 | 1.5048 | 0.00095238 | 2 | 98 | 9.9905
0.0093458 | 3 | 20 | 1.514 | 0.0093458 | 3 | 20 | 1.514 | 0.0046729 | 3 | 99 | 10.0467
1.4803e-16 | 4 | 21 | 1.5 | 1.4803e-16 | 4 | 21 | 1.5 | 0.00090909 | 4 | 100 | 10.0091
0.0029499 | 5 | 22 | 1.4956 | 0.0029499 | 5 | 22 | 1.4956 | 0.0026549 | 5 | 101 | 9.9735
0.0086957 | 6 | 23 | 1.513 | 0.0086957 | 6 | 23 | 1.513 | 0.0043478 | 6 | 102 | 10.0435
0.0056497 | 7 | 24 | 1.5085 | 0.0056497 | 7 | 24 | 1.5085 | 0.0025424 | 7 | 103 | 10.0254
0.0027548 | 8 | 25 | 1.5041 | 0.0027548 | 8 | 25 | 1.5041 | 0.0016529 | 8 | 104 | 10.0165
0.0053763 | 9 | 26 | 1.5081 | 0.0053763 | 9 | 26 | 1.5081 | 0.00080645 | 9 | 105 | 10.0081
0.0026247 | 10 | 27 | 1.5039 | 0.0026247 | 10 | 27 | 1.5039 | 0.0007874 | 10 | 106 | 10.0079
0.0051282 | 11 | 28 | 1.5077 | 0.0051282 | 11 | 28 | 1.5077 | 0.0015385 | 11 | 107 | 10.0154
0.0025063 | 12 | 29 | 1.5038 | 0.0025063 | 12 | 29 | 1.5038 | 0.0030075 | 12 | 108 | 10.0301
0.0024331 | 13 | 30 | 1.4964 | 0.0024331 | 13 | 30 | 1.4964 | 0.0029197 | 13 | 109 | 9.9708
1.4803e-16 | 14 | 31 | 1.5 | 1.4803e-16 | 14 | 31 | 1.5 | 0.00071429 | 14 | 110 | 9.9929
0.002331 | 15 | 32 | 1.5035 | 0.002331 | 15 | 32 | 1.5035 | 0.0020979 | 15 | 111 | 10.021
0.0022676 | 16 | 33 | 1.5034 | 0.0022676 | 16 | 33 | 1.5034 | 0.0013605 | 16 | 112 | 9.9864
0.0044444 | 17 | 34 | 1.5067 | 0.0044444 | 17 | 34 | 1.5067 | 0.002 | 17 | 113 | 10.02
0.004329 | 18 | 35 | 1.5065 | 0.004329 | 18 | 35 | 1.5065 | 0 | 18 | 114 | 10
0 | 19 | 36 | 1.5 | 0 | 19 | 36 | 1.5 | 0.0018987 | 19 | 115 | 9.981
0 | 20 | 37 | 1.5 | 0 | 20 | 37 | 1.5 | 0.0024691 | 20 | 116 | 9.9753
0.0060606 | 21 | 38 | 1.5091 | 0.0060606 | 21 | 38 | 1.5091 | 0.0030303 | 21 | 117 | 10.0303
0.0059172 | 22 | 39 | 1.5089 | 0.0059172 | 22 | 39 | 1.5089 | 0.0029586 | 22 | 118 | 10.0296
0 | 23 | 40 | 1.5 | 0 | 23 | 40 | 1.5 | 0.0022989 | 23 | 119 | 9.977
0 | 24 | 41 | 1.5 | 0 | 24 | 41 | 1.5 | 0.0011236 | 24 | 120 | 9.9888
0.003663 | 25 | 42 | 1.5055 | 0.003663 | 25 | 42 | 1.5055 | 0.00054945 | 25 | 121 | 10.0055
0.0017825 | 26 | 43 | 1.4973 | 0.0017825 | 26 | 43 | 1.4973 | 0.002139 | 26 | 122 | 9.9786
0.0017452 | 27 | 44 | 1.5026 | 0.0017452 | 27 | 44 | 1.5026 | 0.00052356 | 27 | 123 | 10.0052
0 | 28 | 45 | 1.5 | 0 | 28 | 45 | 1.5 | 0.0015306 | 28 | 124 | 9.9847
0.0033333 | 29 | 46 | 1.505 | 0.0033333 | 29 | 46 | 1.505 | 0.0025 | 29 | 125 | 10.025
0.004878 | 30 | 47 | 1.5073 | 0.004878 | 30 | 47 | 1.5073 | 0.0019512 | 30 | 126 | 10.0195
0.0031746 | 31 | 48 | 1.5048 | 0.0031746 | 31 | 48 | 1.5048 | 0.0014286 | 31 | 127 | 10.0143
0.0046512 | 32 | 49 | 1.507 | 0.0046512 | 32 | 49 | 1.507 | 0.0018605 | 32 | 128 | 10.0186
0.0015083 | 33 | 50 | 1.5023 | 0.0015083 | 33 | 50 | 1.5023 | 0.0013575 | 33 | 129 | 9.9864
0.0029499 | 34 | 51 | 1.5044 | 0.0029499 | 34 | 51 | 1.5044 | 1.7764e-16 | 34 | 130 | 10
0 | 35 | 52 | 1.5 | 0 | 35 | 52 | 1.5 | 0.0021552 | 35 | 131 | 9.9784
0.0042194 | 36 | 53 | 1.5063 | 0.0042194 | 36 | 53 | 1.5063 | 0.00042194 | 36 | 132 | 10.0042
0.0013717 | 37 | 54 | 1.5021 | 0.0013717 | 37 | 54 | 1.5021 | 0.00041152 | 37 | 133 | 9.9959
0.0013387 | 38 | 55 | 1.502 | 0.0013387 | 38 | 55 | 1.502 | 0.00080321 | 38 | 134 | 9.992
0.0013072 | 39 | 56 | 1.502 | 0.0013072 | 39 | 56 | 1.502 | 0.00078431 | 39 | 135 | 9.9922
0.0012771 | 40 | 57 | 1.5019 | 0.0012771 | 40 | 57 | 1.5019 | 1.7764e-16 | 40 | 136 | 10
0.0037453 | 41 | 58 | 1.5056 | 0.0037453 | 41 | 58 | 1.5056 | 0.0014981 | 41 | 137 | 10.015
0.0024331 | 42 | 59 | 1.5036 | 0.0024331 | 42 | 59 | 1.5036 | 0.00072993 | 42 | 138 | 9.9927
0.0047619 | 43 | 60 | 1.5071 | 0.0047619 | 43 | 60 | 1.5071 | 0.0017857 | 43 | 139 | 10.0179
0.0034843 | 44 | 61 | 1.5052 | 0.0034843 | 44 | 61 | 1.5052 | 0.0010453 | 44 | 140 | 10.0105
0.0022676 | 45 | 62 | 1.5034 | 0.0022676 | 45 | 62 | 1.5034 | 0.0010204 | 45 | 141 | 10.0102
0.0033223 | 46 | 63 | 1.505 | 0.0033223 | 46 | 63 | 1.505 | 0.0013289 | 46 | 142 | 10.0133
0.0010787 | 47 | 64 | 1.5016 | 0.0010787 | 47 | 64 | 1.5016 | 0.00097087 | 47 | 143 | 9.9903
0.0021097 | 48 | 65 | 1.5032 | 0.0021097 | 48 | 65 | 1.5032 | 0.00063291 | 48 | 144 | 10.0063
0.0020576 | 49 | 66 | 1.5031 | 0.0020576 | 49 | 66 | 1.5031 | 0.00030864 | 49 | 145 | 9.9969
0.002008 | 50 | 67 | 1.503 | 0.002008 | 50 | 67 | 1.503 | 0.00060241 | 50 | 146 | 9.994
0.0019608 | 51 | 68 | 1.5029 | 0.0019608 | 51 | 68 | 1.5029 | 0.00058824 | 51 | 147 | 9.9941
0.0019157 | 52 | 69 | 1.5029 | 0.0019157 | 52 | 69 | 1.5029 | 0.00028736 | 52 | 148 | 10.0029
0.00093371 | 53 | 70 | 1.5014 | 0.00093371 | 53 | 70 | 1.5014 | 0.0014006 | 53 | 149 | 9.986
0.0027397 | 54 | 71 | 1.5041 | 0.0027397 | 54 | 71 | 1.5041 | 0.00054795 | 54 | 150 | 10.0055
0.0017825 | 55 | 72 | 1.5027 | 0.0017825 | 55 | 72 | 1.5027 | 1.7764e-16 | 55 | 151 | 10
0.002611 | 56 | 73 | 1.5039 | 0.002611 | 56 | 73 | 1.5039 | 0.0002611 | 56 | 152 | 10.0026
0.0034014 | 57 | 74 | 1.5051 | 0.0034014 | 57 | 74 | 1.5051 | 0.0010204 | 57 | 153 | 10.0102
0.0016584 | 58 | 75 | 1.5025 | 0.0016584 | 58 | 75 | 1.5025 | 0.00024876 | 58 | 154 | 9.9975
0.0016181 | 59 | 76 | 1.5024 | 0.0016181 | 59 | 76 | 1.5024 | 0.00072816 | 59 | 155 | 9.9927
0.0015798 | 60 | 77 | 1.5024 | 0.0015798 | 60 | 77 | 1.5024 | 0.0007109 | 60 | 156 | 9.9929
0.0015432 | 61 | 78 | 1.5023 | 0.0015432 | 61 | 78 | 1.5023 | 0.00023148 | 61 | 157 | 9.9977
0.0030166 | 62 | 79 | 1.5045 | 0.0030166 | 62 | 79 | 1.5045 | 0.00090498 | 62 | 158 | 10.009
0.0022075 | 63 | 80 | 1.5033 | 0.0022075 | 63 | 80 | 1.5033 | 0.0004415 | 63 | 159 | 10.0044
0.0028736 | 64 | 81 | 1.5043 | 0.0028736 | 64 | 81 | 1.5043 | 0.00043103 | 64 | 160 | 10.0043
0.0035088 | 65 | 82 | 1.5053 | 0.0035088 | 65 | 82 | 1.5053 | 0.00084211 | 65 | 161 | 10.0084
0.0020534 | 66 | 83 | 1.5031 | 0.0020534 | 66 | 83 | 1.5031 | 0 | 66 | 162 | 10
0.002004 | 67 | 84 | 1.503 | 0.002004 | 67 | 84 | 1.503 | 0.0004008 | 67 | 163 | 9.996
0.0019569 | 68 | 85 | 1.5029 | 0.0019569 | 68 | 85 | 1.5029 | 0.00019569 | 68 | 164 | 9.998
0.0031867 | 69 | 86 | 1.5048 | 0.0031867 | 69 | 86 | 1.5048 | 0.00057361 | 69 | 165 | 10.0057
0.0024876 | 70 | 87 | 1.5037 | 0.0024876 | 70 | 87 | 1.5037 | 0 | 70 | 166 | 10
0.0018215 | 71 | 88 | 1.5027 | 0.0018215 | 71 | 88 | 1.5027 | 0 | 71 | 167 | 10
0.0023725 | 72 | 89 | 1.5036 | 0.0023725 | 72 | 89 | 1.5036 | 0.00053381 | 72 | 168 | 10.0053
0.0023148 | 73 | 90 | 1.5035 | 0.0023148 | 73 | 90 | 1.5035 | 0 | 73 | 169 | 10
0.0022599 | 74 | 91 | 1.5034 | 0.0022599 | 74 | 91 | 1.5034 | 0 | 74 | 170 | 10
0.0033113 | 75 | 92 | 1.505 | 0.0033113 | 75 | 92 | 1.505 | 0.00049669 | 75 | 171 | 10.005
0.0026925 | 76 | 93 | 1.504 | 0.0026925 | 76 | 93 | 1.504 | 0 | 76 | 172 | 10
0.002103 | 77 | 94 | 1.5032 | 0.002103 | 77 | 94 | 1.5032 | 0 | 77 | 173 | 10
0.0025681 | 78 | 95 | 1.5039 | 0.0025681 | 78 | 95 | 1.5039 | 0.00061633 | 78 | 174 | 10.0062
0.0025063 | 79 | 96 | 1.5038 | 0.0025063 | 79 | 96 | 1.5038 | 0.00015038 | 79 | 175 | 10.0015
0.0024474 | 80 | 97 | 1.5037 | 0.0024474 | 80 | 97 | 1.5037 | 0.00044053 | 80 | 176 | 10.0044
0.0019102 | 81 | 98 | 1.5029 | 0.0019102 | 81 | 98 | 1.5029 | 0.00028653 | 81 | 177 | 9.9971
0.002331 | 82 | 99 | 1.5035 | 0.002331 | 82 | 99 | 1.5035 | 0.00027972 | 82 | 178 | 9.9972
0.0027322 | 83 | 100 | 1.5041 | 0.0027322 | 83 | 100 | 1.5041 | 0.00013661 | 83 | 179 | 10.0014
0.0017778 | 84 | 101 | 1.5027 | 0.0017778 | 84 | 101 | 1.5027 | 0.00013333 | 84 | 180 | 9.9987
0.0026042 | 85 | 102 | 1.5039 | 0.0026042 | 85 | 102 | 1.5039 | 0.00013021 | 85 | 181 | 10.0013
0.0021177 | 86 | 103 | 1.5032 | 0.0021177 | 86 | 103 | 1.5032 | 0.00038119 | 86 | 182 | 9.9962
0.0024814 | 87 | 104 | 1.5037 | 0.0024814 | 87 | 104 | 1.5037 | 0.00024814 | 87 | 183 | 9.9975
0.0028283 | 88 | 105 | 1.5042 | 0.0028283 | 88 | 105 | 1.5042 | 0.00048485 | 88 | 184 | 10.0048
0.0027613 | 89 | 106 | 1.5041 | 0.0027613 | 89 | 106 | 1.5041 | 0.00047337 | 89 | 185 | 10.0047
0.0023095 | 90 | 107 | 1.5035 | 0.0023095 | 90 | 107 | 1.5035 | 0 | 90 | 186 | 10
0.0026306 | 91 | 108 | 1.5039 | 0.0026306 | 91 | 108 | 1.5039 | 1.7764e-16 | 91 | 187 | 10
0.0018335 | 92 | 109 | 1.5028 | 0.0018335 | 92 | 109 | 1.5028 | 0.00055006 | 92 | 188 | 9.9945
0.0017902 | 93 | 110 | 1.5027 | 0.0017902 | 93 | 110 | 1.5027 | 0.00042965 | 93 | 189 | 9.9957
0.0024484 | 94 | 111 | 1.5037 | 0.0024484 | 94 | 111 | 1.5037 | 0.00020986 | 94 | 190 | 10.0021
0.0027322 | 95 | 112 | 1.5041 | 0.0027322 | 95 | 112 | 1.5041 | 0.00030738 | 95 | 191 | 10.0031
0.002 | 96 | 113 | 1.503 | 0.002 | 96 | 113 | 1.503 | 0 | 96 | 192 | 10