## Slide 1: Phase 2 Cartridge Design Review
September 27, 2013

## Slide 2
1. Yellow dye bled across check valves during loading = bad check valve
on air syringe
2. Manipulating air syringe against closed cartridge
valve pushed yellow
dye towards puck; but backflow occurred when relaxed = bad check valve
on air syringe
Inlet line filled
with blue dye, but not pushed through puck
0.22 mm durapore filter
Sampling
position
Yellow dye made it
to here during step #2
#1
#2
#3
#4
Cartridge
Valve
The check valves do not function properly

## Slide 3
Occasional
bubbles
sweep through
Everything that should happen did happen
with softer check valve balls