McLane Visit

8 December 2008

Mechanical Problems Found:

1) Syringe frame missing valve mounting holes.

2) Rotary Valve gear on motor center hole oversized and loose on adapter

3) Syringe mount needs to be wider to fit over syringe car

4) Rvalve connector loading diagram is inverted

5) Rvalve channels assignments have changed and the drawing needs to be updated to reflect this change.
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6) Grill perforated panel needs to be changed to a “solid” panel that has holes for mounting in the correct position

