FIELD-DEPLOYABLE COLORIMETRIC BIOSENSOR SYSTEM FOR THE RAPID DETECTION OF PATHOGENIC ORGANISMS

By

Janice Duy

B.S. University of Maine, 2005

A DISSERTATION

Submitted in Partial Fulfillment of the

Requirements for the Degree of

Doctor of Philosophy

(Interdisciplinary in Functional Genomics)

The Graduate School

The University of Maine

May 2013

Advisory Committee:

Laurie B. Connell, Research Professor of the School of Marine Sciences, Co-Advisor

Rosemary L. Smith, Professor of Electrical and Computer Engineering, Co-Advisor

Scott D. Collins, Grand Vizier Professor of Chemistry

Keith W. Hutchison, Professor of Biochemistry and Molecular Biology

Michael D. Mason, Associate Professor of Chemical Engineering

DISSERTATION ACCEPTANCE STATEMENT

On behalf of the Graduate Committee for Janice Duy, I affirm that this manuscript is the final and accepted dissertation. Signatures of all committee members are on file with the Graduate School at the University of Maine, 42 Stodder Hall, Orono, Maine.

Dr. Laurie B. Connell, Research Professor of the School of Marine Sciences	    Date

Dr. Rosemary L. Smith, Professor of Electrical & Computer Engineering		    Date

© 2013 Janice Duy

All Rights Reserved

LIBRARY RIGHTS STATEMENT

In presenting this dissertation in partial fulfillment of the requirements for an advanced degree at The University of Maine, I agree that the Library shall make it freely available for inspection. I further agree that permission for “fair use” copying of this dissertation for scholarly purposes may be granted by the Librarian.  It is understood that any copying or publication of this dissertation for financial gain shall not be allowed without my written permission.

Signature:

Date:

FIELD-DEPLOYABLE COLORIMETRIC BIOSENSOR SYSTEM FOR THE RAPID DETECTION OF PATHOGENIC ORGANISMS

By Janice Duy

Dissertation Co-Advisors: Dr. Laurie B. Connell and Dr. Rosemary L. Smith

An Abstract of the Dissertation Presented

in Partial Fulfillment of the Requirements for the

Degree of Doctor of Philosophy

(Interdisciplinary in Functional Genomics)

May 2013

The rapid identification of pathogenic organisms is necessary for recognizing and managing human and environmental health risks.  Numerous detection schemes are available, but most are difficult to employ in non-laboratory settings due to their need for bulky, specialized equipment, multiple reagents, or highly trained personnel.  To address this problem, a rapid, field-compatible biosensor system based on the colorimetric detection of nucleic acid hybrids was developed.

Peptide nucleic acid (PNA) probes were used to capture ribosomal RNA sequences from environmental samples.  Non-target nucleic acids, including single-base mismatches flanked by adenines and uracils, were removed with a micrococcal nuclease digestion step.  Matched PNA-RNA hybrids remained intact and were indicated by the cyanine dye DiSC2(5).  PNA-containing duplexes function as templates for the aggregation of DiSC2(5), visualized as a change in solution color from blue to purple. This transition can be measured as an increase in the solution absorbance at 540 nm (dye aggregate) at the expense of the dye monomer peak at 650 nm. These concomitant spectral changes were used to calculate a “hybridization signal” using the ratio Aaggregate/Amonomer ≈ A540/A650.

Testing with pathogenic environmental samples was accomplished using two model organisms: the harmful algal bloom-causing dinoflagellate Alexandrium species, and the potato wart disease-causing fungus Synchytrium endobioticum.   In both cases, the colorimetric approach was able to distinguish the targets with sensitivities rivaling those of established techniques, but with the advantages of decreased hands-on time and cost.

Assay fieldability was tested with a portable colorimeter designed to quantify the dye-indicated hybridization signal and assembled from commercially available components.  Side-by-side testing revealed no difference in the sensing performance of the colorimeter compared to a laboratory spectrophotometer (Pearson’s r=0.99935). Assay results were obtained within 15 minutes, with a limit of detection down to 10-17 mole.  This quick, inexpensive and robust system has the potential to replace laborious pathogen identification schemes in field environments, and is easily adapted for the detection of different organisms.

# ACKNOWLEDGMENTS

I would like to extend a big thank you to my thesis committee, especially to my advisors Dr. Laurie Connell and Dr. Rosemary Smith.  Throughout my graduate career they have offered useful advice and direction as well as moral support and enthusiasm.  They have also afforded me great intellectual freedom and multiple opportunities to participate in innovative research, multinational collaborations, and global travel.  Drs. Scott Collins, Keith Hutchison, and Michael Mason have also provided invaluable knowledge, discussion, and laboratory help over the years.  Thank you!

Funding for various parts of this work were provided by grants: #0221625 from the National Science Foundation IGERT program to Functional Genomics, NOAA CSCOR MERHAB Program #NA11NOS4780026, #NA05NOS4781232, USDA Biosecurity #2006-55605-16654.  As well, I am grateful for support from the Canadian-American Center New England-Atlantic Provinces-Québec Fellowship	, University of Maine Maine Economic Improvement Fund Dissertation Fellowship, GSBS/Functional Genomics travel funds, and GSG degree-related awards.

I would also like to offer my gratitude to Drs. Paul Millard, Carol Kim, Darrell Donahue, and Joyce Longcore for their technical assistance, helpful conversations, and the use of their labs and equipment.

Finally, I would like to acknowledge my wonderful friends and family for their unwavering love, patience, and support.

TABLE OF CONTENTS

ACKNOWLEDGMENTS	iv

LIST OF TABLES	xii

LIST OF FIGURES	xiii

CHAPTER 1 INTRODUCTION	1

1.1 Motivation and objective	1

1.2 Near-real-time, field-based detection of unamplified nucleic acids	2

1.2.1 Toxic, harmful algal bloom-causing Alexandrium species	3

1.2.2 Synchytrium endobioticum, the causal agent of potato wart disease	5

1.3 Peptide nucleic acid (PNA)	6

1.3.1 PNA probe molecules for nucleic acid hybridization	8

1.3.2 Nucleic acid hybridization using PNA probes	11

1.4 Research approach	14

1.5 Dissertation organization	18

CHAPTER 2 PREPARATION OF SURFACTANT-STABILIZED GOLD NANOPARTICLE-PEPTIDE NUCLEIC ACID CONJUGATES§	20

2.1 Introduction	20

2.2 Background	21

2.3 PNA conjugation to gold nanoparticles	23

2.3.1 Methods and materials	23

2.3.1.1 Chemicals and reagents	23

2.3.1.2 Oligonucleotides	23

2.3.1.3 Gold nanoparticle synthesis and characterization	24

2.3.1.4 Preparation of AuNP-PNA conjugates	25

2.3.1.5 Enzymatic assessment of PNA attachment to AuNPs	25

2.3.1.6 Verification of the functionality of AuNP-PNA conjugates	26

2.3.2 Results and discussion	26

2.3.2.1 Characterization of gold nanoparticle dispersions	26

2.3.2.2 Verification of PNA attachment to gold nanoparticles	30

2.3.2.3 Specific hybridization of AuNP-PNA to complementary DNA	30

2.4 DNA-templated AuNP-PNA assembly	32

2.4.1 Hybridization of AuNP-PNA conjugates to complementary DNA linker	32

2.4.2 PNA sequence effects on gold nanoparticle aggregation	32

2.4.3 Effect of DNA linker concentration on AuNP-PNA aggregation	33

2.4.4 Effect of salt concentration on AuNP assembly	34

2.4.5 Effect of hybridization time and temperature	36

2.5 Chapter summary	37

CHAPTER 3 FIELD-DEPLOYABLE COLORIMETRIC BIOASSAY FOR THE RAPID AND SPECIFIC DETECTION OF HARMFUL ALGAL BLOOM-CAUSING DINOFLAGELLATES†	38

3.1 Introduction	38

3.2 Background	39

3.3 Colorimetric assay optimization	40

3.3.1 PNA probe and dye concentrations	40

3.3.2 Target detection from total RNA samples using nuclease digestion	41

3.4 Assay verification with synthetic RNA	44

3.4.1 Methods and materials	44

3.4.1.1 Oligonucleotides and reagents	44

3.4.1.2 PNA-RNA hybridization and specificity experiments	45

3.4.2 Results and discussion	46

3.4.2.1 Dye indication of PNA-RNA hybridization	46

3.4.2.2 Single-base mismatch discrimination of A. tamarense-based target      RNA sequences	48

3.4.2.3 RNA concentration dependence and discrimination against                       A. ostenfeldii RNA	50

3.5 Testing with total RNA extracted from Alexandrium cell cultures	52

3.5.1 Materials and methods	52

3.5.1.1 Reagents and supplies	52

3.5.1.2 Alexandrium cell culture growth conditions	53

3.5.1.3 Total RNA extraction from algal cells and RNA quality assessment	53

3.5.1.4 Total RNA fragmentation	55

3.5.1.5 PNA hybridization to fragmented total RNA	55

3.5.2 Results and discussion	56

3.5.2.1 Total RNA fragmentation for hybridization to PNA probes	56

3.5.2.2 Optimization of hybridization conditions	58

3.5.2.3 Discrimination between North American                                                    A. tamarense/fundyense/catanella ribotype and A. ostenfeldii	61

3.6 Chapter summary	61

CHAPTER 4 DEVELOPMENT OF A PORTABLE INSTRUMENT FOR FIELD DEPLOYMENT OF COLORIMETRIC BIOASSAY‡	63

4.1 Introduction	63

4.2 Methods and materials	64

4.2.1 Handheld colorimeter design	64

4.2.2 LED testing	67

4.2.3 Colorimeter comparison with a benchtop spectrophotometer	68

4.2.4 Data collection and analysis	69

4.3 Results and discussion	70

4.3.1 Colorimeter design considerations	70

4.3.1.1 Liquid sample holder selection	70

4.3.1.2 Light sources	70

4.3.1.3 Photodetector selection	72

4.3.1.4 Microcontroller selection	73

4.3.1.5 Touchscreen/display	74

4.3.2 Comparison of colorimeter and spectrophotometer response	74

4.4 Chapter summary	78

CHAPTER 5 DETECTION OF THE FUNGAL PHYTOPATHOGEN SYNCHYTRIUM ENDOBIOTICUM USING DiSC2(5)-INDICATED PNA HYBRIDIZATION	79

5.1 Introduction	79

5.2 Collection of potato wart DNA and RNA	80

5.3 Methods and materials	82

5.3.1 Oligonucleotides and reagents	82

5.3.2 DNA and RNA isolation from S. endobioticum wart samples	83

5.3.3 RNA isolation from other sample types	85

5.3.4 18S rDNA PCR amplification and cycle sequencing	86

5.3.5 Nested PCR with S. endobioticum-specific primers	87

5.3.6 In vitro transcription of 18S PCR product and RNA fragmentation	88

5.3.7 Quantitative PCR (qPCR) and reverse transcriptase PCR (RT-qPCR)	89

5.3.8 PNA-RNA hybridization conditions	90

5.4 Results and discussion	91

5.4.1 Optimization of nucleic acid extraction from potato wart	91

5.4.2 PNA probe development from S. endobioticum 18S rDNA sequences	92

5.4.3 S. endobioticum PCR primer development and testing	93

5.4.4 Optimization of qPCR conditions	95

5.4.5 Effect of RNA fragmentation on apparent target copy number	96

5.4.6 Target copy number estimations for potato wart genomic DNA and RNA samples using qPCR	97

5.4.7 Optimization of PNA-RNA hybridization conditions	99

5.4.8 Hybridization to in vitro transcribed RNA and assay detection limit	102

5.4.9 PNA probe specificity to other organisms	105

5.4.10 Hybridization to potato wart RNA	106

5.5 Chapter summary	108

Chapter 6 SUMMARY AND FUTURE CONSIDERATIONS	110

6.1 Contributions of this dissertation	110

6.2 Considerations for future work	112

6.2.1 RNA-templated formation of larger gold nanoparticle assemblies	112

6.2.2 Specific detection of low numbers of toxigenic Alexandrium cells from whole seawater samples	112

6.2.3 Quantitative detection of Synchytrium endobioticum from soil samples	113

BIBLIOGRAPHY	115

APPENDIX A. SENSITIVITY ANALYSIS OF

COLORIMETER PERFORMANCE	127

APPENDIX B. CHARACTERIZATION OF DNA AND RNA EXTRACTED FROM POTATO WART SAMPLES	134

APPENDIX C. EFFECTS OF RNASTABLE® STORAGE MATRIX ON HYBRIDIZATION USING DiSC2(5)	136

APPENDIX D. ALEXANDRIUM CELL COUNTING PROTOCOL	139

APPENDIX E. ALEXANDRIUM RNA EXTRACTION PROTOCOL USING CHEMICAL LYSIS	141

APPENDIX F. PCR AMPLIFICATION OF 18S rRNA GENE FOR SUBSEQUENT IN VITRO TRANSCRIPTION	144

APPENDIX G. IN VITRO TRANSCRIPTION OF 18S rRNA GENE USING PCR-AMPLIFIED dsDNA TEMPLATES	147

APPENDIX H. PNA HYBRIDIZATION TO SYNTHETIC RNA SEQUENCES: QUALITY CONTROL	151

BIOGRAPHY OF THE AUTHOR	157

# LIST OF TABLES

Table 1.1. Thermal stabilities of nucleic acid duplexes.	10

Table 2.1. Oligonucleotide sequences used in Chapter 2.	24

Table 3.1. Oligonucleotide sequences used in Chapter 3.	45

Table 5.1. Oligonucleotide sequences used in Chapter 5.	83

Table 5.2. Target copy numbers from potato wart genomic DNA samples.	98

Table 5.3. Target copy numbers from potato wart RNA samples.	99

Table 5.4. Correlations between RNA quality, target copy number, and colorimetric detection.	107

Table A.1. Normal operating parameters for colorimetric system.	132

Table B.1. Genomic DNA extracted from potato wart samples.	134

Table B.2. Genomic RNA extracted from potato wart samples.	135

# LIST OF FIGURES

Figure 1.1. Structural comparison between PNA and DNA backbones.	8

Figure 2.1. TEM images of gold nanoparticle dispersions.	27

Figure 2.2. Particle size distributions of AuNPs.	28

Figure 2.3. Absorbance spectra of gold nanoparticle dispersions.	29

Figure 2.4. Verification of PNA attachment to AuNPs.	30

Figure 2.5. AuNP-PNA hybridize sequence-specifically to DNA sequences.	31

Figure 2.6. DNA linker hybridization induces AuNP-PNA assembly.	33

Figure 2.7. Effect of DNA linker concentration on the formation of nanoparticle assemblies.	34

Figure 2.8. Effect of salt concentration on AuNP-PNA self-aggregation.	35

Figure 2.9. Effect of salt concentration on DNA-linked AuNP assembly.	36

Figure 2.10. Effect of heating on the formation of DNA-templated AuNP-PNA particle assemblies.	37

Figure 3.1. DiSC2(5) non-specifically indicates PNA-RNA hybridization.	42

Figure 3.2. RNase ONE™ resolves single-base mismatches in PNA-RNA hybrids.	43

Figure 3.3. Optical characteristics of indicator dye in the presence of PNA and PNA-RNA hybrids.	48

Figure 3.4. Micrococcal nuclease digestion of single-base mismatches.	50

Figure 3.5. Quantitative detection of A. tamarense RNA in mixtures with

A. ostenfeldii RNA.	52

Figure 3.6. RNA fragmentation improves hybridization to PNA.	57

Figure 3.7. GuSCN lysis-induced RNA fragmentation.	58

Figure 3.8. Hybridization temperature optimization for Alexandrium RNA.	60

Figure 3.9. Discrimination between related Alexandrium species.	61

Figure 4.1. Colorimeter block diagram.	65

Figure 4.2. Colorimeter unit, inside view.	66

Figure 4.3. Field-ready colorimeter in water-tight case.	67

Figure 4.4. Spectral characteristics of green and red LEDs.	72

Figure 4.5. Normalized photodiode spectral response.	73

Figure 4.6. Raw hybridization signals from spectrophotometer and colorimeter.	75

Figure 4.7. Actual and modeled colorimeter responses.	76

Figure 4.8. Comparison of colorimeter performance to benchtop spectrophotometer.	77

Figure 5.1. Experimental workflow for potato wart samples.	81

Figure 5.2. RNA extracted from potato wart samples show significant degradation.	92

Figure 5.3. Specificity of primer set to S. endobioticum and related chytrid species.	94

Figure 5.4. Effect of RNA fragmentation on apparent target copy number.	97

Figure 5.5. Optimization of hybridization temperature for S. endobioticum binding to PNA 033.	100

Figure 5.6. Further optimization of hybridization and digestion temperature of                 S. endobioticum RNA to PNA 033.	101

Figure 5.7. Effect of micrococcal nuclease concentration on signal-to-noise ratio.	102

Figure 5.8. Minimum detectable number of S. endobioticum copies assessed using in vitro transcribed 18S rRNA.	104

Figure 5.9. Detection of S. endobioticum from extracted RNA.	106

Figure B.1. Concentration-dependent RNAstable® absorbance characteristics.	137

Figure B.2. Raw and corrected hybridization signals for potato and potato wart RNA.	138

# INTRODUCTION

## Motivation and objective

The rapid and specific detection and identification of pathogenic organisms is important for recognizing and managing risks to human and environmental health.  Although many laboratory-based detection methods currently exist and are in use, some applications require the immediate identification of organisms in the field to prevent the spread of disease or to minimize environmental impact.  However, many benchtop techniques can be difficult to employ in field surroundings due to their need for bulky, specialized equipment, multiple reagents, controlled laboratory conditions, or skilled, trained personnel.  Thus, the goal of this work was to create a field-compatible, nucleic acid biosensing system to detect target organisms without the need for special training or equipment.  In particular, colorimetric (color change-based) detection was chosen for ease in both visual and electronic assessment of target abundance.  The specific goals and tasks associated with the development of such a system are:

- Identification of a nucleic acid hybridization-based detection method that is simple, rapid and compatible with the probe and target chemistries required;

- Optimization of assay parameters to achieve the specificity and sensitivity required for target identification in environmental samples; and

- Design and demonstration of a field-deployable instrument for quantifying target abundance using the developed assay.

To demonstrate the applicability of the system to different pathogenic organisms, two unrelated environmental pathogens were used to test and validate the system: the harmful algal bloom-causing dinoflagellate Alexandrium fundyense/tamarense/catanella species complex and the soil-borne obligate parasite Synchytrium endobioticum, which causes potato wart disease.

## Near-real-time, field-based detection of unamplified nucleic acids

Pathogenic microorganisms are biological agents found in natural surroundings that cause disease in plant or animal life-forms.  Certain pathogens that cause severe health, environmental, and/or economic impacts must be identified quickly and accurately to prevent disease transmission, crop losses or animal and human deaths.  Historically, whole-cell visual identification of targets is used as relatively simple instrumentation is required for testing.  However, cell counting is time-consuming and can be confounded by the presence of morphologically identical but possibly benign species.  The advent of nucleic acid sequencing technology ushered in the field of sequence-based detection.  The elegant and precise mechanism of Watson-Crick base-pairing was exploited to develop nucleic acid-based hybridization assays, where a probe molecule is designed to base-pair specifically with its complement in the target organism.

For field-deployable applications, the direct detection of unamplified nucleic acids, specifically ribosomal RNA (rRNA), is advantageous.  Ribosomal RNAs, along with ribosomal proteins, form the ribosomes, which are protein synthesis machinery essential to cell function.  In eukaryotes, rRNAs of different molecular weights have been identified: the 5S, 5.8S, 18S (small subunit, SSU) and 28S (large subunit, LSU).  These are present in all cellular life forms, and are abundant in the cell, with a copy number 100- to 1000-fold greater than ribosomal DNA (rDNA).  These molecules comprise ~90% of the total RNA extracted from the cell.  The large number of rRNAs precludes the need for sequence amplification prior to detection, which decreases both assay complexity, cost, and time to answer.  Second, rRNA sequences have both conserved and variable regions, so organisms can be classified at increasingly finer phylogenetic scale, down to species-level discrimination (Woese 1987).  Most hybridization probes target 18S rRNA sequences, as these were historically better-studied and sequenced due to its reduced complexity (~2000 bases for the 18S rRNA vs. ~5000 bases for the 28S rRNA).

The model organisms used in these studies are still primarily processed in laboratories with trained personnel, and would benefit from the development of rapid detection assays that can be performed with minimal training in resource-limited settings, as detailed in the following sections.

### Toxic, harmful algal bloom-causing Alexandrium species

The dinoflagellate genus Alexandrium includes a number of globally-distributed species that produce a family of neurotoxins, such as saxitoxins and their derivatives (Anderson, Kulis et al. 1990).  In North America, three species encompassing Alexandrium catanella, Alexandrium fundyense, and Alexandrium tamarense are of particular importance as they are found in most coastal waters of the United States and Canada (Anderson 1997).  Shellfish that ingest these toxic algae accumulate high concentrations of these neurotoxins, and are responsible for paralytic shellfish poisoning (PSP) incidents in human and animal consumers (Van Dolah 2000).  Since PSP can cause serious illness or even death if sufficient amounts of toxins are ingested, monitoring the spatial and temporal distribution of this organism is important.  However, accurately pinpointing the size and coverage of toxic Alexandrium populations is problematic since toxins may be produced at very low cell densities, and non-toxic, co-occurring strains can be difficult to distinguish using morphological characteristics (Scholin, Herzog et al. 1994; Touzet, Franco et al. 2007).  These challenges have necessitated a shift from cell identification by microscopic enumeration to molecular analysis methods.  Antibodies to a cell surface protein were developed (Anderson, Kulis et al. 1999), but further testing revealed that it could not discriminate between A. tamarense (saxitoxin-producing) and the co-occurring species A. ostenfeldii (non-saxitoxin-producing) (Anderson, Kulis et al. 2005).  Most Alexandrium molecular detection schemes now use oligonucleotide probes for their enhanced specificity.  Current detection methods which target rRNA sequences include whole-cell hybridization (Anderson, Kulis et al. 2005; Hosoi-Tanabe and Sako 2005), sandwich hybridization (Anderson, Kulis et al. 2005), and PCR-based (Dyhrman, Erdner et al. 2006) assays.  However, these detection schemes remain fairly complex due to the use of molecular labels, multiple reagents, and costly laboratory equipment.  Because the success of many of these methods relies on operator expertise (Godhe, Cusack et al. 2007), their widespread application to real-time HAB monitoring is unsuitable for citizen volunteers, such as those serving in regional US Phytoplankton Monitoring Networks (CCEHBR).

### Synchytrium endobioticum, the causal agent of potato wart disease

Synchytrium endobioticum is a soil-borne fungal parasite that is found worldwide in areas with cool, moist climates (Hampson 1981).  It is an obligate parasite that requires a living host to survive, and infects young sprouts of susceptible potato cultivars (Karling 1964).  Infected tubers develop unsightly and inedible wart-like growths, rendering them unmarketable.  The parasite also releases thick-walled resting spores which are highly infective at less than 1 spore per gram of soil (Chisnall Hampson 1992) and remain viable for over 70 years in the soil (Laidlaw 1985).  Regulations to prevent the spread of this fungus are thus enforced worldwide due to its potentially disastrous economic impact.  In particular, the U.S. Department of Agriculture classifies the fungus as a “select agent” not permitted into the country because of its potential to cause significant economic damage to the potato industry, valued at $4 billion in 2011 (2012).  Meanwhile, the European Plant Protection Agency prohibits crop growing in infested plots for 20 years, or until the fungus cannot be detected in the soil (McNamara and Smith 1998; 1999).

Established methods for the detection of S. endobioticum rely on the microscopic identification of spores by highly-trained personnel, followed by plant infectivity assays which take weeks to obtain a result (McNamara and Smith 1998; 1999; 2004; Agriculture 2007; Standards 2011).  The use of molecular methods has recently been investigated as an alternative to visual confirmation of viable fungal spores.  Endpoint PCR detection using a 543-bp region of the internal transcribed spacer (ITS) was first reported in 2004 (Niepold and Stachewicz 2004).  Another primer set designed to amplify a 472-bp region in the ITS (Lévesque, de Jong et al. 2001) was used for conventional PCR, and this was compared to real-time PCR with primers spanning a 186-bp region (van den Boogert, Gent-Pelzer et al. 2005).  The method was shown to identify S. endobioticum DNA from resting spores and spores spiked into soil.  Real-time PCR primers were subsequently developed to amplify an 84-bp region of the ITS (van Gent-Pelzer, Krijger et al. 2010) and demonstrated a lower detection limit (1 fg DNA compared to 100 fg from van den Boogert et al. (2005)).  Microarray hybridization was also used to detect reverse-transcribed S. endobioticum 18S rRNA extracted from infected potato tubers (Abdullahi, Koerbler et al. 2005).

Despite the sensitivity and specificity of these methods, these are still limited to laboratory use due to the specialized equipment and reagents required by these assays.  The development of a rapid and reliable test that can be performed with minimal training and equipment would be a valuable adduct to the identification methods already available, as suspected S. endobioticum infections can be diagnosed on-site without the lag time associated with sending possibly contaminated samples to specialized laboratories for testing (Agriculture 2007).

## Peptide nucleic acid (PNA)

Most direct detection schemes use DNA oligonucleotide probes; however, the use of DNA in field environments with complex sample matrices is hampered by its susceptibility to degradation from UV irradiation and nucleases.  Hence, non-native nucleic acid analogs which are suited to these “dirty” applications have gained popularity.  In particular, peptide nucleic acids (PNAs) have gained significant attention as hybridization probes for their increased sensitivity, specificity, and resistance to degradation.

Peptide nucleic acid (PNA) is a synthetic molecule originally designed to recognize double-stranded DNA sequences in the 1980s by the groups of biochemist Peter Nielsen and organic chemist Ole Buchardt (Nielsen, Egholm et al. 1991; Egholm, Buchardt et al. 1992).  Computer simulations of B-DNA were used to select a structurally similar candidate molecule that would retain hybridization properties.  It was determined that a new polymer with native nucleobases linked by methylenecarbonyl linkers to a backbone consisting of (2-aminoethylglycine) units would be able to form hybrid duplexes with native DNA and RNA (Nielsen, Egholm et al. 1991; Egholm, Buchardt et al. 1992; Egholm, Buchardt et al. 1993).  As PNA is a polyamide, it has an amine terminus (N-terminus) and a carboxy terminus (C-terminus), and these correspond to the 5′ and 3′ ends of native nucleic acids.  Figure 1.1 shows a structural comparison between PNA and DNA.

DNA

PNA

Figure 1.1. Structural comparison between PNA and DNA backbones.  In PNA, the charged sugar-phosphate backbone is replaced with a neutral pseudopeptide chain. B stands for any nucleobase.  Figure is modified from (Spaulding 2011).

### PNA probe molecules for nucleic acid hybridization

Experiments confirmed that PNA hybridizes to complementary DNA, RNA, and even PNA sequences via canonical Watson-Crick base-pairing (adenine binds to thymine/uracil, while cytosine binds to guanine and vice versa).  PNA is achiral and adapts to the preferred orientation of its complement; PNA-DNA duplexes retain B-form features (Leijon, Gräslund et al. 1994), while PNA-RNA hybrids have A-form structures (Brown, Thomson et al. 1994).  PNA-PNA hybrids also form a double-helix (Wittung, Nielsen et al. 1994) with unique structural features, now termed the P-form (Rasmussen, Kastrup et al.).  Interestingly, PNA can hybridize both in antiparallel or parallel conformation.  Antiparallel hybridization is similar to normal duplex formation, where the N-terminus of the PNA is oriented toward the 3′ end of the DNA/RNA, and the C-terminus faces the 5′ end.  Parallel hybridization of PNA orients the N-terminus to the 5′ end of the corresponding sequence, and this type of hybridization is less thermally stable and much slower than the former (Egholm, Buchardt et al. 1993).  Moreover, PNA was discovered to have unexpected advantages over native nucleic acids in forming these hybrids.

First, the stability of PNA-containing duplexes is much higher than native double helices, with antiparallel conformations of the same duplex type being more thermally stable.  For instance, a 12-base PNA-PNA hybrid has Tm = 67°C, while the corresponding PNA-DNA helix has Tm = 51°C and the same-sequence DNA-DNA hybrid Tm = 33.5°C (Wittung, Nielsen et al. 1994).  As given in Table 1.1 below, the thermal stabilities of PNA-containing hybrids is:

PNA-PNA > PNA-RNA > PNA-DNA

These relative stabilities were determined by measuring the melting temperature, Tm, of annealed oligonucleotides, where Tm is defined as the temperature at which 50% of the oligonucleotides not hybridized.

Table 1.1. Thermal stabilities of nucleic acid duplexes.  A comparison of thermal stabilities of 15-base oligonucleotide hybrids, arranged from most stable to least stable, from (Egholm, Buchardt et al. 1993; Jensen, Ørum et al. 1997).  For PNA-containing hybrids, the PNA sequence is N-TGTACGTCACAACTA-C.  Hybridizations were performed in a buffer containing 100 mM NaCl, 10 mM sodium phosphate, 0.1 mM EDTA, pH 7.

| Hybrid | Tm (°C) |
| --- | --- |
| PNA-PNA | 83.0 |
| PNA-RNA (antiparallel) | 72.3 |
| PNA-RNA (parallel) | 51.2 |
| PNA-DNA (antiparallel) | 69.5 |
| PNA-DNA (parallel) | 56.1 |
| DNA-DNA | 53.3 |
| DNA-RNA | 50.6 |

Next, PNA-containing hybrids are destabilized to a greater extent by mismatched bases.  For instance, a single mismatch in a PNA-DNA hybrid, on average, decreases the Tm by 15°C, while the corresponding average decrease for a DNA-DNA duplex is 11°C (Egholm, Buchardt et al. 1993).  In some cases, this stability change is double that of DNA-DNA complexes (Jensen, Ørum et al. 1997).  Thus, PNA has found wide application in fields where the discrimination of single nucleotide polymorphisms (SNPs) is essential.

PNA can form stable duplexes even at very low salt concentrations that prevent native nucleic acid hybridization due to its neutrally-charged backbone and lack of electrostatic repulsion between probe and target strands (Egholm, Buchardt et al. 1993; Ørum, Nielsen et al. 1995; Tomac, Sarkar et al. 1996).  In a study by Tomac et al. (1996), the stability of same-sequence PNA-DNA and DNA-DNA hybrids were compared as a function of salt concentration (10-500 mM NaCl).  They found that the Tm of DNA homoduplexes increased with NaCl concentration, while PNA-DNA hybrids showed slightly decreased Tm over this salt regime, with both duplexes having similar Tms at 500 mM NaCl.  This low-salt hybridization characteristic can be used to improve target accessibility by disrupting the higher-order structures formed by native nucleic acid sequences.

Finally, PNA is extremely stable in a variety of conditions that degrade DNA and RNA.  In a landmark study by Demidov (1994), PNA sequences did not show significant degradation in human blood serum, nuclear and cytoplasmic extracts from mouse tumor cells, Eschericia coli and Micrococcus luteus extracts, proteinase K, or porcine intestinal peptidase.  PNA also retains its ability to hybridize even in the presence of chaotropic agents, such as high concentrations of guanidinium isothiocyanate (Connell, Ray et al. 2006; Bratcher and Connell 2012).  Sample preparation steps such as nucleic acid purification can consequently be eliminated, reducing both assay complexity and time.

### Nucleic acid hybridization using PNA probes

Improvements in the specificity, sensitivity, and time-to-answer of hybridization-based assays have been achieved with the replacement of DNA probes with PNA sequences.  For instance, DNA probes for whole-cell-based fluorescence in situ hybridization (FISH) protocols must be designed to target only accessible areas of the rRNA sequence (Fuchs, Wallner et al. 1998) because the secondary structure of the rRNA molecule is not disrupted by typical hybridization.  The use of PNA probes in FISH demonstrated that even sites with high rRNA self-complementarity could be detected because buffer ionic strength can be lowered and the hybridization temperature can be raised (Perry-O’Keefe, Rigby et al. 2001).  Early researchers reduced hybridization time from hours (for DNA probes) to 30 minutes (Perry-O’Keefe, Stender et al. 2001), and new QuickFISH™ methods have hybridization times of 15 minutes without any wash steps, leading to a total assay time of <30 minutes (Deck, Anderson et al. 2012) without loss of sensitivity or specificity.  Signal intensities are also improved at least fivefold when using PNA probes (Worden, Chisholm et al. 2000).

PNA have also found use in amplification-based schemes.   The earliest use of PNA in polymerase chain reaction (PCR) was termed “PCR clamping”: a PNA-containing duplex was not amplified because the hybrid could not be recognized by the DNA polymerase.  This method was used to preferentially amplify wild-type targets over mutants, and could discriminate SNPs (Ørum, Nielsen et al. 1993; Ørum 2000).  Applications to real-time PCR were realized with PNA molecular beacons, light-up PNA, and forced intercalation PNA probes.  PNA molecular beacons (MBs) are PNA sequences that are dual-labeled with a fluorophore and a quencher at either end.  Unlike DNA MBs, which require the addition of a self-complementary sequence at the ends to form a “stem-loop” structure which keeps the fluorophore and quencher together (Tyagi and Kramer 1996), PNA MBs are stemless.  The stem-loop structure was found to be unnecessary with PNA (Seitz 2000; Kuhn, Demidov et al. 2002), likely because PNA adopts a compact form in solution due to intra- and intermolecular forces (Achim, A. Armitage et al. 2008).  Light-up probes, meanwhile, have the asymmetric cyanine dye thiazole orange (TO) attached to one end of the PNA sequence (Svanvik, Westman et al. 2000), and are nonfluorescent when unbound to a target strand.  Another promising PNA probe, a “forced intercalation” or FIT-probe, has emerged in recent years.  Instead of an end-tethered fluorescent label, one of the nucleobases in the sequence is replaced with a TO molecule (Köhler and Seitz 2003).  When hybridized to target DNA, the artificial base fluoresces, with signal improvement of up to 26-fold when the TO is paired with an adenine (Köhler, Jarikote et al. 2005).  All three probe types have been validated for use in real-time PCR, and offer the advantages of sequences that are shorter and easier to design than those of DNA probes as well as increased specificity to mismatches (Svanvik, Stahlberg et al. 2000; Petersen, Vogel et al. 2004; Socher, Jarikote et al. 2008).

Surface-based hybridization assays have also benefited from the use of PNA as the immobilized probe layer.  For instance, using a surface plasmon resonance (SPR) system, a PNA probe layer was able to discriminate a single G-G mismatch in short DNA sequences at both 25°C and 40°C, while a corresponding DNA probe layer could only resolve the mutation at the higher temperature (Sato, Fujimoto et al. 2003).  PNA probes were also able to capture target strands over single base-mismatched sequences in low ionic strength solutions including plain water, where DNA probes showed negligible target binding (Park, Germini et al. 2007; Ananthanawat, Vilaivan et al. 2010).  D’Agata et al. (D’Agata, Corradini et al. 2008; D’Agata, Corradini et al. 2010; 2011) recently demonstrated the ultrasensitive (down to 10-21 mole target) and specific detection of unamplified genomic DNA by sandwiching the target DNA between an immobilized PNA probe and DNA-coated gold nanoparticles for signal enhancement.

Non-amplification-based solution hybridization reactions, where the probe molecules are not immobilized, have also gained improvements from the introduction of PNA probes.  For instance, stemless PNA MBs were found to hybridize complementary ssDNA targets within 3 minutes at 25°C, with a signal-to-background ratio of ~10.  Lowering the ionic strength of the hybridization buffer did not change these characteristics, but negatively impacted DNA MB hybridization (Kuhn, Demidov et al. 2002).  Meanwhile, Xi et al. (2003) used E. coli 16S rRNA and reported a detection limit of 12.5 nM with DNA MBs, and 390 pM with PNA MBs (an improvement of 32-fold).  Higher signal intensities from the PNA-based assay were also observed, and hybridization times were shorter, with signals significantly above background after 5 minutes.  Li et al. (2008) built upon this work and demonstrated the viability of using PNA MBs for detecting bacterial targets in environmental samples.  PNA FIT-probes have also been used to detect single mismatches in a structured RNA target sequence at 25°C (Socher, Bethge et al. 2008), which shows promise for RNA detection complex mixtures without external heating requirements.

## Research approach

The detection of pathogenic organisms in near-real-time using platforms and methods that are user-friendly, accurate, and quick has profound applications to human and environmental health outcomes.  Such methods have direct relevance to testing in diverse fields such as food safety, bioterrorism, disease diagnosis, and environmental monitoring.  Classical detection methods for pathogens include cell culturing, immunoassays, microscopy, and polymerase chain reaction (PCR) (Deisingh and Thompson 2002; Lazcka, Del Campo et al. 2007). However, these techniques are generally laboratory-based and require specialized equipment and reagents, trained personnel, and clean, controlled environments. Assays may take up to a week, which can hamper timely decision-making for controlling the spread of pathogens. The development of field-ready techniques which are capable of quickly and accurately identifying harmful organisms is thus important.

Biosensors are uniquely suited to solving this problem.  Biosensors consist of a molecular recognition element coupled with a physicochemical transducer which produces a usable electrical signal.  Although various biorecognition molecules such as receptors, enzymes, antibodies and lectins may be employed, DNA is widely used because it is highly selective and easy to manufacture (Chambers, Arulanandam et al. 2008; Teles and Fonseca 2008).  In addition, DNA strands can be easily attached to various transducer surfaces such as glass, silicon, and gold using simple chemistries (Sassolas, Leca-Bouvier et al. 2008).  The hybridization event between the surface-bound or free probe and the target in the sample is then converted into a detectable signal through various transduction mechanisms (Vercoutere and Akeson 2002; Luong, Male et al. 2008; Sassolas, Leca-Bouvier et al. 2008; Teles and Fonseca 2008).

Transduction mechanisms are broadly categorized as microgravimetric (mass-sensitive), electrochemical, or optical.  Microgravimetric schemes include quartz crystal microbalances (QCMs), microcantilevers, and acoustic wave (AW) devices that detect the added mass of captured targets on sensor surfaces. These platforms are easy to fabricate and miniaturize, but can be difficult to operate in liquids and suffer from decreased sensitivity due to nonspecific adsorption on surfaces (Lazcka, Del Campo et al. 2007; Lucarelli, Tombelli et al. 2008; Teles and Fonseca 2008).  Electrochemical methods measure changes in current, voltage, or impedance due to the hybridization event.  Assay times are short and devices are simple and inexpensive to fabricate, easy to miniaturize, and readily integrated with microfluidics.  However, these transducers usually require redox-active species to achieve adequate selectivity and sensitivity (Mello and Kubota 2002), and are affected by the electrical properties of the sample solution (Palchetti and Mascini 2008).  Optical methods, which can use planar waveguides or optical fibers to direct light to the sensing layer, can measure fluorescence, absorbance, chemiluminescence, and reflectance (Fan, White et al. 2008; Luong, Male et al. 2008). Remote interrogation of dangerous samples can be performed, and sensors can be multiplexed to detect multiple targets simultaneously.

The biosensor system was originally intended for monitoring the harmful algal bloom-causing North American Alexandrium species complex.  The spatial and temporal distribution of these blooms is unpredictable, so frequent sampling by volunteer workers in coastal environments is needed to pinpoint bloom hotspots.  The detection method presented in this work was therefore selected based on: 1) ease of use; 2) translatability to non-laboratory settings; 3) detection accuracy; and 4) cost.  The use of PNA probes was a natural choice for this application because these sequences are robust and have already been demonstrated to hybridize Alexandrium sequences in the extraction buffer (Connell, Ray et al. 2006).  Additionally, PNA offers more specific target recognition than DNA probes, as detailed above.  A solution-based hybridization approach was chosen minimize nonspecific adsorption effects common to surface-based binding schemes (Lazcka, Del Campo et al. 2007; Lucarelli, Tombelli et al. 2008; Teles and Fonseca 2008).

The initial concept for the biosensor system was based on the target-templated aggregation of PNA-functionalized gold nanoparticles (AuNPs).  AuNPs have size-dependent optical characteristics that can be harnessed to produce color changes indicative of target presence (Russ Algar, Massey et al. 2009).  For instance, solutions of DNA-functionalized nanoparticles turn from red to purple-gray when complementary DNA links these probe-coated labels, forming larger particle assemblies (Alivisatos, Johnsson et al. 1996; Mirkin, Letsinger et al. 1996).  The reaction is almost immediately visible as a color change, so end users do not need optical instruments to determine whether the target is present.  If quantitation of the target is necessary, an absorbance detector could be built using a straightforward design and at low cost.  To this end, PNA attachment to AuNPs was accomplished, and the probes retained their ability to bind complementary DNA.  However, DNA-templated aggregation of these PNA-coated particles proved too unpredictable for field-based use and this approach was not further pursued.

The concept of a visible detection method remained attractive, but a suitable optical reporter was needed.  While known for their sensitivity, fluorescently-labeled PNA probes were significantly more expensive than unlabeled sequences (~50% more in 2009), and specialized optical instrumentation is needed to measure fluorescence.  Probes labeled with fluorophores would also need washing steps to remove background fluorescence signals, adding to assay complexity.  Optically active duplex-binding molecules were another option, but most traditional DNA hybridization indicators are unable to recognize PNA-containing hybrids (Wittung, Kim et al. 1994).  A PNA-binding dye was finally discovered 5 years after this study (Smith, Olson et al. 1999), and its use was verified by other researchers with DNA target sequences (Wilhelmsson, Nordén et al. 2002; Komiyama, Ye et al. 2003; Tedeschi, Sforza et al. 2007).  The dye, 3,3'-diethylthiadicarbocyanine iodide (DiSC2(5)), is readily available and inexpensive.  More importantly, DiSC2(5)-indicated hybridization can be perceived with the naked eye as a blue-to-purple color change, while UV-vis spectrophotometry can also be used to quantify hybridization signals.  Since DiSC2(5) is not covalently attached to the probe, unmodified PNA could be purchased at a significant cost savings.  In addition, no studies investigating the affinity of the dye for PNA-RNA hybrids had been reported (as of 2009), so this approach was explored for applicability to a field-deployable assay.

## Dissertation organization

Chapter 2 presents a technique for functionalizing colloidal gold nanoparticles with thiolated peptide nucleic acid probes.  Stable PNA attachment was achieved by first shielding the nanoparticles with a nonionic surfactant prior to probe functionalization.  The PNA-decorated gold nanoparticles were stable in solution and were successfully hybridized to complementary DNA sequences.  However, further experiments revealed that hybridization-induced particle aggregation could not be precisely controlled and thus this approach could not attain the specificity required for this project.

Chapter 3 presents a colorimetric bioassay for the detection of PNA-RNA hybrids.  The model organism for these studies was the toxic North American dinoflagellate Alexandrium fundyense/tamarense/catanella species complex.  Non-target sequences, including those with single-base mismatches flanked by adenines and uracils, were digested with micrococcal nuclease.  Perfectly-matched hybrids were visualized with the cyanine dye 3,3'-diethylthiadicarbocyanine iodide (DiSC2(5)), which turns from blue to purple in the presence of nucleic acid hybrids.  The assay was able to hybridize RNA sequences from the target even in the presence of competing RNA.  Testing with RNA from cell cultures revealed that A. fundyense was detectable at ~3300 cells, while discrimination against the related A. ostenfeldii was achieved at ~6700 cells.

Chapter 4 describes the design and prototyping of a low-cost two-color detector for the portable quantitation of the hybridization signals from the bioassay in Chapter 3.  The device was tested against a benchtop spectrophotometer, and their responses were found to be comparable over the RNA concentration range tested (250-1000 nM).

Chapter 5 demonstrates the performance of the colorimetric assay with total RNA extracted from potato wart samples.  A PCR primer set and a PNA probe were designed to target the 18S rRNA based on available sequence data.  The primer set amplifies a 365-bp region and was validated with both endpoint and real-time PCR.  The PNA probe accurately identified S. endobioticum from potato wart samples containing mostly non-target RNA, with a detection limit down to ~10-17 mole (1.44 pM) for in vitro transcribed RNA.  The colorimetric assay was able to positively identify S. endobioticum-containing total RNA samples from potato wart tissue.

Chapter 6 summarizes the significant results from these studies and proposes future experiments related to this work.

# PREPARATION OF SURFACTANT-STABILIZED GOLD NANOPARTICLE-PEPTIDE NUCLEIC ACID CONJUGATES§

## Introduction

________________________________________

§Sections of this chapter are composed in part from excerpts from Duy, J., Connell, L.B., Eck, W., Collins, S.D., Smith, R.L. “Preparation of surfactant-stabilized gold nanoparticle–peptide nucleic acid conjugates,” 2010. J. Nanopart. Res. 12(7): 2363-2369.

A simple, two-step method of producing stable and functional peptide nucleic acid (PNA)-conjugated gold nanoparticles was investigated for use in biosensing by means of target-mediated aggregation.  The use of a surfactant was explored to prevent the spontaneous aggregation of gold nanoparticles during conjugation with thiolated PNA sequences.  While the chemical stability and target discrimination of PNA probes is superior to that of DNA, their use in metallic nanoparticle systems has been limited by the difficulty of producing stable colloids of nanoparticle-PNA conjugates.  In this work, the nonionic surfactant Tween 20 (polyoxyethylene (20) sorbitan monolaurate) was used to sterically shield gold surfaces prior to the addition of thiolated PNA, producing conjugates which remain dispersed in solution and retain the ability to hybridize to complementary nucleic acid sequences. The conjugates were characterized using transmission electron microscopy, dynamic light scattering, and UV-visible absorbance spectroscopy. PNA attachment to gold nanoparticles was confirmed with an enzyme-linked immunoassay, while the ability of nanoparticle-bound PNA to hybridize to its complement was demonstrated using labeled DNA.  The formation of larger particle assemblies was accomplished by hybridization with a DNA linker, and heating was explored as a means of accelerating the rate of the aggregation process.

## Background

A very sensitive method of optically confirming DNA hybridization is through the use of gold nanoparticle aggregation.  DNA-gold nanoparticle conjugates belong to a class of hybrid materials that merge the specific molecular recognition ability of biomaterials with the size-tunable chemical and physical properties of nanoscale materials.  They have a wide range of applications including biosensing, targeted drug delivery and gene regulation (Rosi and Mirkin 2005).  The conjugation of DNA to gold nanoparticles was first reported in 1996 for the formation of larger particle assemblies organized by sequence-specific DNA hybridization (Alivisatos, Johnsson et al. 1996; Mirkin, Letsinger et al. 1996).  The attachment of DNA to gold nanoparticles is uncomplicated since thiol-terminated DNA strands attach strongly and specifically to gold surfaces, and the negatively-charged phosphate backbones of the strands maintain electrostatic repulsion between individual gold nanoparticles (Mirkin, Letsinger et al. 1996).  Stable gold nanoparticle-DNA conjugate dispersions (where the particles do not aggregate uncontrollably) can thus be produced while maintaining the ability of DNA to hybridize to complementary strands.

However, the limited chemical stability of DNA (Lindahl 1993) compromises its utility in “dirty” field environments or in complex sample matrices.  Peptide nucleic acid (PNA), a DNA mimic, has gained recent attention as a superior probe molecule for such applications.  PNA consists of an uncharged polyamide backbone made of repeating N-(2-aminoethyl)glycine units; nucleobases are joined to the backbone by methylenecarbonyl linkages (Nielsen, Egholm et al. 1991).  PNA hybridizes to complementary DNA and RNA sequences due to comparable intermolecular distances and nucleobase configurations (Nielsen, Egholm et al. 1991; Egholm, Buchardt et al. 1992; Egholm, Buchardt et al. 1993), while its uncharged peptide backbone creates a more discriminating probe molecule.  Its lack of electrostatic repulsion to DNA strands allows PNA to bind complementary sequences even in solutions of low ionic strength, and while PNA-DNA duplexes exhibit a higher thermal stability than corresponding DNA-DNA duplexes, sequence mismatches are more destabilizing (Egholm, Buchardt et al. 1993).  In addition, PNA is chemically robust and is not susceptible to degradation by nucleases or proteases (Demidov, Potaman et al. 1994).

Despite these advantages of PNA, it has not superseded DNA in nanoparticle-based schemes because the attachment of PNA to gold nanoparticles has so far proven difficult (Murphy, Redmond et al. 2004; Lytton-Jean, Gibbs-Davis et al. 2009).  The addition of thiolated PNA to citrate-stabilized gold nanoparticle dispersions results in the immediate agglomeration and precipitation of the particles.  The uncharged, thiolated PNA strands presumably displace citrate anions from the gold surface, allowing these uncovered areas to bind irreversibly to other exposed gold cores, making the dispersions highly unstable.  To date, reports of successful attachment of PNA to gold nanoparticles have required the use of unconventional probe motifs.  Negative charges were imparted to PNA strands either from amino acid residues at the tethered end (Chakrabarti and Klibanov 2003; Murphy, Redmond et al. 2004), or through the creation of PNA-DNA chimeras (Murphy, Redmond et al. 2004).  Although these modifications have allowed for the conjugation of PNA to gold nanoparticles without loss of colloid stability, PNA functionality was compromised.

## PNA conjugation to gold nanoparticles

### Methods and materials

#### Chemicals and reagents

Gold(III) chloride trihydrate, trisodium citrate dihydrate, sodium chloride, Tween 20, and sulfuric acid were purchased from Sigma Aldrich (USA).  Horseradish peroxidase conjugated streptavidin, peroxidase conjugated anti-digoxin, and 3,3',5,5'-tetramethylbenzidine (TMB) were purchased from Pierce (USA).  Milli-Q water with resistance > 18 MΩ/cm was used in all experiments.

#### Oligonucleotides

All oligonucleotide sequences used are shown in Table 2.1.  PNA sequences were purchased from Panagene (Korea).  The DNA linker sequence was obtained from Integrated DNA Technologies (USA).

Table 2.1. Oligonucleotide sequences used in Chapter 2.  Italicized portions of the DNA linker are complementary to PNA 008, while portions in bold are complementary to PNA 010.  E linkers in the PNA sequences were added to prevent self-aggregation, while O spacers extend the immobilized probe sequence from the surface.  The UniR sequence contains the degenerate bases M (either adenosine or cytosine) and W (either adenosine or thymine)

| Name | Sequence (N-C/5'-3') |
| --- | --- |
| PNA UniR | N-HS-OOOO-CAGCMGCCGCGGTAATWC-EE-Biotin-C |
| PNA 008 | N-Biotin-EE-GTGCAACAATCCCACC-OOOO-SH-C |
| PNA 010 | N-Biotin-EE-CAACACCATGTGAACTG-OOOO-SH-C |
| DNA linker | 5'-GGTGGGATTGTTGCAC-TTTTTTT-Digoxigenin-TTTTTTT-CAGTTCACATGGTGTTG-3' |

#### Gold nanoparticle synthesis and characterization

Gold nanoparticles were synthesized by the reduction of 1 mM gold(III) chloride trihydrate with 30 mM trisodium citrate dihydrate (Grabar, Freeman et al. 1995).  The resulting particles were characterized using TEM, DLS, and UV-visible absorbance spectroscopy.  A Philips CM-10 transmission electron microscope (accelerating voltage 92 kV) equipped with a Gatan BioScan model 792 digital camera was used to image gold nanoparticles.  DLS measurements were performed at 20°C using a Zetasizer Nano ZS (Malvern Instruments, UK) with a 4 mW He-Ne laser (633 nm wavelength), and the solvent refractive index was set to 1.330 while the complex index of refraction of the particles was set to n = 0.2 and k = 3.32.  Absorbance spectroscopy was performed using a NanoVue spectrophotometer (GE Healthcare) using a 4 µL sample solution volume.

#### Preparation of AuNP-PNA conjugates

Solutions of 10% (v/v) Tween 20 in water (final concentration 84.6 mg/mL) and citrate-stabilized gold nanoparticles (final concentration 3.97 nM) were mixed and allowed to react for 30 minutes in a rotisserie-type shaker at room temperature.  Next, 860 pmol of biotin-labeled, thiol-terminated PNA (final concentration 132.3 nM) was added and incubated at room temperature in a shaker for 24 hours.  The resulting dispersion was rinsed twice by centrifugation at 15700g for 15 minutes, then resuspended in a solution of 0.02% (v/v) Tween 20 in water.

#### Enzymatic assessment of PNA attachment to AuNPs

PNA attachment to gold nanoparticles was confirmed using the free biotin label at the end of each PNA sequence.  Gold nanoparticle dispersions were standardized to 25% of the original (citrate-stabilized) concentration at 506 nm (Liu, Atwater et al. 2007).  Horseradish peroxidase-conjugated streptavidin (SA-HRP), 990 µL of 300 ng/mL concentration in 0.02% Tween 20, was added to 10 µL of diluted gold nanoparticle dispersion and allowed to react for 30 minutes in a shaker at room temperature.  Excess SA-HRP was removed through five rinses by centrifugation with 0.02% Tween 20 in water.  The resulting pellet from the last centrifugation was mixed with 100 µL TMB and incubated for 5 minutes.  The formation of a blue product from the reaction of TMB in the presence of horseradish peroxidase was measured at 370 nm and 650 nm, the major absorbance peaks of oxidized TMB.

#### Verification of the functionality of AuNP-PNA conjugates

Digoxigenin-labeled DNA (10 pmol, 10 nM final concentration) in a solution of 200 mM sodium phosphate dibasic, pH 8.1 with 0.02% Tween 20, was added to 6.48 nM of gold nanoparticle dispersion and allowed to hybridize for 30 minutes at room temperature.  The excess DNA was removed by five rinses by centrifugation with BSA blocker (bovine serum albumin 0.5% (w/v), Tween 20 0.05% (v/v) in phosphate buffered saline pH 7.4) to prevent nonspecific antibody adsorption in the next step.  The remaining pellet was reacted with 1:1000 (v/v) horseradish peroxidase-conjugated anti-digoxin (anti-dig HRP) in BSA blocker for 30 minutes at room temperature.  Excess anti-dig HRP was removed by rinsing and centrifugation as described above.  The pellet was mixed with TMB and solution absorbance peaks were read at 370 nm and 650 nm.

### Results and discussion

#### Characterization of gold nanoparticle dispersions

TEM images (Figure 2.1) of the effect of surfactant adsorption onto gold nanoparticles before and after the addition of thiolated PNA are shown in contrast to citrate-stabilized particles.  Figure 2.1a and Figure 2.1c, respectively, show surfactant- and citrate-stabilized particles before the addition of thiolated PNA.  The surfactant-stabilized particles remain dispersed, with minimal aggregation, after PNA attachment (Figure 2.1b), while the citrate-stabilized particles formed large crosslinked aggregates (Figure 2.1d).  The AuNP-PNA conjugates were also stable in buffers containing various salt concentrations (from 0 to 0.3 M NaCl) for several months (data not shown).  Mean diameters obtained from the TEM images were 14.26 ± 2.46 nm (citrate-stabilized particles), 16.25 ± 1.14 nm (surfactant-modified), and 16.05 ± 1.62 nm (PNA-modified).

Figure 2.1. TEM images of gold nanoparticle dispersions.  Surfactant-stabilized dispersions (a) before and (b) after the addition of thiolated PNA.  Citrate-stabilized dispersions (c) before and (d) after the addition of thiolated PNA.  The scale bar in each image is 100 nm.

The particle size distributions of the gold nanoparticle systems obtained by DLS were monomodal, with narrow peaks (Figure 2.2), indicating disperse particles.  The citrate-stabilized particles yielded a DLS-determined mean diameter (Z-average diameter) of 24.09 nm, while the surfactant-modified particles yielded a somewhat larger mean diameter of 29.67 nm without the formation of any additional peaks.  The AuNP-PNA conjugates yielded much larger mean diameters of 98.10 nm, 93.42 nm, and 102 nm for PNA UniR, PNA 008, and PNA 010, respectively.  Since the analysis of TEM images of gold nanoparticles before and after the surfactant-mediated attachment of PNA revealed no apparent differences in particle diameters, and no fused gold cores, this discrepancy in diameter enlargement could be due to the formation of small aggregates in solution, possibly due to the weak adsorption of the biotin label (which contains a sulfide group) to gold surfaces (Love, Estroff et al. 2005).  In contrast, DLS particle size distributions of gold nanoparticle dispersions functionalized with the same PNA sequences but labeled with a Cy5 fluorophore yielded a Z-average diameter of 36 nm.

Figure 2.2. Particle size distributions of AuNPs.  Particle size distributions by number of gold nanoparticle solutions show increased particle sizes in the samples exposed to thiolated PNA, indicating the formation of PNA-conjugated gold nanoparticles.

Normalized optical absorption spectra for the gold nanoparticle dispersions are shown in Figure 2.3.  The absorbance peak of the unmodified particles was at ~520 nm, while that of the surfactant-stabilized particles was shifted to 523 nm. The absorbance peak of the AuNP-PNA conjugates was further shifted to ~530 nm, with a slight increase in the absorbance values at longer wavelengths (>600 nm).  These shifts are consistent with previously reported results for surfactant- (Aslan and Perez-Luna 2002) and DNA-coated (Xu and Craig 2005) gold nanoparticles.  These observed absorbance shifts are most likely due to changes in the dielectric properties of the layer immediately surrounding the gold nanoparticles (Mulvaney 1996; Schmitt, Machtle et al. 1999; Eck, Helm et al. 2001), and to some small amount of particle aggregation.  Again, biotin-gold interactions could be responsible for this effect, as absorbance spectra obtained from gold nanoparticles functionalized with Cy5-labeled PNA only exhibited a 2-nm peak shift, without absorbance increases at longer wavelengths.

Figure 2.3. Absorbance spectra of gold nanoparticle dispersions.  The surfactant- and PNA-treated particles exhibit slight shifts in absorbance peak wavelengths, due to the alteration of the dielectric constant immediately surrounding the particles.  Nanoparticles with PNA also show increases in absorbance values at longer (>600 nm) wavelengths, possibly due to interactions between the gold surface and biotinylated PNA ends which lead to a small amount of aggregation.

#### Verification of PNA attachment to gold nanoparticles

The addition of TMB to AuNP-PNA conjugate dispersions treated with SA-HRP resulted in large increases in absorbance at both 370 nm and 650 nm compared to both citrate-stabilized and Tween 20-treated gold nanoparticles (Figure 2.4), indicative of the oxidation of TMB and hence the presence of attached PNA.  The changes in peak absorbance values differed for the three PNA sequences tested, with increases of 12, 10, and 5 times that of the Tween 20-stabilized particles for the PNA UniR, PNA 008, and PNA 010, respectively.

Figure 2.4. Verification of PNA attachment to AuNPs.  The attachment of biotinylated PNA to gold nanoparticles was confirmed using the oxidation of tetramethylbenzidine (TMB) by horseradish peroxidase-conjugated streptravidin (SA-HRP).  SA-HRP was allowed to bind to biotinylated PNA and unreacted molecules were removed by washing before the addition of TMB, which is oxidized by horseradish peroxidase.  TMB develops absorbance peaks at 370 nm and 650 nm when it is oxidized.

#### Specific hybridization of AuNP-PNA to complementary DNA

Increases in TMB absorbance peak values due to the specific hybridization of PNA probe-coated gold nanoparticles are shown in Figure 2.5.  Background levels of absorbance at 370 nm and 650 nm were observed for gold nanoparticle dispersions without complementary DNA, and for unmodified and negative control PNA UniR-coated nanoparticles with target DNA. AuNP-PNA conjugates allowed to hybridize with complementary DNA exhibited greater than a 2.5-fold increase in absorbance at 370 nm over the background level.  Observed differences in the amount of hybridized DNA target to PNA 008 and PNA 010 could be ascribed to sequence-dependent PNA characteristics, which may have affected particle surface coverage (Demers, Mirkin et al. 2000) and hybridization kinetics (Kushon, Jordan et al. 2001; Peterson, Heaton et al. 2001; Armitage 2003).  The presence of free surfactant molecules in solution should not have affected DNA recognition since detergents in concentrations of up to 0.1% are routinely used in hybridization buffers for PNA hybridization (Williams, Stender et al. 2002).

Figure 2.5. AuNP-PNA hybridize sequence-specifically to DNA sequences.  Hybridization of complementary digoxigenin-labeled DNA to AuNP-PNA conjugates verified using the oxidation of tetramethylbenzidine (TMB) by horseradish peroxidase-conjugated anti-digoxin (anti-dig HRP).  The development of TMB absorbance peaks at 370 nm and 650 nm indicate DNA hybridization.

## DNA-templated AuNP-PNA assembly

### Hybridization of AuNP-PNA conjugates to complementary DNA linker

The AuNP-PNA were placed in hybridization buffer by first centrifuging the original solutions (220 µL in a 1.7 mL polypropylene microcentrifuge tube) at 15700g for 7 minutes.  The supernatant was removed and the pellet was resuspended in an equal volume of hybridization buffer.  Nanoparticles conjugated to both PNA 008 and 010 were mixed in equal amounts (100 µL of each type of AuNP-PNA) in a 1.7 mL polypropylene microcentrifuge tube.  DNA linker (various final concentrations) was then added to the mixture.  The total volume of the solution was brought to 1 mL through the addition of 715 µL of hybridization buffer (typically phosphate buffered saline with 0.02% Tween 20, abbreviated as PBST).  Hybridization was allowed to proceed at room temperature (~20°C) in a dry heat block.  In testing the effect of heating on particle assembly, the mixtures were prepared as above, then placed into a 50°C water bath for 5 minutes.  The tubes were then transferred to a room-temperature water bath.  Spectrophotometric and dynamic light scattering measurements were performed as in section 2.3.1.3.

### PNA sequence effects on gold nanoparticle aggregation

DLS measurements of AuNP-PNA mixtures with and without DNA linker addition are shown in Figure 2.6.  The hybridization of DNA linker to both PNA sequences shifted the size distribution of the particles to larger diameters.  However, larger-sized particles were observed in the mixture of the two PNA conjugates.  This indicates self-aggregation of the mixture of PNA-coated particles.  An analysis of the probe sequences revealed that PNA 008 and PNA 010 have some complementarity to each other and could be linking nanoparticles prior to the addition of target DNA.  No differences were observed in the absorbance spectra of the two-probe and one-probe solutions (no red-to-purple color change associated with large particle assemblies), so the PNA probe sequences were not redesigned to eliminate the complementarity.

Figure 2.6. DNA linker hybridization induces AuNP-PNA assembly.  Hydrodynamic diameter of AuNP-PNA mixture (solid line) and mixture with complementary DNA linker (dashed line).

### Effect of DNA linker concentration on AuNP-PNA aggregation

The amount of linker DNA required to form large particle assemblies was evaluated by adding various concentrations of linker to the system.  From Figure 2.7, 100 nM of linker is sufficient to aggregate the particles.  Increasing the DNA concentration to 1000 nM decreased the measured aggregation.  This could be due to the large number of negatively-charged DNA strands repelling each other and preventing the formation of larger particle networks.

Figure 2.7. Effect of DNA linker concentration on the formation of nanoparticle assemblies.  Assemblies were allowed to form in 0.3 M PBST for 18 hours.

### Effect of salt concentration on AuNP assembly

Attempts at forming AuNP-PNA assemblies with the DNA linker in pure water were unsuccessful.  The monovalent salt concentration (NaCl) of the hybridization buffer was varied from 0 to 300 mM, and the solutions were incubated at room temperature (~20°C) until a color change was observed, at approximately 15 hours.  The normalized absorption spectra for AuNP-PNA mixtures with increasing salt concentrations is shown in Figure 2.8.  Over the incubation timecourse, only the solution with 0.3M NaCl showed aggregation.  The same experiment was repeated with the addition of linker DNA (Figure 2.9).  Here, color changes attributed to the formation of DNA-linked nanoparticle assemblies were observed in solutions with at least 0.2 M NaCl.  This result, while unexpected, was confirmed by others (Lytton-Jean, Gibbs-Davis et al. 2009).  At 0.3 M NaCl, however, the AuNP-PNA mixture without DNA linker also turned purple, denoting aggregation.  Presumably the nanoparticles are not fully covered by the PNA strands, and the high salt concentration facilitated the irreversible fusing of the gold cores.

Figure 2.8. Effect of salt concentration on AuNP-PNA self-aggregation.  No self-aggregation was observed for all salt concentrations below 0.3 M NaCl over 15 hours at room temperature.

Figure 2.9. Effect of salt concentration on DNA-linked AuNP assembly.  DNA-templated nanoparticle assembly was successful at salt concentrations at or above 0.2 M NaCl.

### Effect of hybridization time and temperature

In an effort to reduce the time required for the formation of DNA-linked nanoparticles, the effect of raising the hybridization temperature to 50°C was tested.  In the first experiment, this method produced a visible color change (aggregation) within 5 minutes in DNA-containing solutions with 0.2 M NaCl.  The absorbance spectra for these solutions are shown in Figure 2.10.  The control reaction with no DNA did not show any aggregation over 20 h.  Although very encouraging, numerous attempts to replicate this experiment were unsuccessful.  It was concluded that this reaction was too unpredictable to be useful as a field-deployable sensing method.

Figure 2.10. Effect of heating on the formation of DNA-templated AuNP-PNA particle assemblies.  AuNPs conjugated to two different PNA sequences were mixed with 100 nM of DNA linker, then heated to 50°C for 5 minutes in a water bath.  Absorbance spectra were taken immediately after removal from the water bath.

## Chapter summary

In this work, stable and functional gold nanoparticle-PNA (AuNP-PNA) conjugates created using an intermediate surfactant stabilization step are described for the first time. The conjugation procedure is uncomplicated and requires no special modifications to commercially available thiolated PNA sequences. The resulting conjugates retain both the spectral characteristics of gold nanoparticles and the ability of PNA to hybridize to its complement.  However, attempts to control the linkage of AuNPs functionalized with two different PNA sequences by using a synthetic ssDNA sequence resulted in unpredictable nanoparticle aggregation, and was deemed unsuitable for field use.

# FIELD-DEPLOYABLE COLORIMETRIC BIOASSAY FOR THE RAPID AND SPECIFIC DETECTION OF HARMFUL ALGAL BLOOM-CAUSING DINOFLAGELLATES†

## Introduction

A colorimetric bioassay was investigated for the identification and quantification of RNA targets.  Colorimetric sensing is inexpensive, rapid, and easy to perform.  It has the potential to replace laborious identification schemes in field environments for monitoring the spatial and temporal distribution of harmful algal blooms.

________________________________________

†Sections of this chapter are composed in part from excerpts from Duy, J., Smith, R.L., Collins, S.D., and Connell, L.B. “A field-deployable colorimetric bioassay for the rapid and specific detection of ribosomal RNA,” 2012. Biosens. Bioelectron. http://dx.doi.org/10.1016/j.bios.2012.05.039

In this approach, peptide nucleic acid (PNA) probes are used to capture RNA targets, and a micrococcal nuclease digestion is performed to remove all non-target nucleic acids, including single base mismatches flanked by adenines or uracils.  Perfectly-matched PNA-RNA hybrids remained intact and were detected using the symmetrical cyanine dye 3,3'-diethylthiadicarbocyanine iodide (DiSC2(5)).  Assay applicability to complex samples was demonstrated using mixtures containing RNA sequences from two related, harmful algal bloom-causing Alexandrium species.  Target RNA was detected even in mixtures with mismatched sequences in excess of the perfect match.  Hybridization to RNA extracted from cultured Alexandrium cells was also demonstrated, with a detection limit of ~3300 cells.

## Background

The detection of PNA-DNA hybrids using the indicator dye DiSC2(5) was pioneered by Bruce Armitage’s group in the late 1990s (Seifert, Connor et al. 1999; Smith, Olson et al. 1999).  Wittung et al. (1994) reported that commonly-used DNA intercalators such as ethidium bromide, 8-methylpsoralen, and ruthenium complexes with planar aromatic ligands did not bind as expected in PNA-DNA and PNA-PNA hybrids.  However, the minor groove binders distamycin and diamidinophenylindole (DAPI) showed some affinity for PNA-DNA duplexes with central AATA/TTAT base sequence, but no affinity for corresponding PNA-PNA duplexes.  The Armitage group’s studies of cyanine dye interactions with DNA duplexes revealed that certain dyes, such as DiSC2(5) prefer to bind within the minor groove of hybrids with alternating A/T sequences, such as ATATA/TATAT (Seifert, Connor et al. 1999).

The use of DiSC2(5) as an indicator for PNA-DNA hybridization gained popularity in the years following its initial discovery.  Wilhelmsson et al. (2002) utilized the colorimetric assay and the higher thermal stability of perfectly-matched PNA-DNA hybrids to resolve mismatches; PNA probes and DNA sequences with 0, 1, 2, and 3 mismatches were mixed with dye and then heated to 95°C, and the perfectly-matched PNA-DNA sequence exhibited a color change before any of the mismatched sequences.  Komiyama et al. (2003) achieved single-base mismatch discrimination in hybridized DNA by using either nuclease S1 or mung bean nuclease, which both degrade single-stranded nucleic acids, to digest mismatched DNA sequences prior to the addition of DiSC2(5).  The T(PNA)-G(DNA) mismatch was harder to resolve than other types of mismatches, but was identified by increasing the dye concentration.  They verified the applicability of this method in recognizing SNPs by using 80-base asymmetric PCR-amplified (single-stranded) DNA from the apolipoprotein E (apoE) gene.

Another approach for improving the specificity of the dye aggregation on the PNA-DNA template used the inclusion of millimolar concentrations of succinyl-β-cyclodextrin in the reaction mixture (Tedeschi, Sforza et al. 2007).  The use of this method on wild-type and mutant 77-base apoE single-stranded PCR products distinguished a single base (T to C) substitution.  While no defined mechanism for the increased specificity was determined, the authors postulated that the cyclodextrin provided a more favorable environment for the dye due to the presence of negative charges on the molecule, as well as its hydrophobic interior.

## Colorimetric assay optimization

### PNA probe and dye concentrations

Initial experiments with the colorimetric assay were undertaken using synthetic DNA oligonucleotide sequences.  The PNA probe and cyanine dye concentrations used were taken from existing literature (Komiyama, Ye et al. 2003) to verify the color change attributed to PNA-DNA hybridization.  Preliminary tests of PNA-RNA duplex formation using this assay also used these probe and dye concentrations.

Tests revealed that optimizing the target detection limit required a balance between probe, dye, and methanol concentrations.  For instance, the lower limit of detection is set by the probe concentration; experiments with synthetic RNA oligonucleotides showed that targets were reliably detected when the PNA:RNA ratio was at or above 10 (that is, 100 nM target for a 1 µM probe concentration).  The PNA probe concentration was lowered to 1 µM after extensive testing in order to detect a minimal target concentration.  Lower probe concentrations were also tested, but colorimetric signals for these could not be resolved with the NanoDrop ND-1000.

Colorimetric detection using DiSC2(5) is also influenced by the dye concentration relative to concentration of target duplexes (dye:duplex ratio), and by the methanol concentration.  If the dye concentration is too high (relative to PNA-RNA concentration), the signal decreases because more of the dye is in monomer form and absorbs at 650 nm.  A dye concentration that is too low, however, will produce a lower signal-to-noise ratio.  Methanol concentrations that are too high will inhibit dye aggregation in matched PNA-RNA duplexes, as the dye will remain in monomeric form in the solvent.  Dye:duplex ratios from 0.2:1 to 100:1 were tested along with methanol concentrations from 5-30%.  The signal-to-noise ratio was maximized and relatively constant when the dye:duplex ratio was between 10:1 and 20:1, with methanol concentrations less than 20%.  All subsequent testing used a dye:duplex ratio of 10:1, with a methanol concentration of 5% (v/v); that is, for a PNA probe concentration of 1 µM, the dye concentration is 10 µM.

### Target detection from total RNA samples using nuclease digestion

PNA hybridization to synthetic RNA sequences revealed the inherent lack of specificity of the indicator dye (Figure 3.1).  The hydrophobic nature of the dye favored its aggregation even on mismatched PNA-RNA duplexes, presumably since these provide sites that are more hydrophobic than the aqueous buffer.

Figure 3.1. DiSC2(5) non-specifically indicates PNA-RNA hybridization.

Single-base mismatch resolution of PNA-DNA hybrids was reported through the use of S1 nuclease (Komiyama, Ye et al. 2003).  This procedure was attempted for both mismatched PNA-DNA and PNA-RNA hybrids, but both yielded unsatisfactory results: the nuclease digested the perfectly-matched hybrids while mismatched duplexes were unaffected (data not shown).  In addition, the procedure required multiple buffer adjustments and increased hybridization and incubation times (total assay time was greater than 1 hour).  Hence, this approach was not further pursued.

RNase ONE™ ribonuclease (Promega Corp., Madison, WI, USA) was tested to increase the specificity of the assay.  This nuclease digests 3′ of all ribonucleotides (non-base-specific) and could be used without any buffer changes (Brewer, Murray et al. 1992).  Single base mismatches in PNA-RNA hybrids were successfully resolved using this nuclease (Figure 3.2).

Figure 3.2. RNase ONE™ resolves single-base mismatches in PNA-RNA hybrids.  Background-level signals were obtained from mismatched PNA-RNA duplexes after digestion with RNase ONE™.

However, testing with total RNA extracted from Alexandrium cell cultures demonstrated the need for a nuclease that could also eliminate dsRNA.  Micrococcal nuclease (MNase) was then selected as a candidate nuclease due to its ability to digest both single- and double-stranded DNA and RNA with relative non-specificity (Alexander, Heppel et al. 1961; von Hippel and Felsenfeld 1964; Dingwall, Lomonossoff et al. 1981; Hörz and Altenburger 1981; Drew 1984).  In addition, this nuclease is significantly less expensive than RNase ONE™.

## Assay verification with synthetic RNA

### Methods and materials

#### Oligonucleotides and reagents

A PNA sequence complementary to the large-subunit ribosomal RNA of the North American Alexandrium tamarense/fundyense/catanella ribotype, NA1 (Anderson, Kulis et al. 1999) was purchased from PNA Bio, Inc. (Thousand Oaks, CA, USA).  RNA oligonucleotides were purchased from Integrated DNA Technologies (Coralville, IA, USA).  The probe and target sequences are given in Table 3.1.  The PNA probes were resuspended in molecular biology-grade water, while RNA stock solutions were made in The RNA Storage Solution (1 mM sodium citrate, pH 6.4) from Ambion (Austin, TX, USA).  Micrococcal nuclease was purchased from New England BioLabs (Ipswich, MA, USA).  The dye 3,3'-diethylthiadicarbocyanine iodide (dye content 98%) was purchased from Sigma Aldrich (St. Louis, MO, USA) and 2 mM stock solutions were prepared in methanol.

Table 3.1. Oligonucleotide sequences used in Chapter 3.  Mismatches are in bold and underlined.

| Oligonucleotide | Base sequence N-terminus to C-terminus for PNA, 5' to 3' for RNA | Mismatch location Flanking bases |
| --- | --- | --- |
| PNA 024 | EE-GTG CAA CAC TCC CAC C-EE* | N/A |
| A. tamarense RNA (AlexCap′) | GGU GGG AGU GUU GCA C |  |
| C-A mismatch | GGU GGG AAU GUU GCA C | A and U |
| C-C mismatch | GGU GGG ACU GUU GCA C |  |
| C-U mismatch | GGU GGG AUU GUU GCA C |  |
| A-A mismatch | GGU GGG AGA GUU GCA C | G and G |
| A-C mismatch | GGU GGG AGC GUU GCA C |  |
| A-G mismatch | GGU GGG AGG GUU GCA C |  |
| A. ostenfeldii RNA (Aost′) | GGU GAG AUU GUU GCG U | Multiple bases |
| S. endobioticum RNA (Se18s′) | CCG AAC GAG ACA UUG A |  |

* E moieties are solubility enhancers(Gildea, Casey et al. 1998).

#### PNA-RNA hybridization and specificity experiments

Nucleic acid hybridization was performed in 0.2 mL DNase- and RNase-free thin-walled polypropylene PCR tubes (Molecular BioProducts, San Diego, CA, USA) in a total volume of 20 µL.  Note that all quantities given refer to final concentrations in solution.  PNA probes (1 µM) and RNA oligonucleotides (varying concentrations) were mixed in hybridization buffer (50 mM Tris-HCl, 5 mM CaCl2, pH 7.9).  The mixture was kept at 25°C for 5 minutes in a dry heat block, after which micrococcal nuclease (100 Kunitz units) was added and the solution was incubated for a further 5 minutes.  DiSC2(5) in methanol (10 µM) was then mixed in, and absorbance spectra were recorded immediately after dye addition.  Spectrophotometer readings were made using 2 µL of sample with a NanoDrop ND-1000 (Thermo Fisher Scientific Inc., Waltham, MA, USA).

In order to construct a concentration response curve for the target RNA, nucleic acid hybridization was performed as described above, with A. tamarense RNA diluted to final concentrations of 100, 250, 500, 750, and 1000 nM.  Assay specificity for A. tamarense in the presence of competing A. ostenfeldii sequences was assessed with samples containing both target and non-target Alexandrium RNA.  The target A. tamarense RNA concentration was varied (100, 250, 500, 750, and 900 nM), and A. ostenfeldii (non-target) RNA was added to bring the total RNA concentration of each sample to 1000 nM; for instance, 900 nM A. ostenfeldii RNA was added to the 100 nM A. tamarense concentration.

### Results and discussion

#### Dye indication of PNA-RNA hybridization

The dye DiSC2(5) has characteristic absorption peaks at 650 nm and 580 nm in hybridization buffer.  When added to a solution containing the PNA probe alone, the positions of the dye peaks are unaltered (Figure 3.3, solid line).  However, when added to a solution containing PNA-RNA hybrids, a new absorption peak at ~540 nm is observed in the dye spectrum (dashed line).  This shift in peak absorbance (~110 nm) is immediately visible as a color change in the solution from blue (PNA probe only) to purple (PNA-RNA hybrid) (Figure 3.3, insets).  This new peak is attributed to the development of higher-order dye aggregates forming in the minor groove of PNA-RNA hybrids.  In PNA-DNA hybrids, DiSC2(5) aggregation is the result of cooperative binding of face-to-face dye molecules (Smith, Olson et al. 1999).  While slight conformational differences in the double-helical structures of PNA-DNA and PNA-RNA hybrids exist (Eriksson and Nielsen 1996), the position of the induced absorbance peak suggests that the mechanism of DiSC2(5) aggregation is the same for both types of hybrids (Hannah and Armitage 2004; Hannah, Gil et al. 2005; Tomlinson, Frezza et al. 2006).  Since the dye aggregate peak at 540 nm grows at the expense of the monomer peak at 650 nm, a “hybridization signal” can be defined as the ratio of these two values.

One advantage of this method is the rapid response associated with using a PNA probe.  While PNA-RNA hybridization was allowed to proceed for five minutes for protocol reproducibility, hybridization was complete within seconds ((Rose 1993) and unpublished data).  In contrast, most nucleic acid-based detection schemes require much longer incubation periods (Mothershed and Whitney 2006; O’Connor and Glynn 2010).

Figure 3.3. Optical characteristics of indicator dye in the presence of PNA and PNA-RNA hybrids.  The cyanine dye DiSC2(5) in the presence of PNA alone has characteristic peaks at 650 nm and 580 nm (solid line) and produces a blue solution.  PNA-RNA hybrids induce the formation of a blue-shifted peak at ~540 nm in the dye spectrum, with a concomitant decrease in the 650 nm peak (dashed line), and the solution changes color to purple.  PNA and RNA concentrations used were 1 µM, while DiSC2(5) concentration was 10 µM.  Absorbance values were taken at 25°C with a Beckman-Coulter DU-640 spectrophotometer, 1 cm pathlength.

#### Single-base mismatch discrimination of A. tamarense-based target RNA sequences

The ability of the assay to recognize point mutations in RNA was evaluated by constructing artificial mismatches in the center of the target sequence (Table 3.1).  The hybridization signal from perfectly-matched hybrids (Figure 3.4, “PM”) was unaffected by micrococcal nuclease digestion.  No mismatches were resolved without the addition of nuclease (Figure 3.4, unshaded bars), but point mutations between adenines and uracils could be reduced to background levels after nuclease treatment. The C(PNA)-C(RNA) mismatch was an exception, and produced a signal only slightly above background.  Mismatches flanked by guanines, however, could not be completely digested away in the experiment timeframe of 5 minutes (Figure 3.4, shaded bars).  These findings are consistent with prior work on micrococcal nuclease: the enzyme shows a marked preference for digesting AT-rich regions, and its action is much reduced when confronted with guanines and cytosines (von Hippel and Felsenfeld 1964; Dingwall, Lomonossoff et al. 1981; Hörz and Altenburger 1981; Drew 1984).  It should also be noted that the A(PNA)-N(RNA) mismatches have been reported as some of the least destabilizing types to a PNA-RNA duplex (Jensen, Ørum et al. 1997).  These factors should be considered when using this method to identify point mutations, especially since the binding characteristics of each probe sequence must be empirically determined (Pozhitkov, Noble et al. 2006).

Other investigators have used the single-strand-specific endonuclease S1 to digest all types of single-nucleotide mismatches in PNA-DNA hybrids (Komiyama, Ye et al. 2003; Ye, Miyajima et al. 2007).  However, this approach requires additional steps such as buffer change/salt addition, a lengthy incubation time, and a stop solution.  Furthermore, the applicability of S1 nuclease is limited to unstructured nucleic acid sequences, which precludes its use with genomic DNA or RNA (Ren, Zhou et al. 2004).  Micrococcal nuclease completely digests single- and double-stranded nucleic acids without degrading PNA-RNA hybrids.  Hence, target sequences in environmental samples do not need to be pre-hybridized or amplified, as the nuclease will remove non-target DNA and RNA from the mixture.

Figure 3.4. Micrococcal nuclease digestion of single-base mismatches.  Perfectly-matched and single-base mismatched RNA sequences showed high hybridization signals to the PNA probe (light bars).  The effect of micrococcal nuclease on these hybrids is denoted by the shaded bars.  Point mutations sandwiched by adenines and uracils (C-N base mismatches, inside dashed box) were successfully differentiated from the perfect match.  Guanine-flanked mismatches (A-N mismatches, inside dotted box) showed reduced hybridization signals.  PM indicates the perfectly-matched PNA-RNA duplex, while base mismatches are given by N(PNA)-N(RNA).  Nucleic acid concentrations used were 1 µM, while DiSC2(5) concentration was 10 µM.  Absorbance values were taken at 25°C with a NanoDrop ND-1000 spectrophotometer.

#### RNA concentration dependence and discrimination against A. ostenfeldii RNA

The utility of this method in quantifying the amount of target RNA present, even in mixtures with similar, non-target sequences, is shown in Figure 3.5.  The hybridization signal A540/A650 was used to construct concentration curves for A. tamarense RNA alone, and for the same RNA mixed with A. ostenfeldii sequences.  Without nuclease digestion, the signals from the mixed samples were uniformly high (hollow circles), as the A. ostenfeldii RNA sequence has only four base mismatches to the probe (Table 3.1), and the hybridization was performed at room temperature (25°C).  Nevertheless, nuclease treatment of the mixed-RNA samples (filled circles) yielded hybridization signals that were virtually identical to those from A. tamarense RNA only (triangles).  RNA mixtures containing 250 nM and 500 nM of the target sequence produced lower hybridization signals than the target-only samples.  This is likely due to the low hybridization temperature; presumably the large numbers of A. ostenfeldii RNA sequences competitively bound the PNA probes.  In that case, the mismatched RNA, as well as unbound target strands, would be digested by the nuclease.

Figure 3.5. Quantitative detection of A. tamarense RNA in mixtures with A. ostenfeldii RNA.  A concentration curve with target RNA only (triangles) was constructed.  The same concentrations of target RNA were mixed with interfering non-target RNA, allowed to hybridize to the PNA probe, and were treated (filled circles) or not treated (hollow circles) with micrococcal nuclease.  The plotted lines are intended to guide the eye.  The probe concentration used was 1 µM, and the dye concentration was 10 µM.  Absorbance values were taken at 25°C with a NanoDrop ND-1000 spectrophotometer.

## Testing with total RNA extracted from Alexandrium cell cultures

### Materials and methods

#### Reagents and supplies

Alexandrium cell cultures were obtained from the Provasoli-Guillard National Center for Marine Algae and Microbiota (East Boothbay, Maine, USA) and from Dave Kulis.  Guanidinium lysis buffer (1.5 M GuSCN, 50 mM Tris, 15 mm EDTA, 2% sarkosyl and 0.2% SDS (v/v), pH 8.9) was kindly supplied by Amber Bratcher.  Filter paper (5.0 µm pore size, 25 mm diameter, SVLP membrane was obtained from Millipore (EMD Millipore, Billerica, MA, USA).  Disposable plastic syringes with Luer-Lok™ tips were purchased from BD Biosciences (Franklin Lakes, NJ, USA).  Sterile filter units (0.45 µm pore size, 25 mm diameter) were obtained from Thermo Fisher Scientific (Waltham, MA, USA).  The RNA Clean & Concentrator-25™ Kit was purchased from Zymo Research Corp. (Irvine, California, USA).  The FlashGel™ System, FlashGel™ 1.2% agarose RNA cassettes, FlashGel™ RNA Marker, and formaldehyde sample buffer were obtained from Lonza Group Ltd. (Basel, Switzerland).  RNA Fragmentation Reagents were purchased from Life Technologies Corp. (Grand Island, New York, USA).

#### Alexandrium cell culture growth conditions

Cultures were grown in flasks filled with nutrient-replete L1 media (Guillard and Hargraves 1993) made with 0.1 µm-filtered and autoclaved seawater from Boothbay Harbor, Maine, USA.  Cultures were maintained at 15°C incubator with a 12h:12h light:dark cycle.

#### Total RNA extraction from algal cells and RNA quality assessment

RNA was extracted from algal cultures in exponential growth phase.  Prior to extraction, 1 mL of cell culture was mixed with 10 µL of Lugol’s (iodine) solution in a 1.5 mL polypropylene microcentrifuge tube.  The fixed cells were then placed into a Sedgewick-Rafter counting chamber, and at least 100 squares or 400 cells were counted to extrapolate the cell concentration of the culture.  The detailed protocol is found in Appendix D.

The extraction procedure used was modified from existing protocols for sandwich hybridization assays (Scholin, Marin et al. 1999; Jones, Preston et al. 2008).  The protocol is detailed in Appendix E. Briefly, a disposable plastic syringe was used to collect algal cells from the culture flask.  The solution was passed through a filter unit containing a 5.0 µm filter membrane, and the collected cells were carefully transferred to a 2.0 mL polypropylene tube containing at least 500 µL of the GuSCN lysis buffer.  The tube was vortexed briefly to suspend the cells in the lysis buffer, then placed into a heat block set to 85°C.  The tube was finger-vortexed (flicked) for 6 times after a 2.5-minute incubation, then immediately returned to the heat block and incubated for a further 2.5 minutes.  The solution was then removed from heat and allowed to cool briefly at room temperature, then the lysate was passed through a sterile 0.45 µm filter assembly to remove cell debris.  The extracted RNA was then purified with the RNA Clean & Concentrator-25™ Kit following the manufacturer’s protocol, with the addition of one extra 400-µL wash step (in consultation with Zymo Research Corp. technical support) to improve lysis buffer removal from the samples.  RNA was eluted in The RNA Storage Solution and kept on ice until use.

Alternatively, RNA was extracted from filtered cells using the ZR Fungal/Bacterial RNA Miniprep™ Kit (Zymo Research Corp.), following manufacturer’s protocol.  RNA was eluted in The RNA Storage Solution and kept on ice until use.

RNA quantity and quality were spectrophotometrically assayed with a NanoDrop ND-1000.  In addition, the quality of extracted RNA was determined by agarose gel electrophoresis using the FlashGel™ RNA System.

#### Total RNA fragmentation

Intact total RNA was fragmented using the RNA Fragmentation Reagents.  Briefly, RNA (up to 10 µg) was diluted with molecular biology-grade water to 9 µL in a 0.2 mL thin-walled polypropylene tube.  The 10X RNA Fragmentation Reagent (1 µL) was added and the solution was mixed, then heated to 95°C for 1 minute.  The Stop Solution (1 µL) was mixed in immediately, and samples were kept on ice until use.

#### PNA hybridization to fragmented total RNA

Nucleic acid hybridization was performed as described in section 3.4.1.2, with modifications to the hybridization time and temperature conditions.  After the PNA and RNA were mixed (18 µL combined volume), the mixture was heated to 90°C for 3 minutes to denature the nucleic acids, then hybridization was allowed to proceed at 60°C for 5 minutes.  Micrococcal nuclease was then mixed in, and the solution was incubated at 37°C for 5 minutes.  DiSC2(5) dye in methanol was then added immediately, and solution absorbance spectra were obtained with the NanoDrop ND-1000.

### Results and discussion

#### Total RNA fragmentation for hybridization to PNA probes

Multiple nucleic acid hybridization studies have shown that the fragmentation of total nucleic acids is necessary for successful probe-target duplex formation (Small, Call et al. 2001; Nelson, Liles et al. 2002; Mehlmann, Townsend et al. 2005; Metfies, Huljic et al. 2005; Liu, Guo et al. 2007).  The same result was obtained with total RNA extracted from Alexandrium cell cultures.  Figure 3.6 shows hybridization signals obtained from intact and fragmented A. tamarense (isolate PW06) total RNA, to a perfectly complementary PNA probe.  The results suggest that PNA-RNA hybridization is hindered when using intact RNA as both nuclease-treated and undigested samples show much lower signals than those from PNA incubated with fragmented RNA.  RNA folding simulations showed that the target region is completely bound to a complementary strand.  Most likely the higher-order structures formed by the intact RNA molecules prevent efficient probe hybridization in the short time frame required for the assay.

Figure 3.6. RNA fragmentation improves hybridization to PNA.  Enhanced hybridization signals were obtained when total RNA extracted from A. tamarense (isolate PW06) was fragmented prior to hybridization.

The size distribution of the fragmented RNA is shown in Figure 3.7.  The size distribution of the fragments is between 28-300 bases, as determined using the gel analysis tool of ImageJ (Abramoff, Magelhaes et al. 2004).  The fragmentation protocol used in subsequent experiments was selected to reflect this size distribution, and the optimal incubation parameters were found to be 1 minute at 95°C using The RNA Fragmentation Reagents.

4

3

2

1

Figure 3.7. GuSCN lysis-induced RNA fragmentation.  The lysis protocol used for Alexandrium nucleic acid extraction induces RNA shearing, with a size distribution between 28-300 bases long.   Lane 1: New England BioLabs Low Range ssRNA Ladder (50-1000 bases); 2: Zymo Research small-RNAmall-RNA™ Marker (21-29 bases); 3: A. ostenfeldii CCMP1773 RNA; 4: A. fundyense GTCA 28.

#### Optimization of hybridization conditions

The neutral backbone of the PNA molecule allows for the formation of stable nucleic acid hybrids within a very short timeframe of seconds to minutes (Rose (1993) and unpublished data).  Unlike DNA and RNA, however, only one generalized study on the predicted melting temperature (Tm) of PNA-containing hybrids have been undertaken (Giesen, Kleider et al. 1998).  The predicted duplex Tm is defined as the temperature at which the amounts of base-paired duplexes and unpaired single strands are in equilibrium, and is used as starting point for determining the optimal hybridization temperature.  The empirically derived a formula for the calculation of the Tm of PNA-DNA hybrids from Giesen et al. (1998) is given as:

Tm=20.79+0.83*Tm,nnDNA-26.13*fpyr+0.44*length

where Tm,nnDNAis the calculated melting temperature for a corresponding DNA-DNA duplex of the same sequence, using the nearest-neighbor model (SantaLucia, Allawi et al. 1996), fpyr is the fraction of pyrimidines (cytosines and thymines) in the sequence, and length is the number of bases in the strand.

Using this formula, the predicted melting temperature for the Alexandrium-specific PNA probe to a complementary DNA target is 80°C.  PNA-RNA hybrids of the same sequence have empirical Tm values ~4°C above that of PNA-DNA duplexes (Egholm, Buchardt et al. 1993; Jensen, Ørum et al. 1997), so the estimated Tm for the desired hybrid is 84°C.  Preliminary studies with fragmented total RNA using lower incubation temperatures (25-45°C) yielded very low hybridization signals and were not further pursued.  Figure 3.8 shows the effect of three higher hybridization temperatures on the formation of PNA-RNA duplexes.

Figure 3.8. Hybridization temperature optimization for Alexandrium RNA.  Fragmented total RNA from A. tamarense PW06 was used.  PNA-RNA mixtures were heated to 90°C for 3 min, then the temperature was lowered for hybridization (5 min).

Fragmented total RNA from A. tamarense isolate PW06 was used in the assay.  No significant differences were observed between the hybridization efficiencies at the three temperatures shown, and a temperature of 60°C was chosen for all subsequent experiments.  Duplex formation was also tested as a function of incubation time, but no differences in signal were observed for hybridization times from 5-60 minutes.  To decrease total assay time, a hybridization duration of 5 minutes was implemented for successive tests.

#### Discrimination between North American A. tamarense/fundyense/catanella ribotype and A. ostenfeldii

The applicability of the assay to RNA extracted from cell cultures was tested with the related, co-occurring dinoflagellates A. tamarense (PW06) and A. ostenfeldii (CCMP1773).  Detection of targets was possible at the 5 ng/µL RNA concentration, equivalent to ~3300 cells, while discrimination between the two species was achieved at 10 ng/µL, about 6700 cells (Figure 3.9).

Figure 3.9. Discrimination between related Alexandrium species.  PNA-RNA were mixed, denatured at 90°C for 3 min, hybridized at 60°C for 5 min, 10 Ku MNase digestion at 60°C for 5 min.

## Chapter summary

This chapter reports the first demonstration of using the cyanine dye DiSC2(5) to detect PNA-RNA hybrids in a rapid and highly specific colorimetric assay.  Peptide nucleic acid (PNA) probes targeted to the toxigenic North American Alexandrium dinoflagellate species were used to capture complementary RNA sequences.  PNA hybridization to RNA targets was almost instantaneous, and a short digestion with micrococcal nuclease eliminated all other non-target nucleic acids, including single nucleotide mismatches sandwiched by adenines or uracils.  PNA-RNA duplexes were indicated by the color change (from blue to purple) of the symmetrical cyanine dye 3,3'-diethylthiadicarbocyanine iodide (DiSC2(5)).  This color change was due to the formation of a blue-shifted dye aggregate absorbance peak at the expense of the original dye monomer peak, enabling the amount of target RNA in the sample to be quantified by the ratio of these absorbance values.  Preliminary validation of the applicability of the assay to complex samples was carried out with mixtures containing both A. tamarense (target) and A. ostenfeldii (non-target) RNA oligonucleotides.  A. tamarense RNA could be detected even in the presence of excess A. ostenfeldii sequences, indicating the suitability of this method in detecting unamplified targets in environmental samples.

Assay hybridization parameters were optimized for use with total RNA extracted from cell cultures.  Notably, extracted RNA must be fragmented to improve hybridization signals, and the hybridization temperature was empirically determined to be about ~25°C below the estimated Tm.  RNA extracted from the co-occurring, morphologically indiscernible species A. tamarense and A. ostenfeldii were detectable at ~3300 cells, and the species were differentiated at ~6700 cells within 15 minutes.

# DEVELOPMENT OF A PORTABLE INSTRUMENT FOR FIELD DEPLOYMENT OF COLORIMETRIC BIOASSAY‡

## Introduction

A portable two-color absorbance detector for use with the DiSC2(5) colorimetric hybridization assay was designed and built using off-the-shelf electronic components and a custom-designed test cell.  A microcontroller module was programmed to automate the measurement of solution absorbance values and was used with a digital display for user output.   The response of the colorimeter was modeled using available optical data, and the model accuracy was confirmed using absorbance data from a side-by-side test with a benchtop spectrophotometer.  Normalized concentration curves from both instruments were compared, and the colorimeter response was found to be equivalent to that of the spectrophotometer.  This inexpensive device has the potential to be used in field environments for quantifying PNA hybridization to target RNA sequences using the DiSC2(5)-indicated colorimetric assay.

________________________________________

‡Sections of this chapter are composed in part from excerpts from:

Duy, J., Smith, R.L., Collins, S.D., and Connell, L.B. “A field-deployable colorimetric bioassay for the rapid and specific detection of ribosomal RNA,” 2012. Biosens. Bioelectron. http://dx.doi.org/10.1016/j.bios.2012.05.039

and

Duy, J., Smith, R.L., Collins, S.D., and Connell, L.B. “Low-cost colorimeter development for the field-based detection of harmful algal blooms,” 2011. OCEANS 2011, MTS-IEEE OCEANS ’11 Kona, p. 1-5.

## Methods and materials

### Handheld colorimeter design

A prototype two-wavelength colorimeter was assembled primarily from off-the-shelf parts.  LED light sources were obtained from Vishay Semiconductors (Malvern, PA) and were selected to encompass the dye monomer (red, Vishay TLDR5800, λpeak = 652 nm) and dye aggregate (green, Vishay TLHP5800, λpeak = 556 nm) wavelengths.  A high-output-current, dual-channel 10-bit programmable digital-to-analog converter (MAX5550, Maxim, Sunnyvale, CA, USA) was used as a constant-current source for the LEDs.  The photodetector chosen for this system was a programmable light-to-frequency converter (TSL230BR-LF, TAOS, Inc., Plano, TX).  User input and display were handled with a 3.2″ TFT LCD module (ITDB02-3.2S, ITead Studio, Shenzen, China).  For side-by-side instrument evaluation, 1-cm pathlength ultra-micro cuvettes (BrandTech Scientific, Inc., Essex, CT, USA) were used to fit both the colorimeter and a Beckman Coulter DU-640 benchtop spectrophotometer (Brea, CA, USA).  A cuvette holder with integrated positioning slots for the LEDs and photodetector was designed and sent out for 3D printing (Shapeways, New York, NY, USA).  An Arduino Mega 2560 was used to control all devices, and for data collection and signal processing.  The colorimeter block diagram is shown in Figure 4.1.  The current embodiment of the device was designed and assembled by Corey Hirn for increased portability and water-resistance.  An inside view of the instrument with major components labeled is pictured in Figure 4.2, while the operating version, in a water-tight Pelican 1060 Micro Case (Pelican Products, Inc., Torrance, CA, USA), is pictured in Figure 4.3.

Figure 4.1. Colorimeter block diagram.  A functional block diagram of the colorimeter is shown.  A microcontroller unit is used to run a touchscreen display and optical measurement components.  A simplified drawing of the cuvette holder (side view) is included, with the positions of the LEDs (light emitting diodes) and PD (photodetector) relative to the liquid sample illustrated.

Figure 4.2. Colorimeter unit, inside view.  The current version of the device is housed in a 3D-printed plastic housing designed for water-resistance (white polymer).  The major components of the instrument labeled above are: A) Arduino Mega 2560 microcontroller module, underneath B) a custom PCB shield containing C) the MAX5550 constant-current source; D) a custom 3D-printed cuvette holder holding the E) green and red Vishay LEDs and F) TAOS, Inc. light-to-frequency photodetector; and G) the ITead LCD module for user input and display.  The holder for a 9-V battery is on the back side of the case (not pictured).

Figure 4.3. Field-ready colorimeter in water-tight case.  The working colorimeter is shown inside a Pelican 1060 Micro Case (external dimensions approximately 9.4″ x 5.6″ x 2.6″).  The main screen of the program is shown.

### LED testing

The spectral characteristics of the two Vishay LEDs were obtained with the kind assistance of Dr. Nuri Emanetoglu.  Briefly, an Ocean Optics USB650 spectrometer controlled with SpectraSuite software (Dunedin, FL, USA) was used to record the emission spectra of the light sources.  A Keithley 2400 Broad Purpose SourceMeter (Keithley Instruments Inc., Cleveland, OH, USA) was used to control the current into the sources.  The instrument was used as a constant-current source to drive the LEDs; for the green LED, IF = 25 mA and IF = 1 mA for the red LED to match the output characteristics under typical colorimeter working conditions.

### Colorimeter comparison with a benchtop spectrophotometer

PNA probes (1 µM) and RNA (varying concentrations) were mixed in hybridization buffer and incubated for 5 minutes at 25°C.  Micrococcal nuclease (100 Kunitz units) and DiSC2(5) dye (10 µM) were added as in section 3.4.1.2.  The reaction volume of each sample was 200 µL.  Absorbance values from 220-750 nm were collected using the Beckman-Coulter DU-640 spectrophotometer.

For colorimeter testing, the photodetector was set to 10X sensitivity and 2X frequency scaling (Inc. 2006).  The green LED was driven with the maximum forward current, IF = 25 mA.  For instrument setup (software-controlled), the irradiance measured by the photodetector from the green LED through a blanking solution was recorded, and the forward current for the red LED was adjusted to produce the same response (typically IF ≈ 1 mA).

For sample reading, each LED was powered for 1 second and the photodetector output was recorded using pulse accumulation for that duration.  An off time of 20 ms was introduced before the next LED was turned on to allow the photodetector signal to return to zero.  Output values displayed to the user reflect two successive reads that are within 0.1% of each other.  Each liquid sample was read three times consecutively to determine instrument reproducibility.

### Data collection and analysis

Absorbance data from the spectrophotometer were used directly to obtain the raw hybridization signal, by taking the ratio of the absorbance peak magnitudes for the dye aggregate (λ = 540 nm) and monomer (λ = 650 nm):

Equation 4.1

For colorimeter-based measurements, the “blank” output frequency, Iblank, of each LED was recorded.  The absorbance of each sample was then calculated as:

Equation 4.2

The raw hybridization signal, taken from the calculated absorbance values given in Equation 4.2, was defined as:

Equation 4.3

Normalization of the hybridization signals was performed to enable comparisons between different platforms.  The raw signals calculated in Equation 4.3 were transformed to a 0-to-1 scale using the following equation:

Hnorm=Hsample,raw-Hmin,rawHmax,raw-Hmin,raw                                Equation 4.4

Data analysis was performed using OriginPro 8.6 (OriginLab Corp., Northampton, MA, USA).

## Results and discussion

### Colorimeter design considerations

#### Liquid sample holder selection

One of the potential advantages of this colorimetric assay is a low per-test cost, achieved by reducing reagent consumption.  Preliminary tests were thus performed using rectangular glass capillary tubing (Vitrocom, Mountain Lakes, NJ, USA, 0.3-1 mm liquid pathlength).  However, this approach proved to have several disadvantages: 1) the capillaries had to be individually cut and sealed at one end prior to use; 2) the capillaries had to be silanized (Seed 2001) to prevent nucleic acid adsorption to the chamber walls; 3) the shorter pathlength decreased assay sensitivity.  Because of these drawbacks, further testing was conducted with a 1-cm-pathlength cuvette.  The BrandTech ultra-micro cuvette, with a minimum fill volume of 70 µL, was selected to minimize sample volume requirements.  In addition, this conventional cuvette size could be used for side-by-side comparison with laboratory spectrophotometers.

#### Light sources

The illumination sources used for this application must encompass the dye monomer and peak absorbance wavelengths, at 650 nm and 540 nm, respectively.  Light emitting diodes (LEDs) were chosen for their (relatively) small size, low cost, low power consumption, and fairly narrow bandwidth.

While red LEDs of considerable brightness were easily found in the wavelength range of interest, corresponding green LEDs were very difficult to obtain.  Initial experiments were conducted with Stanley 1211C BRBG bi-color surface-mount LEDs, with peak wavelengths at 558 nm (green) and 647 nm (red) (Stanley Electric Co. 2010).  This bi-color package does not require the alignment of both sources to the photodetector, has a compact footprint, and produced detectable signals through the 1-mm rectangular capillary tube.  However, the low luminous light intensities of the LEDs (11.7 mcd for the red and 2.4 mcd for the green) and wide angular beam spread made it incompatible with the 1-cm pathlength design that was eventually adopted.

Two discrete high-power LEDs from Vishay Semiconductors were used in the current colorimeter design (red, Vishay TLDR5800, λpeak = 652 nm and green, Vishay TLHP5800, λpeak = 556 nm).  The emission spectra for these sources are shown in Figure 4.4.  These are standard T-1¾, 5 mm LED packages with considerably higher luminous intensities (2500 mcd for red, 25 mcd for green) and smaller beam spread (±4°).  Again, however, the green LED is much less powerful than the red, and possibly greater assay sensitivity may be achieved with a green laser diode.

Figure 4.4. Spectral characteristics of green and red LEDs.  Emission spectra were measured for the green and red Vishay LEDs.  Both sources were powered with a source –measure unit on constant-current output, with IF(green) = 25 mA and IF(red) = 1 mA.

#### Photodetector selection

The TAOS, Inc. TSL230BR Programmable Light-to-Frequency converter was selected as the photodetector for several reasons: 1) large output range without the need for a separate amplifier; 2) high signal resolution not limited by 5-V digital I/O ports; 3) ease of interfacing with a microcontroller; and 4) single-supply, low power operation.  The silicon photodiode response curve was provided by the manufacturer and is shown in Figure 4.5.

Figure 4.5. Normalized photodiode spectral response.

#### Microcontroller selection

The Arduino Uno (Smart Projects, Italy) based on the ATmega320 microprocessor was originally chosen as the microcontroller for 1) ease of software development due to the C-like programming environment and 2) ease of prototyping because of its popularity with electronics hobbyists (many pre-designed components are available for the Arduino).  The controller was upgraded to the Arduino Mega 2560 (based on the more powerful ATmega2560 processor) for use with the touchscreen/display which requires more I/O ports.

#### Touchscreen/display

A touchscreen/display unit was chosen to eliminate the need for other input components such as individual buttons or a separate keypad.  In addition, a single module is easier to integrate into a water-resistant case design.  The particular display, a 3.2″ TFT LCD (ITDB02-3.2S, ITead Studio) was selected for ease of interfacing with the Arduino, and for its large display size relative to cost.

### Comparison of colorimeter and spectrophotometer response

Samples of A. tamarense RNA at concentrations of 100, 250, 500, 750, and 1000 nM were prepared as described in section 3.4.1.2.  The hybridization signal, H, for each sample was calculated for the prototype colorimeter and with the Beckman-Coulter DU-640, using identical samples.  The signal from the blank (no RNA) sample was subtracted from the response at each concentration, and graphed in Figure 4.6.

Figure 4.6. Raw hybridization signals from spectrophotometer and colorimeter.  Hybridization signals were calculated for the spectrophotometer using Equation 4.1, while colorimeter responses were derived from Equation 4.2 and Equation 4.3.  For each concentration, triplicate samples were run.

Both instruments show increasing, concentration-dependent hybridization signals; however, the colorimeter response is lower for all samples.  This discrepancy can be attributed to a) the spectral widths of the individual light sources (Figure 4.4) and b) the wavelength-dependent response of the photodetector (Figure 4.5).  Using these assumptions, the colorimeter response was modeled as:

Rred,green=Ared,greenλ lred,greenλ pλ dλ 		Equation 4.5

Where Aλ is the measured absorbance, lλ is the LED intensity, and pλ is the photodetector response at each wavelength.  The integrand R represents the actual absorbance spectrum as seen through the limitations of the colorimeter.  The modeled hybridization signal can then be calculated as:

Hmodel=RgreenRred 			Equation 4.6

The actual and modeled colorimeter responses are shown in Figure 4.7.  Agreement between the responses indicates that the model is valid and that the differences between the spectrophotometer and colorimeter responses to the same samples are indeed due to the spectral width of the LEDs and to the non-uniform photodetector response at the illuminated wavelengths.

Figure 4.7. Actual and modeled colorimeter responses.  Actual colorimeter values (gray bars) were compared to theoretical colorimeter absorbance values modeled by Equation 4.5.

To enable comparison between the instruments, the concentration curves obtained from each device were normalized to span the range from 0 (no RNA) to 1 (1000 nM RNA), as given in Equation 4.4.  A Pearson’s correlational analysis was performed on the two sets of data (Figure 4.8).  At a significance level of p = 0.05, the correlation coefficient r = 0.99935: the colorimeter response is virtually identical to that of the spectrophotometer for the same samples.  This indicates that the inexpensive and portable colorimeter can replace a benchtop instrument for out-of-laboratory hybridization assays indicated by DiSC2(5).

Figure 4.8. Comparison of colorimeter performance to benchtop spectrophotometer.  The hybridization signals obtained from the colorimeter and from a Beckman-Coulter DU-640 spectrophotometer (see text for equations) were compared on a scatter plot, with a line drawn through r = 1.  The correlation between the signals obtained, at a significance level of 0.05, is r = 0.99935.  Each point represents three independent samples at the RNA concentration tested.

## Chapter summary

A handheld absorbance detector for use with the DiSC2(5) colorimetric hybridization assay was designed and built from commercially-available electronic components.  The performance of the device was compared to that of a benchtop spectrophotometer using identical samples, and the responses from both instruments were found to be comparable.  Raw signals from both instruments showed that signal discrepancies were due to both the spectral width of the light sources and the wavelength-dependent photodetector characteristics, but that this did not affect the response over the concentrations tested.  This inexpensive device has the potential to be used in field environments for quantifying PNA hybridization to target RNA sequences using the DiSC2(5)-indicated colorimetric assay.

# DETECTION OF THE FUNGAL PHYTOPATHOGEN SYNCHYTRIUM ENDOBIOTICUM USING DiSC2(5)-INDICATED PNA HYBRIDIZATION

## Introduction

Synchytrium endobioticum is a parasitic fungus considered the most important potato pathogen worldwide.  Catastrophic impacts to agriculture and the economy can result if the fungus is spread to uninfested areas, as it is very infective and can survive in spore form for upwards of 70 years.  To date, the identification of S. endobioticum relies on visual identification by highly trained personnel.  In this work we tested a rapid hybridization method using peptide nucleic acid (PNA) probes and a symmetrical cyanine dye for the detection of S. endobioticum from potato wart samples.  Target DNA and RNA were extracted from infected wart tissue, and a segment of the 18S ribosomal RNA gene was sequenced for several S. endobioticum pathotypes.  Sequence information was used to design a both a specific PCR primer set and PNA probe.  The primer set identified S. endobioticum from all available samples and did not show any cross-reactivity to the potato host or to related chytrids.  The PNA probe was used in a rapid colorimetric hybridization assay that detected target RNA down to 10-17 mole within 15 minutes of total hands-on time.  This simple colorimetric technique has the potential to be used for rapid testing in non-laboratory settings with untrained staff to identify cases of potato wart and to prevent the dissemination of S. endobioticum spores.

## Collection of potato wart DNA and RNA

Synchytrium endobioticum, as a USDA select agent, is not permitted into the country because of its potentially catastrophic effects on the country’s potato industry, currently valued at $4 billion per year in 2011 (2012).  The presence of potato wart fungus (disease or resting spores) can close a field to agricultural cultivation for a minimum of 10 years, during which time no profits are realized.  Thus, all sample collection and nucleic acid extraction were performed outside of the US, at the Canadian Food Inspection Agency (CFIA) in St. John’s, Newfoundland, during a 4-day laboratory expedition.  Potato wart is endemic in Newfoundland, and the CFIA Plant Health Laboratory maintains a plot infested with S. endobioticum in order to test newly-developed potato varieties (cultivars) for resistance to the fungus.  They are responsible for testing for suspected incidences of potato wart disease from potato samples sent in by home growers and commercial producers.  They also perform soil tests on quarantined fields to determine if S. endobioticum spores are present in the plots.

Fresh wart samples were not available during the experimental timeframe, hence preserved (dried or frozen) wart samples from previous years were used.  Sample pre-treatment (rehydration or defrosting), spore liberation, and nucleic acid extraction methods were all evaluated and adjusted on-site as no samples were available for testing at the US laboratory.  The CFIA also did not have any protocols related to the extraction of genetic material, as S. endobioticum spores are identified visually.

The main purpose of the trip was to collect as much RNA as possible for testing with the colorimetric bioassay.  The performance of several RNA extraction kits was compared for lysis efficienty and extracted RNA quality (spectrophotometrically determined), and this kit was used for all subsequent samples.  However, the quality and quantity of the RNA samples was inadequate for protocol development and testing.  Genomic DNA was also collected for sequencing, but proved to be invaluable in preliminary testing and protocol refinement as it was used to synthesize in vitro transcribed 18S rRNA.  A simplified workflow is presented below (Figure 5.1), with detailed information available in the chapter sections given.

Figure 5.1. Experimental workflow for potato wart samples.  DNA and RNA were extracted from potato wart samples, then subject to different assessments, as detailed in the indicated chapter sections (in red).

DNA and RNA quantification and quality assessments were performed upon return to the US as limited equipment were available for testing these at the CFIA laboratory.  Spore counts (for calculating D/RNA content per cell) were attempted at the host laboratory, but were uniformly unsuccessful, highlighting both the high level of expertise of the CFIA personnel and the need for more accessible identification methods.

## Methods and materials

### Oligonucleotides and reagents

All PNA sequences used in this study were purchased from PNA Bio (Thousand Oaks, CA, USA). DNA primers and RNA oligonucleotides were purchased from Integrated DNA Technologies (Coralville, IA, USA). The probe and target sequences are given in Table 5.1. The PNA probes were resuspended in molecular biology-grade water, while DNA stock solutions were made in nuclease-free 10 mM Tris buffer, pH 8 from Ambion (Austin, TX, USA). Formamide (100%) was obtained from Thermo Fisher Scientific (Waltham, MA, USA).  Micrococcal nuclease and 10X micrococcal nuclease reaction buffer (500 mM Tris-HCl, 50 mM CaCl2) were obtained from New England BioLabs, Inc. (Ipswich, MA, USA).  The dye 3,3'-diethylthiadicarbocyanine iodide (dye content 98%) was purchased from Sigma Aldrich (St. Louis, MO, USA) and 2 mM stock solutions were prepared in methanol.

Table 5.1. Oligonucleotide sequences used in Chapter 5.  The first entry is the PNA probe, while primers are all DNA.  The T7 promoter sequence (for in vitro transcription of PCR products) is underlined.

| Oligonucleotide | Base sequence N-terminus to C-terminus for PNA, 5′ to 3′ for DNA | Description |
| --- | --- | --- |
| PNA 033 | ATT GTT CCC GTC ACT A | S. endobioticum 18S rRNA-specific PNA probe, starts at position 1124 (relative to S. cerevisiae) |
| PNS1 | CCA AGC TTG AAT TCG TAG TCA TAT GCT TGT C | 18S rDNA universal eukaryotic forward primer |
| T7-PNS1 | TAA TAC GAC TCA CTA TAG GG CCA AGC TTG AAT TCG TAG TCA TAT GCT TGT C | T7 promoter-appended 18S rDNA forward primer |
| NS8 | TCC GCA GGT TCA CCT ACG GA | 18S rDNA universal eukaryotic reverse primer |
| Se18S816f | CTA TGC CAA CTA CGG ATC GGA CG | S. endobioticum-specific 18S rDNA forward primer |
| Se18S1136r | GTC CCT CTA AGA AGT GCT ACA AGG AAC AAA | S. endobioticum-specific 18S rDNA reverse primer |

### DNA and RNA isolation from S. endobioticum wart samples

Frozen and dried wart samples obtained from an infested plot in Avondale, Newfoundland, Canada were kindly provided by the CFIA Plant Health Laboratory in St. John’s, Newfoundland.  Frozen wart samples were thawed at room temperature, while dried wart samples were rehydrated at room temperature for either ~10 minutes or overnight.  The wart was then ground to a uniform slurry with a mortar and pestle, then rinsed with water through a stack of metal sieves (850, 250, 150, 106, 75, and 25 µm mesh sizes).  The final (25 µm) fraction was collected and suspended in enough water to rinse all material from the sieve.  This fraction was used for kit-based DNA extraction.  The kits tested were the ZR Fungal/Bacterial DNA Miniprep Kit (Zymo Research, Inc., Irvine, CA, USA), the FastDNA SPIN Kit (MP Biomedicals, Santa Ana, CA, USA), and the MasterPure™ Yeast DNA Purification Kit (Epicentre® Biotechnologies, Madison, WI, USA).  For wart samples from a single known potato cultivar, 0.2 g (wet weight) of thawed wart was placed into a Shredder PULSE Tube with 750 µL of the kit lysis buffer and ground to a slurry using a PCT Shredder (Pressure BioSciences, Inc., South Easton, MA, USA).  The mixture was then transferred to the kit lysis tube and subjected to a 5-minute beadbeating step using a 2-mL tube vortex adapter (MO BIO Laboratories, Inc., Carlsbad, CA, USA).  DNA isolation was completed using the kit manufacturer’s protocol.

For RNA isolation, warts were prepared as above and used with the ZR Fungal/Bacterial RNA Midiprep Kit (Zymo Research, Inc.).  For each extraction, 1 mL of suspended cells was pelleted by centrifugation.  The supernatant was replaced with the kit lysis buffer, and a 5-minute beadbeating step using a 2-mL vortex adapter was performed before following the manufacturer’s protocol without further modifications.  RNA was eluted into molecular biology-grade water and preserved (dried) using an RNAstable® 96-well plate (Biomatrica, San Diego, CA, USA), according to manufacturer’s instructions.  RNA was rehydrated with 11 µL of molecular biology-grade water per well prior to use.  RNA quantity and quality were evaluated spectrophotometrically with a NanoDrop ND-1000.  RNA samples were sized using gel electrophoresis with pre-cast gels.  RNA samples were denatured prior to gel loading by heating to 65°C for 5 minutes then snap-cooling on ice.  A 1.2% FlashGel RNA cassette (Lonza Group Ltd., Basel, Switzerland) with a FlashGel™ RNA marker (0.5-9 kb) was initially used to evaluate the RNA, but sample RNA sizes were shorter than the marker and could not be resolved.  Samples (300 ng) were re-run on a pre-cast E-Gel EX 4% agarose gel (Life Technologies, Inc., Carlsbad, CA, USA) with shorter markers (Low Range ssRNA Ladder, 50-1000 bases, and microRNA Marker, 17-25 bases, both from New England BioLabs Inc., Ipswich, MA, USA).

### RNA isolation from other sample types

RNA from uninfected potato tubers was extracted a using a protocol modified from Bowrin et al. (2013).  Whole tubers were washed, surface-sanitized by spraying with DRNase Free (Argos Technologies, Inc., Elgin, IL, USA), then rinsed with DEPC-treated treated water and dried.  All implements used were also treated with DRNase Free to eliminate RNase contamination.  A sterile scalpel was used to slice the potato (with skin).  These slices were weighed and immersed in 100% formamide (1:1 w/v), and ground with a mortar and pestle.  The slurry was then transferred to 2 mL polypropylene microcentrifuge tubes (Eppendorf), and centrifuged for 10 minutes at 14,000g.  The pellet was discarded and the supernatant was passed through a sterile 0.45 µm filter assembly.  Sodium acetate was added to a concentration of 0.3M, then 0.6 volume of isopropanol was mixed in.  The mixture was incubated at -20°C for 30 minutes, then centrifuged at 14,000g for 10 minutes to pellet the RNA.  The supernatant was removed, and the RNA was purified using the RNA Clean and Concentrator-25 Kit (Zymo Research Inc.), and eluted into The RNA Storage Solution.  A NanoDrop ND-1000 was used to quantitate the RNA, and nucleic acid quality was checked with agarose gel electrophoresis.

Total RNA was extracted from Saccharomyces cerevisiae strain Cry Havoc using a rapid guanidinium isothiocyanate (GuSCN)-based lysis method (Scholin, Marin et al. 1999; Jones, Preston et al. 2008).  Briefly, cells were pelleted by centrifugation, and the supernatant discarded.  GuSCN lysis buffer (0.5-1 mL) was mixed with the yeast cells, and the mixture was heated to 85°C for 5 minutes.  Cell debris was filtered out using a 0.45 µm syringe filter, then RNA in the filtrate was purified using the Zymo RNA Clean and Concentrator-25™ Kit according to manufacturer’s instructions with, the addition of one extra 400-µL wash step (in consultation with Zymo Research Corp. technical support) to improve lysis buffer removal from the samples.  RNA was eluted in The RNA Storage Solution.

RNA from the chytrids JEL341 and JEL517A were extracted using the Zymo ZR Fungal/Bacterial RNA Miniprep™ Kit according to manufacturer’s instructions.

### 18S rDNA PCR amplification and cycle sequencing

The small-subunit ribosomal RNA gene was amplified for sequencing using the primer pairs PNS1 (forward)/NS8 (reverse), predicted to amplify 1791 bp, nearly the full length of the region (Binder, Hibbett et al. 2001).  The Phusion® Hot Start Flex 2X Master Mix (New England BioLabs, Inc.) was used for DNA amplification.  For each 25 µL reaction, the final concentrations of components were as follows: 1X Phusion Hot Start buffer (1.5 mM MgCl2), 200 µM of each dNTP, 0.5 µM each of the forward and reverse primer, 0.5 U Phusion DNA polymerase, and 50 ng genomic DNA template.  The PCR conditions were: 98°C 45 s (initial denaturation), then 35 cycles of 98°C 10 s, 70°C 30 s, 72°C 1:15 minutes.  A final extension step was performed for 10 minutes at 72°C.  An MJ Research PTC-200 Thermal Cycler (MJ Research, Inc., Watertown, MA, USA) was used to provide the required cycling parameters.  A detailed protocol is given in Appendix F.

PCR products (8 independent reactions) were pooled and purified using the Wizard® SV Gel and PCR Clean-Up System (Promega Corp., Madison, WI, USA).  Amplicon sizes were confirmed using agarose gel electrophoresis (1.8% gel) before sequencing.  Nucleic acid concentrations were assayed with a NanoDrop ND-1000 (Thermo Fisher Scientific Inc., Waltham, MA, USA).  Cycle sequencing was performed by the University of Maine DNA Sequencing Facility.  Sequences were edited using Geneious Pro 5.6.3 (Biomatters Ltd., Auckland, NZ).

### Nested PCR with S. endobioticum-specific primers

Forward and reverse primers specific for S. endobioticum (Table 5.1) were manually selected using the 18S rRNA gene alignments from publicly available and sequenced potato wart samples.  Nested PCR with these primers was performed on the amplified small-subunit RNA gene from wart tissue samples and from the closely-related chytrids JEL341 and JEL517A.  Amplification reactions were carried out as in section 5.3.4, with the following modifications: template PCR product concentration was 10 ng, and the PCR conditions were 98°C 45 s, then 20 cycles of 98°C 10 s, 72°C 15 s (annealing and extension for 2-step PCR).  A final extension step was performed for 10 minutes at 72°C.  The PCR products were visualized using agarose gel electrophoresis.

### In vitro transcription of 18S PCR product and RNA fragmentation

Limited numbers of potato wart RNA samples were available for testing, so target RNA for the optimization of hybridization parameters was generated using in vitro transcription of the 18S rRNA gene.  PCR amplification of the gene was carried out using the T7 promoter-appended PNS1 (forward)/NS8 (reverse) primer pair (Table 5.1).  For each 25 µL reaction, the final concentrations of each component, as well as the PCR conditions, were the same as in section 5.3.4.  Reaction products were purified using the Wizard® SV Gel and PCR Clean-Up system, and amplicon size was confirmed using agarose gel electrophoresis.  DNA concentration was measured using a NanoDrop ND-1000.

In vitro transcription of the rDNA PCR products was performed with the T7 High Yield RNA Synthesis Kit (New England BioLabs, Inc.) using the manufacturer protocol and 1 µg of DNA template.  The step-by-step protocol is given in Appendix G.  The transcribed RNA was purified from all other reaction components using ammonium acetate.  Briefly, the reaction solution was mixed with 1 volume 5M ammonium acetate (Ambion) to a final concentration of 2.5M ammonium acetate.  The mixture was incubated on ice for 15 minutes, then centrifuged at 10,000g for 15 minutes.  The RNA pellet was washed with 70% ethanol to remove salts, air-dried, and resuspended in The RNA Storage Solution.  RNA quality was assessed using agarose gel electrophoresis and with a NanoDrop ND-1000.

In vitro transcribed (IVT) RNA was fragmented using RNA Fragmentation Reagents (Ambion) prior to hybridization.  RNA products were diluted to 5 µg in 9 µL with MB-grade water in 0.2 mL DNase- and RNase-free thin-walled polypropylene PCR tubes (Molecular BioProducts, San Diego, CA, USA).  The 10X RNA Fragmentation Solution (1 µL) was added, and the mixtures were incubated at 95°C for 1 minute, after which 1 µL of the Stop Solution was added.

### Quantitative PCR (qPCR) and reverse transcriptase PCR (RT-qPCR)

The presence and abundance of S. endobioticum in the collected potato wart samples was confirmed using qPCR (for DNA samples) and RT-qPCR (for RNA samples).  The specific S. endobioticum forward and reverse primers used for nested endpoint PCR (Table 5.1) were used to amplify a 365-bp amplicon for qPCR.  Genomic DNA templates were diluted with molecular biology-grade water.  qPCR reactions were performed in separate wells of a Multiplate™ clear, low-profile, unskirted 96-well plate covered with with Microseal® ‘B’ adhesive seals (both from Bio-Rad Laboratories, Inc.).  Each 10 µL reaction contained 1 ng of cDNA, 500 nM each forward and reverse primer, and 5 µL SsoFast™ EvaGreen® Supermix (Bio-Rad Laboratories, Inc.).  Reactions were run on a Bio-Rad CFX96™ Real-Time PCR Detection System programmed to 98°C for 2 minutes (initial denaturation), 40 cycles of 98°C for 5 s, 60°C for 15 s, and melting curve analysis from 65-95°C at 0.5°C increments.  Each plate contained triplicates of each sample, as well as no-template controls, negative controls (potato, JEL341 and JEL517A), and copy number standards.

For RT-qPCR, RNAstable®-preserved RNA samples were first rehydrated following manufacturer’s instructions with 11 µL of molecular biology-grade water per well.  RNA concentrations were determined using a NanoDrop ND-1000, and 500 ng of each RNA sample was reverse-transcribed into complementary DNA (cDNA) using the iScript™ Reverse Transcription Supermix for RT-qPCR (Bio-Rad Laboratories, Inc., Hercules, CA, USA).  Reverse transcription products were diluted with molecular biology-grade water, and qPCR was performed as described for genomic DNA samples, except that RNA standards were used to determine copy number.  Data were analyzed using Bio-Rad CFX software.

### PNA-RNA hybridization conditions

Nucleic acid hybridization was performed in thin-walled 0.2 mL PCR tubes in a total volume of 20 µL.  A 1 µL volume of 20 µM PNA stock (1 µM final concentration) was mixed with 2 µL of 10X micrococcal nuclease reaction buffer (1X MRB is 50 mM Tris-HCl, 5 mM CaCl2, pH 7.9).  Fragmented total RNA transcripts diluted in water were added to bring the total volume to 18 µL.  Nucleic acid secondary structure was disrupted by heating to 90°C for 3 minutes, then hybridization was allowed to proceed for 5 minutes in the thermocycler.  Micrococcal nuclease (MNase, 1 µL) was added and the solution was incubated at the same temperature for a further 5 minutes.  DiSC2(5) in methanol (1 µL, 10 µM final concentration) was then mixed in, and absorbance spectra were recorded immediately after dye addition.  Spectrophotometer readings were made with a NanoDrop ND-1000 using 2 µL of sample.

For preliminary optimization experiments, the same procedure was used, except that either synthetic RNA complementary to the PNA probe or in vitro transcribed potato wart 18S rRNA was used as the sample.

## Results and discussion

### Optimization of nucleic acid extraction from potato wart

Preliminary experiments with both dried and frozen wart samples showed that nucleic acid quality varied with sample type and preparation.  DNA quality was highest when frozen and thawed wart samples were first homogenized and then lysed by beadbeating.  Examination of cell lysates after this step revealed intact S. endobioticum resting spores; increasing the beadbeating time over 5 minutes increased the number of lysed cells but resulted in lower quality DNA yields (Table B.1).

RNA extraction was maximized from rehydrated and sieved wart samples.  Thawed wart samples, whether ground or wet-sieved, caused gelling of the lysis buffer and lower RNA recovery.  The high polysaccharide content in potatoes has been reported as the cause for low RNA extraction efficiency (Kumar, Iyer et al. 2007). Various beadbeating times were tested, and five minutes of processing yielded near-100% cell lysis (data not shown).  RNA samples (rehydrated from RNAstable® storage) analyzed spectrophotometrically (Table B.2) and with agarose gel electrophoresis (Figure 5.2) revealed significant degradation.  Most of the low-molecular-weight RNA smear was distributed between <17 to 150 bases from gel analysis using ImageJ software.  This degradation is most likely the result of the extraction procedure.

| a) | b) |
| --- | --- |
| 150 nt M1 M2 A B C D E F | M1 M2 G H I J K L |

Figure 5.2. RNA extracted from potato wart samples show significant degradation.  The gel lanes are labeled alphabetically and correspond to the samples given in Table 5.3 and Table B.2.  M1 is the NEB Low Range ssRNA Ladder (50-1000 bases), while M2 is the NEB microRNA Marker (17-25 bases).  Samples were heat-denatured (65°C for 5 min) and run on a 4% E-Gel EX pre-cast agarose gel.

### PNA probe development from S. endobioticum 18S rDNA sequences

As of this writing, only two 18S rRNA gene sequences for S. endobioticum were available in GenBank.  Thus, to minimize PCR-generated artifacts in the amplification products to be used for sequencing, the proofreading Phusion® polymerase was used in all PCR reactions.  This polymerase a derivative of the enzyme from Pyrococcus fusorius (Pfu) and has a reported error rate of 9.5 X 10-7 errors/bp (Frey and Suppmann 1995), much lower than that of either Pfu (1.6 X 10-6 errors/bp (Lundberg, Shoemaker et al. 1991)) or Taq (2.0 X 10-5 errors/bp (Lundberg, Shoemaker et al. 1991)).

The 18S rDNA sequences obtained were aligned with the two existing S. endobioticum GenBank entries.  Short 16-base candidate probe sequences were selected for minimal secondary structure formation and low self-hybridization characteristics.  These structural constraints are requirements for PNA probe development since PNA-PNA hybridization is favored over PNA-RNA duplex formation (Egholm, Buchardt et al. 1993; Jensen, Ørum et al. 1997), and nonspecific dye aggregation to PNA probe sequences has been reported (Tedeschi, Sforza et al. 2007).  
The 18S rRNA database ProbeBase (Loy, Horn et al. 2003) was used to verify the in silico specificity of the candidate probes.  The PNA sequence selected showed complementarity only to S. endobioticum rRNA entries.

### S. endobioticum PCR primer development and testing

Specific S. endobioticum forward and reverse primers were designed to 1) to include the PNA target region and 2) amplify only S. endobioticum sequences.  The S. endobioticum 18S rDNA sequences from GenBank (Accession Nos. AY854021.1 and AJ784274.1) were aligned with 18S rDNA sequences from potato and other Synchytrium species.  The specificity of the candidate primers were tested in silico using Primer-BLAST (default settings) and did not show any cross-reactivity to other sequences.

The 18S rDNA amplified from potato wart DNA samples was sequenced but a BLAST search of the resulting sequences yielded only potato DNA.  Nested PCR was then attempted using the S. endobioticum primers with the 18S rDNA PCR products as templates.  Gel electrophoresis yielded the expected band, and the DNA was confirmed as S. endobioticum after cycle sequencing and alignment with the existing 18S rDNA sequence.  This result indicates that, while S. endobioticum DNA is indeed present in the potato wart sample, the majority of the DNA is from the potato host.

The specificity of the primer set was tested with the PNS1/NS8-amplified PCR products from the related chytrids Chytridiomyces sp. JEL341 and JEL517A.  The primer set did not produce the expected 365-bp product with these templates, as shown in Figure 5.3.  This demonstrates that the primer set is specific for S. endobioticum and can be used to identify the fungus.

300 bp

5

6

7

2

1

3

4

Figure 5.3. Specificity of primer set to S. endobioticum and related chytrid species.  The primer set Se18S816f/Se18S1136r (Table 5.1) was used to amplify a putative 365-base S. endobioticum-specific sequence in the 18S rRNA gene.  A 1.2% agarose gel was used to size the PCR products.  Lanes: 1) FlashGel DNA ladder; 2) No DNA negative control; 3) Potato wart DNA 2; 4) Potato wart DNA 8; 5) Potato wart DNA 9; 5) JEL341 DNA; 7) JEL517A DNA.  The arrow points to the 300 bp fragment of the ladder.

### Optimization of qPCR conditions

Primers for real-time PCR are ideally chosen to amplify a 50-150 bp region to take advantage of the efficiency of the reaction.  However, no other suitable primer sets that include the desired probe region were found using in silico analysis.  Thus, the S. endobioticum endpoint PCR primers were adapted for qPCR use by increasing the recommended annealing/extension time from 5 seconds to 15 seconds (Dr. David A. Palmer, Bio-Rad technical support, personal communication).  The optimal annealing/extension temperature (60°C) was found using potato wart, JEL341 and JEL517A DNA templates in a gradient PCR setup.  At this temperature, no amplification of either JEL341 or JEL517A was observed, and the presence of potato wart was detected earliest (as measured by the crossing threshold value, Cq).

A standard curve PCR reaction was used to evaluate the efficiency and linearity of the reaction, with potato wart 18S rDNA as templates (102-106 copies).  The (software-calculated) efficiency E = 98.9%, R2 = 0.997, and slope = -3.347, which are all acceptable values for qPCR.  These reaction conditions were used for all succeeding qPCR experiments.

One interesting finding from the qPCR data is that S. endobioticum target sequences (using the wart-specific nested PCR primers) are detected from unamplified genomic DNA samples using qPCR, but are undetectable when using conventional endpoint amplification without nested PCR.  The sensitivity difference is even more apparent when the amount of input DNA is compared: 50 ng of genomic template was used for endpoint PCR, while 1 ng was used for qPCR.  While other authors have reported enhanced sensitivity when using qPCR compared to endpoint PCR, this may be attributed to many factors, such as detection chemistry, primer sequences used, polymerase choice, DNA template variability, and nucleic acid extraction method (Bastien, Procop et al. 2008).  Regardless, qPCR appears to be a superior choice to conventional PCR for this application, as no pre-amplification of the 18S rRNA gene is necessary for the detection of S. endobioticum from potato wart DNA samples.  The use of qPCR over nested PCR also reduces both reagent consumption and time to detection.

### Effect of RNA fragmentation on apparent target copy number

The RNA extracted from wet-sieved potato wart showed significant fragementation, with a majority of the RNA smear in the range of <17 to ~150 bases.  Since target copy number is calculated using a primer set that amplifies a 365-bp product, the copy numbers from the RNA samples will most likely be underestimated (Fleige and Pfaffl 2006; Gingrich, Rubio et al. 2008).  To assess the effect of RNA fragmentation on apparent target abundance, intact and fragmented in vitro transcribed wart 18S rRNA were used as templates for the RT-qPCR reaction.  Figure 5.4 shows the effect of fragmentation on the reported target abundance.  Amplification of the S. endobioticum PCR product is underreported by 70-95% as a result of RNA degradation.

a)

b)

Figure 5.4. Effect of RNA fragmentation on apparent target copy number.  a) Absolute target copy numbers from intact and fragmented in vitro transcribed potato wart RNA. b) Apparent copy number from fragmented samples, as a percentage of the original copy number from intact RNA.

### Target copy number estimations for potato wart genomic DNA and RNA samples using qPCR

Target copy numbers in extracted potato wart nucleic acids were estimated using qPCR, and are shown in Table 5.2 (DNA) and Table 5.3 (RNA).  S. endobioticum was detected in all potato wart genomic DNA samples, regardless of storage condition (frozen or dried), indicating that both treatments preserve potato DNA and chytrid resting spores.  The exception to this is DNA 7, which was extracted from 0.2 g of frozen, thawed Arran Victory (potato cultivar) wart.  The thawed sample appeared to be more decomposed and spongier compared to other frozen warts, and DNA yield and quality were poor (1.76 ng/µL).  No further information about storage conditions was available, but it is possible that the DNA in this sample was already degraded prior to freezing.  No direct comparisons regarding DNA yield and target quantity could be made with these samples, as sporangial counts were unavailable prior to nucleic acid extraction.

Table 5.2. Target copy numbers from potato wart genomic DNA samples.  Copy numbers were estimated using absolute quantification with a standard curve during qPCR.

| DNA sample | Target copy number/ng (X 103) |
| --- | --- |
| 1 | 4.51 |
| 2 | 137 |
| 3 | 18.8 |
| 4 | 407 |
| 5 | 529 |
| 6 | 37.8 |
| 7 | 0 |
| 8 | 96.8 |
| 9 | 264 |

S. endobioticum was detected in all RNA samples using qPCR.  However, the actual target abundances in RNA samples are likely higher than the numbers established using qPCR, due to the degradation of the samples and the relatively long amplicon used to determine copy number ((Fleige and Pfaffl 2006; Gingrich, Rubio et al. 2008) and Figure 5.4).

Table 5.3. Target copy numbers from potato wart RNA samples.  Copy numbers were estimated using absolute quantification with a standard curve using RT-qPCR.

| RNA sample | Target copy number/ng (X 103) |
| --- | --- |
| A | 0.868 |
| B | 1.444 |
| C | 16.97 |
| D | 4.177 |
| E | 15.25 |
| F | 69.03 |
| G | 11.83 |
| H | 22.53 |
| I | 36.42 |
| J | 21.2 |
| K | 3.011 |
| L | 10.01 |

### Optimization of PNA-RNA hybridization conditions

All hybridization experiments were performed with fragmented RNA samples, as PNA hybridization to these targets is improved when the long, structured transcript is sheared into smaller sizes ((Small, Call et al. 2001; Nelson, Liles et al. 2002; Mehlmann, Townsend et al. 2005; Liu, Guo et al. 2007) and unpublished data).  Hybridization temperature was selected to 1) maximize target binding over non-target capture; and 2) preserve MNase activity.  The predicted Tm for the PNA probe hybridized to complementary DNA given by the equation (Giesen, Kleider et al. 1998) is 68.3°C.  Hybridization to the same RNA sequence increases the Tm by ~4°C (Egholm, Buchardt et al. 1993; Jensen, Ørum et al. 1997), so the estimated Tm for the PNA-RNA hybrid is 72.3°C.  Various hybridization temperatures were tested from 10-50°C below the Tm.  In vitro transcribed potato wart 18S rRNA was used for testing to account for higher-order RNA structures present in natural samples.  Hybridization was most efficient from 10-30°C below the melting temperature, but nuclease activity decreased at temperatures above 60°C (Figure 5.5), so the highest temperature (63°C) was not used for further testing.

Figure 5.5. Optimization of hybridization temperature for S. endobioticum binding to PNA 033.  In vitro transcribed 18S rRNA was used to determine the best hybridization temperature to PNA 033.  After hybridization, 10 Ku of micrococcal nuclease was used to digest excess RNA at the same temperature prior to DiSC2(5) dye addition.

Hybridization signals at 43°C and 53°C appeared to be comparable, but tests using only synthetic RNA oligonucleotides revealed that hybridization at the higher temperature resulted in digestion of the perfectly-matched RNA off the PNA probe (Figure 5.6).  Thus, 43°C was chosen as the hybridization/digestion temperature for all subsequent studies to prevent the inadvertent removal of target sequences.

Figure 5.6. Further optimization of hybridization and digestion temperature of S. endobioticum RNA to PNA 033.  Perfectly-complementary synthetic RNA sequences were used to evaluate the effect of hybridization/digestion temperature.  Micrococcal nuclease (10 Ku) was incubated with the mixture after hybridization.

Nuclease concentration was also selected to maximize the signal to noise ratio.  The lower limit of detection using synthetic RNA oligonucleotides was in the 100 nM (picomole) range, but preliminary experiments with in vitro transcribed RNA revealed lower numbers of detectable targets (in the high picomolar range).  Based on the estimated number of target copies in the RNA samples (Table 5.3), total RNA concentrations of 20-30 ng/µL, equivalent to ~107 targets (~166 X 10-18 mole) per 20 µL reaction were selected for optimizing the nuclease concentration.  For both samples, a nuclease concentration of 10 Ku produced the best signal-to-noise ratio (5.93 for 20 ng/µL, and 4.57 for 30 ng/µL), as shown in Figure 5.7.

| a) | b) |
| --- | --- |

Figure 5.7. Effect of micrococcal nuclease concentration on signal-to-noise ratio.  Micrococcal nuclease (10-50 Ku) was added after hybridization and incubated for 5 minutes at 43°C.  Total RNA concentrations of a) 20 ng/µL and b) 30 ng/µL were tested to approximate the amounts of S. endobioticum genomic RNA needed for detection.  The gray patterned bars represent yeast total RNA only, while the colored bars are yeast RNA with 250 nM synthetic RNA oligonucleotides spiked in.  The dashed line represents the signal obtained from a no-RNA control.  MNase concentrations of 10 Ku yielded the best signal-to-noise ratios for both amounts of RNA.

### Hybridization to in vitro transcribed RNA and assay detection limit

The lower limit of detection for S. endobioticum was assessed using in vitro transcribed RNA and the optimized hybridization conditions, as shown in Figure 5.8.  Varying amounts (5, 10, 15, 20 ng/µL) IVT RNA from genomic DNA samples 2, 8, and 9 were spiked into a yeast RNA background, for a total RNA concentration of 20 ng/µL.  The concentration curves are plotted as the number of target copies for each concentration tested (estimated from RT-qPCR, Table 5.3).  The dashed line is the highest signal obtained from a pure yeast sample (20 ng/µL), and is used as the threshold for S. endobioticum detection.

The minimum detectable number of targets was 17.3 X 106 for DNA 8 IVT RNA; this is equivalent to 2.87 X 10-17 mole (1.44 pM in a 20 µL reaction volume).  DNA 2 IVT RNA was detectable at 127.7 X 106 copies (21.2 X 10-17 mole/10.6 pM), while DNA  9 IVT RNA was indistinguishable from the background at 44.8 X 106 copies (7.44 X 10-17 mole/3.72 pM).  From these differences, it is evident that copy number alone is not sufficient to predict the hybridization signal obtained from the colorimetric method.  It has been shown that target accessibility to the probe sequence is the major determinant of hybridization signal intensity (Small, Call et al. 2001; Behrens, Rühland et al. 2003; Chandler, Newton et al. 2003; Peplies, Glöckner et al. 2003; Metfies and Medlin 2008).  Slight sequence differences in the various S. endobioticum pathotypes, which can influence higher-order nucleic acid structures, could be one source of the variability observed.  Another source of signal discrepancy could be the size distribution of the fragmented RNA; the procedure used in these studies produced RNA from 30-300 bases, but RNA sizes outside of this range have been observed.  Previous studies have shown that hybridization signals were highest with fragments between 20-150 bases (Mehlmann, Townsend et al. 2005; Liu, Guo et al. 2007).

a) DNA 2

b) DNA 8

c) DNA 9

Figure 5.8. Minimum detectable number of S. endobioticum copies assessed using in vitro transcribed 18S rRNA.  Target copy numbers were estimated using qPCR results of intact samples of a) DNA 2, b) DNA 8 and c) DNA 9 in vitro transcribed RNA.  For all samples, the total RNA concentration was 20 ng/µL, with the IVT RNA comprising 5, 10, 15, and 20 ng/µL spiked into a S. cerevisiae yeast RNA background.  The dashed line is the background level generated using 20 ng/µL yeast RNA only.

### PNA probe specificity to other organisms

The PNA probe sequence was targeted only to S. endobioticum, and was not predicted to bind to either JEL341 or JEL517A based on available sequence alignments and public database information.  However, it is well known that in silico-validated probe sequences must be tested empirically to determine actual hybridization characteristics (Perry-O’Keefe, Rigby et al. 2001; Hugenholtz, Tyson et al. 2002; Nelson, Liles et al. 2002; Abdullahi, Koerbler et al. 2005; Brehm-Stecher, Hyldig-Nielsen et al. 2005; He, Wu et al. 2005; Peplies, Lachmund et al. 2006; Pozhitkov, Noble et al. 2006).

The specificity of the PNA probe sequence was tested by hybridization to synthetic (oligonucleotide and in vitro transcribed) and extracted RNA.  The probe did not exhibit hybridization to RNA from firefly luciferase gene, Alexandrium spp., S. cerevisiae (yeast), S. tuberosum (potato).  It did cross-react to total RNA from organisms closely-related to S. endobioticum, JEL341 and JEL517A.  JEL341 shares 89.3% similarity to S. endobioticum, while JEL517A shares 83.6% identity, based on alignments of their 18S rRNA gene sequences.

The hybridization of the PNA probe to these related chytrids, while not ideal, is not anticipated to interfere with the detection of S. endobioticum from potato wart samples.  JEL341 and JEL517A are found with chickweed, and this plant grows in hot, humid environments that are hostile to S. endobioticum; consequently, these related chytrids are not expected to be present in any samples containing the desired target.

### Hybridization to potato wart RNA

Potato wart RNA samples were hybridized to PNA 033 and hybridization signals were calculated after correction for RNAstable® content (Appendix B).  Samples were classified as S. endobioticum-positive or negative based on the highest possible signal from potato RNA alone (Figure 5.9 and Table 5.4).

Figure 5.9. Detection of S. endobioticum from extracted RNA.  The dashed line denotes the highest possible signal from potato (host) RNA.  All RNA concentrations were 30 ng/µL (600 ng total per reaction).

Table 5.4. Correlations between RNA quality, target copy number, and colorimetric detection.

| RNA sample | Target copy number per reaction (X 106) | 260/230 | 260/280 | Sample origin | Detection |
| --- | --- | --- | --- | --- | --- |
| A | 0.521 | 0.10 | 1.55 | Dried rehydrated wart | + |
| B | 0.856 | 0.85 | 1.59 | Dried wart | + |
| C | 10.18 | 1.29 | 1.89 | Dried wart | + |
| D | 2.51 | 1.35 | 1.83 | Dried wart | + |
| E | 9.15 | 1.46 | 1.95 | Dried wart | - |
| F | 41.41 | 1.31 | 1.79 | Frozen wart | + |
| G | 7.1 | 1.42 | 1.89 | Dried wart | + |
| H | 13.52 | 1.18 | 1.78 | Dried wart | + |
| I | 21.85 | 1.43 | 1.88 | Dried wart | + |
| J | 12.72 | 1.22 | 1.81 | Dried wart | + |
| K | 1.8 | 0.28 | 1.55 | Dried rehydrated wart | - |
| L | 10.01 | 1.89 | 1.88 | Dried wart | + |

Of the 12 available potato wart samples, 9 (83.3%) were correctly identified as containing S. endobioticum (2 false negatives).  No clear correlations between (apparent) target copy number and hybridization signal were found, consistent with the results from the in vitro transcribed RNA.  The irregular fragment size distribution (Figure 5.2) may have contributed to the false negatives obtained, as other studies have determined that a narrow range of sizes (20-150 bases) produced the highest hybridization signals (Mehlmann, Townsend et al. 2005; Liu, Guo et al. 2007).  Spectrophotometrically measured RNA quality was also not a good indicator of hybridization success, but the low 260/230 ratios obtained from the extraction procedure denote residual impurities that could have interfered with the assay.  Attempts to repurify the RNAstable®-containing samples using either ammonium acetate precipitation or spin columns were unsuccessful.

The total amount of RNA in each sample was 600 ng (30 ng/µL concentration in a 20 µL volume), and this is in concurrence with other reports of minimum total RNA required for species detection from environmental samples (Small, Call et al. 2001; Abdullahi, Gryshan et al. 2011).  Compared to the one study using potato wart RNA in microarray hybridization (Abdullahi, Koerbler et al. 2005), the amount used in this assay is 76% less.  While amplification-based schemes require much lower starting target quantities, the short assay time and minimal expertise, reagent and equipment requirements of this method are clear advantages in rapid detection that can be performed by untrained workers.  Currently, suspected cases of potato wart fungus must be sent out to specialized laboratories for confirmation; the use of the simple colorimetric assay presented here can significantly reduce the time to disease diagnosis and prevent the spread of S. endobioticum to other locations.

## Chapter summary

This work demonstrates the rapid and specific identification of Synchytrium endobioticum from total RNA extracted from resting spores in potato wart samples.  Cellular DNA was extracted from infected potatoes, and a portion of the 18S rRNA gene was sequenced to confirm the identity of the fungus, as well as for probe and primer development.  A primer set spanning a 365-bp region of the gene was validated using both endpoint and real-time PCR, and did not show cross-reactivity to either the potato host or to related chytrid sequences.

A peptide nucleic acid (PNA) probe was designed to capture 18S rRNA sequences, and nontarget nucleic acids (mostly from the potato host) were digested using micrococcal nuclease.  Hybridization was indicated by the cyanine dye 3,3'-diethylthiadicarbocyanine iodide (DiSC2(5)), and detection of targets down to ~10-17 mole from in vitro transcribed RNA was possible using this assay.  Testing with known wart-positive extracted RNA samples correlated with qPCR detection, and few false negatives.  The incorrectly identified samples may be attributed to suboptimal target accessibility, and could be improved by either increasing RNA amount or fragmentation prior to hybridization.  This rapid technique uses unamplified nucleic acids and requires minimal instrumentation and training.  It has the potential to quickly identify suspected outbreaks of S. endobioticum and can be used to prevent the spread of the disease to uninfected areas.

# SUMMARY AND FUTURE CONSIDERATIONS

## Contributions of this dissertation

This work details the application of PNA-based nucleic acid hybridization schemes to the detection of RNA from environmental samples, for the development of user-friendly, inexpensive, and field-deployable methods.  Chapter 2 presents the first report of the conjugation of functional PNA sequences to citrate-stabilized gold nanoparticles.  Hybridization to complementary DNA was achieved, and attempts at forming larger DNA-linked particle assemblies were moderately successful.

The cyanine dye DiSC2(5), previously only used to indicate PNA-DNA hybrids, was successfully utilized for the visualization of PNA-RNA hybrids.  In Chapter 3 sequence information from the toxic dinoflagellate Alexandrium fundyense/tamarense/catanella species complex was used to optimize probe, dye, and methanol concentrations for eventual application to cellular RNA targets; all subsequent experiments were performed with 1 µM PNA probe, 10 µM DiSC2(5), and 10% methanol.  Assay specificity was improved to single-base-mismatch resolution of bases flanked by adenines and uracils through the use of microccocal nuclease, a relatively non-base-specific enzyme that digests both single- and double-stranded nucleic acids.  Tests with RNA oligonucleotides revealed that the assay is linear for target concentrations from 100 nM to 1 µM and could distinguish A. fundyense/tamarense/catanella sequences even when mixed with increasing amounts of sequences from the related species A. ostenfeldii.  The assay was tested with total RNA from Alexandrium cell cultures and could detect ~3300 cells from background.  At ~6700 cells, A. fundyense cells could be distinguished from those of A. ostenfeldii.

A handheld colorimeter was designed to quantitate the hybridization signals from the colorimetric assay in non-laboratory settings, and design considerations and assembly details are given in Chapter 4.  The device was built primarily from commercially available electronics parts and is significantly less expensive than laboratory spectrophotometers.  Side-by-side testing against a benchtop spectrophotometer showed equivalent responses over the RNA concentration range tested (250 nM to 1 µM).  Raw hybridization signals from the colorimeter showed lower sensitivity compared to the laboratory instrument; this was attributed to the finite spectral widths of the LED light sources used, and to the wavelength-dependent photodiode responsivity.

The suitability of the colorimetric method for the rapid detection of other organisms is discussed in Chapter 5.  Here, the fungal potato pathogen Synchytrium endobioticum was used to test the specificity and sensitivity of the assay.  Specific PCR primers and a PNA probe were designed from the 18S rRNA gene sequence.  The primer set amplified a 365-bp region and was validated for use with both endpoint and real-time PCR.  The primers were specific to S. endobioticum and did not cross-react with either potato or related chytrid sequences.  The PNA probe was used to detect in vitro transcribed RNA down to ~10-17 mole (1.44 pM), but this detection limit appears to be strain-dependent.  Total RNA from potato wart samples could be positively identified at 600 ng, with few false negatives, within 15 minutes.

## Considerations for future work

### RNA-templated formation of larger gold nanoparticle assemblies

The results from Chapter 2 indicate that PNA-decorated nanoparticles can be aggregated using the specificity of nucleic acid hybridization.  However, numerous attempts to reproduce this data were unsuccessful, even with the same batch of reagents.  The controlled aggregation of PNA-decorated gold nanoparticles is not fully understood and produced inconsistent results.  Thus, this method was not further pursued.  However, a re-examination of the data has revealed that PNA sequence design may be critical to the success of the assay.  If two PNA sequences are to be used as probes, both sequences must be checked for self-complementarity, and complementarity to each other (probe sequences binding to one another).  PNA-PNA hybridization is favored over either PNA-RNA or PNA-DNA formation so care must be taken in sequence design for a two-PNA probe system.

Another possibility is the use of other nucleic acid mimics as probes.  In particular, morpholino oligos would be well-suited for this application.  Similar to PNA probes, morpholinos offer high specificity and increased resistance to nucleases.  They are also more soluble in water, so they should be less likely to self-aggregate in solution.

### Specific detection of low numbers of toxigenic Alexandrium cells from whole seawater samples

Alexandrium cells can be detected above the background at ~3300 cells.  This requires the filtration of 33 L of seawater to achieve the necessary detection limit of 100 cells/L.  While other methods have demonstrated detection limits that are even higher than this, for practical reasons it would be advantageous to lower the detectable number of cells.

The use of fluorescently-tagged probes, specifically a PNA forced intercalation (FIT) probe, is one potential way of improving the detection limit of the assay, due to the high (>26-fold change) signal amplification associated with probe hybridization to the target.  These probes have been shown to resolve single mismatches at 25°C in large excess of nontarget sequences, so the assay can be isothermal and nuclease digestion will be unnecessary.  Another advantage of using these labeled probes is the decrease in the amount of PNA needed per assay, as the detection limit is not set by the dye concentration relative to the number of duplexes in solution.  As well, dye concentration will not need to be optimized for different probe concentrations.

Another area of improvement is in the development of a field-compatible lysis protocol.  The current RNA extraction procedure requires the use of a chemical buffer that is incompatible with DiSC2(5) colorimetric detection and will likely quench fluorescent tags.  Alternatively, magnetic separation of the extracted RNA and resuspension in hybridization buffer could be investigated.

### Quantitative detection of Synchytrium endobioticum from soil samples

We have demonstrated that the colorimetric assay can be used to detect S. endobioticum from potato wart samples.  A natural extension to this study is to validate the usefulness of the assay to nucleic acids extracted from soil samples.  Current regulations for re-opening a quarantined field for agricultural cultivation require substantial soil testing to demonstrate the absence of viable fungal spores.  If possible, more probe sequences should be designed using available sequence data, then validated experimentally for specificity and sensitivity improvements.

# BIBLIOGRAPHY

(1999). “Synchytrium endobioticum: soil tests and descheduling of previously infested plots.” EPPO Bulletin 29(3): 225-231.

(2004). “Synchytrium endobioticum.” EPPO Bulletin 34(2): 213-218.

(2012). Potatoes 2011 Summary. N. A. S. S. United States Department of Agriculture.

Abdullahi, I., Y. Gryshan, et al. (2011). “Amplification-free detection of grapevine viruses using an oligonucleotide microarray.” Journal of Virological Methods 178(1-2): 1-15.

Abdullahi, I., M. Koerbler, et al. (2005). “The 18S rDNA sequence of Synchytrium endobioticum and its utility in microarrays for the simultaneous detection of fungal and viral pathogens of potato.” Applied Microbiology and Biotechnology 68(3): 368-375.

Abramoff, M. D., P. J. Magelhaes, et al. (2004). “Image processing with ImageJ.” Biophotonics Int 11(7): 36-42.

Achim, C., B. A. Armitage, et al. (2008). “Peptide Nucleic Acids (PNAs).”

Agriculture, U. S. D. o. (2007). Recovery Plan for Potato Wart Disease Caused by Synchytrium endobioticum (Schilberszky) Percival. U. S. D. o. Agriculture.

Alexander, M., L. A. Heppel, et al. (1961). “The Purification and Properties of Micrococcal Nuclease.” The Journal of Biological Chemistry 236(11): 3014-3019.

Alivisatos, A. P., K. P. Johnsson, et al. (1996). “Organization of 'nanocrystal molecules' using DNA.” Nature 382: 609-611.

Ananthanawat, C., T. Vilaivan, et al. (2010). “Comparison of DNA, aminoethylglycyl PNA and pyrrolidinyl PNA as probes for detection of DNA hybridization using surface plasmon resonance technique.” Biosens Bioelectron 25(5): 1064-1069.

Anderson, D., D. Kulis, et al. (2005). “Identification and enumeration of Alexandrium spp. from the Gulf of Maine using molecular probes.” Deep Sea Research Part II: Topical Studies in Oceanography 52(19-21): 2467-2490.

Anderson, D. M. (1997). “Bloom dynamics of toxic Alexandrium species in the northeastern U.S.” Limnology and Oceanography 42(5, part 2): 1009-1022.

Anderson, D. M., D. M. Kulis, et al. (1999). “Detection of the toxic dinoflagellate Alexandrium fundyense (Dinophyceae) with oligonucleotide and antibody probes: Variability in labeling intensity with physiological condition.” Journal of Phycology 35(4): 870-883.

Anderson, D. M., D. M. Kulis, et al. (1990). “Dynamics and physiology of saxitoxin production by the dinoflagellates Alexandrium spp.” Marine Biology 104: 511-524.

Armitage, B. A. (2003). “The impact of nucleic acid secondary structure on PNA hybridization.” Drug Discovery Today 8: 222-228.

Aslan, K. and V. H. Perez-Luna (2002). “Surface modification of colloidal gold by chemisorption of alkanethiols in the presence of a nonionic surfactant.” Langmuir 18: 6059-6065.

Bastien, P., G. W. Procop, et al. (2008). “Quantitative Real-Time PCR Is Not More Sensitive than “Conventional” PCR.” Journal of Clinical Microbiology 46(6): 1897-1900.

Behrens, S., C. Rühland, et al. (2003). “In Situ Accessibility of Small-Subunit rRNA of Members of the Domains Bacteria, Archaea, and Eucarya to Cy3-Labeled Oligonucleotide Probes.” Applied and Environmental Microbiology 69(3): 1748-1758.

Binder, M., D. S. Hibbett, et al. (2001). “Phylogenetic Relationships of the Marine Gasteromycete Nia vibrissa.” Mycologia 93(4): 679-688.

Bowrin, V., J. Rouse-Miller, et al. (2013). “Formamide-based RNA isolation at above zero temperatures from high starch cassava tubers.” Phytochemical Analysis 24(1): 93-96.

Bratcher, A. R. and L. B. Connell (2012). The Use of Peptide Nucleic Acids in Surface Plasmon Resonance for Detection of Red Tide Algae. Molecular Biological Technologies for Ocean Sensing. S. M. Tiquia-Arashiro, Humana Press: 135-150.

Brehm-Stecher, B. F., J. J. Hyldig-Nielsen, et al. (2005). “Design and evaluation of 16S rRNA-targeted peptide nucleic acid probes for whole-cell detection of members of the genus Listeria.” Applied and Environmental Microbiology 71(9): 5451-5457.

Brewer, G., E. Murray, et al. (1992). “RNase ONE™: Advantages for nuclease protection assays.” Promega Notes 38(1).

Brown, S. C., S. A. Thomson, et al. (1994). “NMR solution structure of a PNA complexed with RNA.” Science 265: 778-780.

CCEHBR. “Phytoplankton Monitoring Network.”   Retrieved 08/03/2010, from http://www.chbr.noaa.gov/pmn/default.aspx.

Chakrabarti, R. and A. M. Klibanov (2003). “Nanocrystals modified with peptide nucleic acids (PNAs) for selective self-assembly and DNA detection.” J. Am. Chem. Soc. 125: 12531-12540.

Chambers, J. P., B. P. Arulanandam, et al. (2008). “Biosensor recognition elements.” Curr Issues Mol Biol 10: 1-10.

Chandler, D. P., G. J. Newton, et al. (2003). “Sequence versus Structure for the Direct Detection of 16S rRNA on Planar Oligonucleotide Microarrays.” Applied and Environmental Microbiology 69(5): 2950-2958.

Chisnall Hampson, M. (1992). “A bioassay for Synchytrium endobioticum using micropropagated potato plantlets.” Canadian Journal of Plant Pathology 14(4): 289-292.

Connell, L. B., J. Ray, et al. (2006). “Enhanced detection levels in a semi-automated sandwich hybridisation assay using a peptide nucleic acid (PNA) signal probe.” African Journal of Marine Science 28(2): 237-239.

D’Agata, R., G. Breveglieri, et al. (2011). “Direct Detection of Point Mutations in Nonamplified Human Genomic DNA.” Analytical Chemistry 83(22): 8711-8717.

D’Agata, R., R. Corradini, et al. (2010). “Ultrasensitive detection of non-amplified genomic DNA by nanoparticle-enhanced surface plasmon resonance imaging.” Biosensors and Bioelectronics 25(9): 2095-2100.

D’Agata, R., R. Corradini, et al. (2008). “Ultrasensitive detection of DNA by PNA and nanoparticle-enhanced surface plasmon resonance imaging.” ChemBioChem 9(13): 2067-2070.

Deck, M. K., E. S. Anderson, et al. (2012). “Multicenter evaluation of the Staphylococcus QuickFISH method for simultaneous identification of Staphylococcus aureus and coagulase-negative staphylococci directly from blood culture bottles in less than 30 minutes.” Journal of Clinical Microbiology 50(6): 1994-1998.

Deisingh, A. K. and M. Thompson (2002). “Detection of infectious and toxigenic bacteria.” Analyst 127: 567-581.

Demers, L. M., C. A. Mirkin, et al. (2000). “A fluorescence-based method for determining the surface coverage and hybridization efficiency of thiol-capped oligonucleotides bound to gold thin films and nanoparticles.” Anal. Chem. 72: 5535-5541.

Demidov, V. V., V. N. Potaman, et al. (1994). “Stability of peptide nucleic acids in human serum and cellular extracts.” Biochem. Pharmacol. 48(6): 1310-1313.

Demidov, V. V., V. N. Potaman, et al. (1994). “Stability of peptide nucleic acids in human serum and cellular extracts.” Biochemical Pharmacology 48(6): 1310-1313.

Dingwall, C., G. P. Lomonossoff, et al. (1981). “High sequence specificity of micrococcal nuclease.” Nucleic Acids Research 9(12): 2659-2673.

Drew, H. R. (1984). “Structural Specificities of Five Commonly Used DNA Nucleases.” Journal of Molecular Biology 176: 535-557.

Dyhrman, S. T., D. Erdner, et al. (2006). “Molecular quantification of toxic Alexandrium fundyense in the Gulf of Maine using real-time PCR.” Harmful Algae 5(3): 242-250.

Eck, D., C. A. Helm, et al. (2001). “Plasmon resonance measurements of the adsorption and adsorption kinetics of a biopolymer onto gold nanocolloids.” Langmuir 17: 957-960.

Egholm, M., O. Buchardt, et al. (1993). “PNA hybridizes to complementary oligonucleotides obeying the Watson-Crick hydrogen-bonding rules.” Nature 365: 566-568.

Egholm, M., O. Buchardt, et al. (1993). “PNA hybridizes to complementary oligonucleotides obeying the Watson-Crick hydrogen-bonding rules.” Nature 365: 566-568.

Egholm, M., O. Buchardt, et al. (1992). “Peptide nucleic acids (PNA). Oligonucleotide analogues with an achiral peptide backbone.” J. Am. Chem. Soc. 114: 1895-1897.

Egholm, M., O. Buchardt, et al. (1992). “Peptide Nucleic Acids (PNA). Oligonucleotide Analogues with an Achiral Peptide Backbone.” Journal of the American Chemical Society 114: 1895-1897.

Eriksson, M. and P. E. Nielsen (1996). “PNA-nucleic acid complexes. Structure, stability and dynamics.” Quarterly Reviews of Biophysics 29(4): 369-394.

Fan, X., I. M. White, et al. (2008). “Sensitive optical biosensors for unlabeled targets: A review.” Anal Chim Acta 620: 8-26.

Fleige, S. and M. W. Pfaffl (2006). “RNA integrity and the effect on the real-time qRT-PCR performance.” Mol Aspects Med 27(2-3): 126-139.

Frey, B. and B. Suppmann (1995). “Demonstration of the Expand™ PCR System's Greater Fidelity and Higher Yields with a lacI-based PCR Fidelity Assay.” Biochemica 2: 34-35.

Fuchs, B. M., G. Wallner, et al. (1998). “Flow cytometric analysis of the in situ accessibility of Escherichia coli 16S rRNA for fluorescently labeled oligonucleotide probes.” Applied and Environmental Microbiology 64(12): 4973-4982.

Giesen, U., W. Kleider, et al. (1998). “A formula for the thermal stability (Tm) prediction of PNA/DNA duplexes.” Nucleic Acids Research 26(21): 5004-5006.

Gildea, B. D., S. Casey, et al. (1998). “PNA Solubility Enhancers.” Tetrahedron Letters 39(1): 7255-7258.

Gingrich, J., T. Rubio, et al. (2008). Effect of RNA Degradation on Data Quality in Quantitative PCR and Microarray Experiments. Hercules, CA, Bio-Rad Laboratories, Inc.

Godhe, A., C. Cusack, et al. (2007). “Intercalibration of classical and molecular techniques for identification of Alexandrium fundyense (Dinophyceae) and estimation of cell densities.” Harmful Algae 6(1): 56-72.

Grabar, K. C., R. G. Freeman, et al. (1995). “Preparation and characterization of Au colloid layers.” Anal. Chem. 67: 735-743.

Guillard, R. R. L. and P. E. Hargraves (1993). “Stichochrysis immobilis is a diatom, not a chrysophyte.” Phycologia 32: 234-236.

Hampson, M. C. (1981). “Potato wart caused by Synchytrium endobioticum: past and future emphases in research.” Canadian Journal of Plant Pathology 3(2): 65-122.

Hannah, K. C. and B. A. Armitage (2004). “DNA-Templated Assembly of Helical Cyanine Dye Aggregates: A Supramolecular Chain Polymerization.” Accounts of Chemical Research 37: 845-853.

Hannah, K. C., R. R. Gil, et al. (2005). “1H NMR and Optical Spectroscopic Investigation of the Sequence-Dependent Dimerization of a Symmetrical Cyanine Dye in the DNA Minor Groove.” Biochemistry 44(48): 15924-15929.

He, Z., L. Wu, et al. (2005). “Empirical Establishment of Oligonucleotide Probe Design Criteria.” Applied and Environmental Microbiology 71(7): 3753-3760.

Hörz, W. and W. Altenburger (1981). “Sequence specific cleavage of DNA by micrococcal nuclease.” Nucleic Acids Research 9(12): 2643-2658.

Hosoi-Tanabe, S. and Y. Sako (2005). “Rapid detection of natural cells of and (Dinophyceae) by fluorescence in situ hybridization.” Harmful Algae 4(2): 319-328.

Hugenholtz, P., G. Tyson, et al. (2002). Design and Evaluation of 16S rRNA-Targeted Oligonucleotide Probes for Fluorescence In Situ Hybridization. Gene Probes. M. Muro and R. Rapley, Humana Press. 179: 29-42.

Inc., T. A. O. S. (2006). TSL230R-LF, TSL230AR-LF, TSL230BR-LF. Plano, TX: 1-8.

Jensen, K. K., H. Ørum, et al. (1997). “Kinetics for Hybridization of Peptide Nucleic Acids (PNA) with DNA and RNA Studied with the BIAcore Technique.” Biochemistry 36: 5072-5077.

Jones, W. J., C. M. Preston, et al. (2008). “A robotic molecular method for in situ detection of marine invertebrate larvae.” Mol Ecol Resour 8(3): 540-550.

Karling, J. S. (1964). Synchytrium, Academic Press.

Köhler, O., D. V. Jarikote, et al. (2005). “Forced Intercalation Probes (FIT Probes): Thiazole Orange as a Fluorescent Base in Peptide Nucleic Acids for Homogeneous Single-Nucleotide-Polymorphism Detection.” ChemBioChem 6(1): 69-77.

Köhler, O. and O. Seitz (2003). “Thiazole orange as fluorescent universal base in peptide nucleic acids.” Chemical Communications 0(23): 2938-2939.

Komiyama, M., S. Ye, et al. (2003). “PNA for One-Base Differentiating Protection of DNA from Nuclease and Its Use for SNPs detection.” Journal of the American Chemical Society 125: 3758-3762.

Kuhn, H., V. V. Demidov, et al. (2002). “Hybridization of DNA and PNA Molecular Beacons to Single-Stranded and Double-Stranded DNA Targets.” Journal of the American Chemical Society 124(6): 1097-1103.

Kumar, G. N., S. Iyer, et al. (2007). “Extraction of RNA from fresh, frozen, and lyophilized tuber and root tissues.” Journal of Agricultural and Food Chemistry 55(5): 1674-1678.

Kushon, S. A., J. P. Jordan, et al. (2001). “Effect of secondary structure on the thermodynamics and kinetics of PNA hybridization to DNA hairpins.” J. Am. Chem. Soc. 123: 10805-10813.

Laidlaw, W. M. R. (1985). “A method for the detection of the resting sporangia of potato wart disease (Synchytrium endobioticum) in the soil of old outbreak sites.” Potato Research 28(2): 223-232.

Lazcka, O., F. J. Del Campo, et al. (2007). “Pathogen detection: A perspective of traditional methods and biosensors.” Biosens Bioelectron 22: 1205-1217.

Leijon, M., A. Gräslund, et al. (1994). “Structural characterization of PNA-DNA duplexes by NMR. Evidence for DNA in a B-like conformation.” Biochemistry 33: 9820-9825.

Lévesque, C. A., S. N. de Jong, et al. (2001). Molecular phylogeny and detection of Synchytrium endobioticum, the causal agent of potato wart. The Canadian Phytopathological Society Annual Meeting, Ontario, Canada.

Li, X., E. Morgenroth, et al. (2008). “Quantitative rRNA-Targeted Solution-Based Hybridization Assay Using Peptide Nucleic Acid Molecular Beacons.” Applied and Environmental Microbiology 74(23): 7297-7305.

Lindahl, T. (1993). “Instability and decay of the primary structure of DNA.” Nature 362: 709-715.

Liu, W. T., H. Guo, et al. (2007). “Effects of target length on the hybridization efficiency and specificity of rRNA-based oligonucleotide microarrays.” Appl Environ Microbiol 73(1): 73-82.

Liu, X., M. Atwater, et al. (2007). “Extinction coefficient of gold nanoparticles with different sizes and different capping ligands.” Colloids Surf. B 58: 3-7.

Love, J. C., L. A. Estroff, et al. (2005). “Self-assembled monolayers of thiolates on metals as a form of nanotechnology.” Chem. Rev. 105: 1103-1169.

Loy, A., M. Horn, et al. (2003). “probeBase: an online resource for rRNA-targeted oligonucleotide probes.” Nucleic Acids Research 31(1): 514-516.

Lucarelli, F., S. Tombelli, et al. (2008). “Electrochemical and piezoelectric DNA biosensors for hybridisation detection.” Anal Chim Acta 609: 139-159.

Lundberg, K. S., D. D. Shoemaker, et al. (1991). “High-fidelity amplification using a thermostable DNA polymerase isolated from Pyrococcus furiosus.” Gene 108(1): 1-6.

Luong, J. H. T., K. B. Male, et al. (2008). “Biosensor technology: Technology push versus market pull.” Biotechnol Adv 26: 492-500.

Lytton-Jean, A. K. R., J. M. Gibbs-Davis, et al. (2009). “Highly cooperative behavior of peptide nucleic acid-linked DNA-modified gold-nanoparticle and comb-polymer aggregates.” Advanced Materials 21: 706-709.

Lytton-Jean, A. K. R., J. M. Gibbs-Davis, et al. (2009). “Highly Cooperative Behavior of Peptide Nucleic Acid-Linked DNA-Modified Gold-Nanoparticle and Comb-Polymer Aggregates.” Advanced Materials 21(6): 706-709.

McNamara, D. G. and I. M. Smith (1998). “National control measures for Synchytrium endobioticum1.” EPPO Bulletin 28(4): 507-511.

Mehlmann, M., M. B. Townsend, et al. (2005). “Optimization of fragmentation conditions for microarray analysis of viral RNA.” Anal Biochem 347(2): 316-323.

Mello, L. D. and L. T. Kubota (2002). “Review of the use of biosensors as analytical tools in the food and drink industries.” Food Chem 77: 237-256.

Metfies, K., S. Huljic, et al. (2005). “Electrochemical detection of the toxic dinoflagellate with a DNA-biosensor.” Biosensors and Bioelectronics 20(7): 1349-1357.

Metfies, K. and L. K. Medlin (2008). “Feasibility of Transferring Fluorescent In Situ Hybridization Probes to an 18S rRNA Gene Phylochip and Mapping of Signal Intensities.” Applied and Environmental Microbiology 74(9): 2814-2821.

Mirkin, C. A., R. L. Letsinger, et al. (1996). “A DNA-based method for rationally assembling nanoparticles into macroscopic materials.” Nature 382: 607-609.

Mothershed, E. A. and A. M. Whitney (2006). “Nucleic acid-based methods for the detection of bacterial pathogens: Present and future considerations for the clinical laboratory.” Clinica Chimica Acta 363: 206-220.

Mulvaney, P. (1996). “Surface plasmon spectroscopy of nanosized metal particles.” Langmuir 12: 788-800.

Murphy, D., G. Redmond, et al. (2004). “Hybridization and melting behavior of peptide nucleic acid (PNA) oligonucleotide chimeras conjugated to gold nanoparticles.” Helv. Chim. Acta 87: 2727-2734.

Nelson, B. P., M. R. Liles, et al. (2002). “Label-free detection of 16S ribosomal RNA hybridization on reusable DNA arrays using surface plasmon resonance imaging.” Environmental Microbiology 4(11): 735-743.

Nielsen, P. E., M. Egholm, et al. (1991). “Sequence-selective recognition of DNA by strand displacement with a thymine-substituted polyamide.” Science 254(5037): 1497-1500.

Niepold, F. and H. Stachewicz (2004). “PCR-detection of Synchytrium endobioticum (Schilb.) Perc.” Journal of Plant Diseases and Protection 111(4): 313-321.

O’Connor, L. and B. Glynn (2010). “Recent advances in the development of nucleic acid diagnostics.” Expert Review of Medical Devices 7(4): 529-539.

Ørum, H. (2000). “PCR clamping.” Current Issues in Molecular Biology 2(1): 27-30.

Ørum, H., P. E. Nielsen, et al. (1993). “Single base pair mutation analysis by PNA directed PCR clamping.” Nucleic Acids Research 21(23): 5332-5336.

Ørum, H., P. E. Nielsen, et al. (1995). “Sequence-specific purification of nucleic acids by PNA-controlled hybrid selection.” BioTechniques 19(3): 472-480.

Palchetti, I. and M. Mascini (2008). “Electroanalytical biosensors and their potential for food pathogen and toxin detection.” Anal Bioanal Chem 391: 455-471.

Park, H., A. Germini, et al. (2007). “Effect of ionic strength on PNA-DNA hybridization on surfaces and in solution.” Biointerphases 2(2): 80.

Peplies, J., F. O. Glöckner, et al. (2003). “Optimization Strategies for DNA Microarray-Based Detection of Bacteria with 16S rRNA-Targeting Oligonucleotide Probes.” Applied and Environmental Microbiology 69(3): 1397-1407.

Peplies, J., C. Lachmund, et al. (2006). “A DNA Microarray Platform Based on Direct Detection of rRNA for Characterization of Freshwater Sediment-Related Prokaryotic Communities.” Applied and Environmental Microbiology 72(7): 4829-4838.

Perry-O’Keefe, H., S. Rigby, et al. (2001). “Identification of indicator microorganisms using a standardized PNA FISH method.” Journal of Microbiological Methods 47(3): 281-292.

Perry-O’Keefe, H., H. Stender, et al. (2001). “Filter-based PNA in situ hybridization for rapid detection, identification and enumeration of specific micro-organisms.” Journal of Applied Microbiology 90(2): 180-189.

Petersen, K., U. Vogel, et al. (2004). “Short PNA molecular beacons for real-time PCR allelic discrimination of single nucleotide polymorphisms.” Molecular and Cellular Probes 18(2): 117-122.

Peterson, A. W., R. J. Heaton, et al. (2001). “The effect of surface probe density on DNA hybridization.” Nucleic Acids Res. 29: 5163-5168.

Pozhitkov, A., P. A. Noble, et al. (2006). “Tests of rRNA hybridization to microarrays suggest that hybridization characteristics of oligonucleotide probes for species discrimination cannot be predicted.” Nucleic Acids Research 34(9): e66.

Rasmussen, H., J. S. Kastrup, et al. Crystal structure of a peptide nucleic acid (PNA) duplex at 1.7 A resolution, Nat Struct Biol. 1997 Feb;4(2):98-101.

Ren, B., J.-M. Zhou, et al. (2004). “Straightforward detection of SNPs in double-stranded DNA by using exonuclease III/nuclease S1/PNA system.” Nucleic Acids Research 32(4): e42.

Rose, D. J. (1993). “Characterization of Antisense Binding Properties of Peptide Nucleic Acids by Capillary Gel Electrophoresis.” Analytical Chemistry 65(24): 3545-3549.

Rosi, N. L. and C. A. Mirkin (2005). “Nanostructures in biodiagnostics.” Chem. Rev. 105: 1547-1562.

Russ Algar, W., M. Massey, et al. (2009). “The application of quantum dots, gold nanoparticles and molecular switches to optical nucleic-acid diagnostics.” TrAC Trends in Analytical Chemistry 28(3): 292-306.

SantaLucia, J., Jr., H. T. Allawi, et al. (1996). “Improved nearest-neighbor parameters for predicting DNA duplex stability.” Biochemistry 35(11): 3555-3562.

Sassolas, A., B. D. Leca-Bouvier, et al. (2008). “DNA biosensors and microarrays.” Chem Rev 108: 109-139.

Sato, Y., K. Fujimoto, et al. (2003). “Detection of a K-ras point mutation employing peptide nucleic acid at the surface of a SPR biosensor.” Colloids and Surfaces B: Biointerfaces 27(1): 23-31.

Schmitt, J., P. Machtle, et al. (1999). “Preparation and optical properties of colloidal gold monolayers.” Langmuir 15: 3256-3266.

Scholin, C. A., M. Herzog, et al. (1994). “Identification of group- and strain-specific genetic markers for globally distributed Alexandrium (Dinophyceae). II. Sequence analysis of a fragment of the LSU rRNA gene.” Journal of Phycology 30: 999-1011.

Scholin, C. A., R. Marin, et al. (1999). “DNA probes and a receptor-binding assay for detection of Pseudo-nitzschia (Bacillariophyceae) species and domoic acid activity in cultured and natural samples.” Journal of Phycology 35(6): 1356-1367.

Seed, B. (2001). Silanizing Glassware. Current Protocols in Cell Biology, John Wiley & Sons, Inc.

Seifert, J. L., R. E. Connor, et al. (1999). “Spontaneous assembly of helical cyanine dye aggregates on DNA nanotemplates.” Journal of the American Chemical Society 121: 2987-2995.

Seitz, O. (2000). “Solid-Phase Synthesis of Doubly Labeled Peptide Nucleic Acids as Probes for the Real-Time Detection of Hybridization.” Angewandte Chemie International Edition 39(18): 3249-3252.

Small, J., D. R. Call, et al. (2001). “Direct Detection of 16S rRNA in Soil Extracts by Using Oligonucleotide Microarrays.” Applied and Environmental Microbiology 67(10): 4708-4716.

Smith, J. O., D. A. Olson, et al. (1999). “Molecular Recognition of PNA-Containing Hybrids:  Spontaneous Assembly of Helical Cyanine Dye Aggregates on PNA Templates.” Journal of the American Chemical Society 121(12): 2686-2695.

Socher, E., L. Bethge, et al. (2008). “Low-Noise Stemless PNA Beacons for Sensitive DNA and RNA Detection.” Angewandte Chemie International Edition 47(49): 9555-9559.

Socher, E., D. V. Jarikote, et al. (2008). “FIT probes: Peptide nucleic acid probes with a fluorescent base surrogate enable real-time DNA quantification and single nucleotide polymorphism discovery.” Analytical Biochemistry 375(2): 318-330.

Spaulding, D. C. (2011). “Peptide Nucleic acids: Robust Probe Hybridization Technology.” 2013, from http://www.biosyn.com/tew/Peptide-Nucleic-acids-Robust-Probe-Hybridization-Technology.aspx.

Standards, S. o. P. H. D. (2011). National Diagnostic Protocol for Potato Wart, caused by Synchytrium endobioticum. D. o. A. Australian Government, Fisheries and Forestry.

Stanley Electric Co., L. (2010). 1211C Series Bi-color Type (1.6 X 1.5 mm): 9-10.

Svanvik, N., A. Stahlberg, et al. (2000). “Detection of PCR products in real time using light-up probes.” Analytical Biochemistry 287(1): 179-182.

Svanvik, N., G. Westman, et al. (2000). “Light-up probes: thiazole orange-conjugated peptide nucleic acid for detection of target nucleic acid in homogeneous solution.” Analytical Biochemistry 281(1): 26-35.

Tedeschi, T., S. Sforza, et al. (2007). “Fast and easy colorimetric tests for single mismatch recognition by PNA–DNA duplexes with the diethylthiadicarbocyanine dye and succinyl-β-cyclodextrin.” Journal of Biochemical and Biophysical Methods 70(5): 735-741.

Teles, F. R. R. and L. P. Fonseca (2008). “Trends in DNA biosensors.” Talanta 77: 606-623.

Tomac, S., M. Sarkar, et al. (1996). “Ionic Effects on the Stability and Conformation of Peptide Nucleic Acid Complexes.” Journal of the American Chemical Society 118: 5544-5552.

Tomlinson, A., B. Frezza, et al. (2006). “A structural model for cyanine dyes templated into the minor groove of DNA.” Chemical Physics 325(1): 36-47.

Touzet, N., J. M. Franco, et al. (2007). “Characterization of nontoxic and toxin-producing strains of Alexandrium minutum (Dinophyceae) in Irish coastal waters.” Applied and Environmental Microbiology 73(10): 3333-3342.

Tyagi, S. and F. R. Kramer (1996). “Molecular beacons: probes that fluoresce upon hybridization.” Nature Biotechnology 14(3): 303-308.

van den Boogert, P. H. J. F., M. P. E. Gent-Pelzer, et al. (2005). “Development of PCR-based Detection Methods for the Quarantine Phytopathogen Synchytrium endobioticum, Causal Agent of Potato Wart Disease.” European Journal of Plant Pathology 113(1): 47-57.

Van Dolah, F. M. (2000). “Marine algal toxins: Origins, health effects, and their increased occurrence.” Environmental Health Perspectives 108(Supplement 1): 133-141.

van Gent-Pelzer, M. E., M. Krijger, et al. (2010). “Improved real-time PCR assay for detection of the quarantine potato pathogen, Synchytrium endobioticum, in zonal centrifuge extracts from soil and in plants.” European Journal of Plant Pathology 126(1): 129-133.

Vercoutere, W. and M. Akeson (2002). “Biosensors for DNA sequence detection.” Curr Opin Chem Biol 6: 816-822.

von Hippel, P. H. and G. Felsenfeld (1964). “Micrococcal Nuclease as a Probe of DNA Conformation.” Biochemistry 3(1): 27-39.

Wilhelmsson, L. M., B. Nordén, et al. (2002). “Genetic screening using the colour change of a PNA-DNA hybrid-binding cyanine dye.” Nucleic Acids Research 30: e3.

Williams, B., H. Stender, et al. (2002). PNA fluorescent in situ hybridization for rapid microbiology and cytogenetic analysis. Methods in Molecular Biology. P. E. Nielsen. Totowa, Humana Press Inc.: 181-194.

Wittung, P., S. K. Kim, et al. (1994). “Interactions of DNA binding ligands with PNA-DNA hybrids.” Nucleic Acids Research 22(24): 5371-5377.

Wittung, P., P. E. Nielsen, et al. (1994). “DNA-like double helix formed by peptide nucleic acid.” Nature 368(6471): 561-563.

Woese, C. R. (1987). “Bacterial evolution.” Microbiological Reviews 51(2): 221-271.

Worden, A. Z., S. W. Chisholm, et al. (2000). “In situ hybridization of Prochlorococcus and Synechococcus (marine cyanobacteria) spp. with RRNA-targeted peptide nucleic acid probes.” Applied and Environmental Microbiology 66(1): 284-289.

Xi, C., M. Balberg, et al. (2003). “Use of DNA and Peptide Nucleic Acid Molecular Beacons for Detection and Quantification of rRNA in Solution and in Whole Cells.” Applied and Environmental Microbiology 69(9): 5673-5678.

Xu, J. and S. L. Craig (2005). “Thermodynamics of DNA hybridization on gold nanoparticles.” J. Am. Chem. Soc. 127: 13227-13231.

Ye, S., Y. Miyajima, et al. (2007). “Combination of peptide nucleic acid beacon and nuclease S1 for clear-cut genotyping of single nucleotide polymorphisms.” Analytical Biochemistry 363(2): 300-302.

##### SENSITIVITY ANALYSIS OF COLORIMETER PERFORMANCE

Major parameters affecting colorimeter design

Recall from Chapter 4 that the raw colorimeter hybridization signal is given by:

H=AgreenAred				        Equation A.1

The absorbance values, meanwhile, are calculated directly from the photodetector output.  The photodetector is a light-to-frequency converter, so incident illumination at the photodetector surface is measurable at the output as frequency values.  However, the silicon-based photodetector is more sensitive to longer (red) wavelengths (Figure 4.5).  In addition, the red LED is more powerful than the green LED.  Thus, the green LED is driven with the maximum forward current value, IF≈ 25 mA, and the forward current of the red LED is adjusted until the irradiance, E (in µW/cm2) seen by the photodetector is equivalent.  That is,

Eblankred=Eblankgreen				Equation A.2

Where the irradiance from each LED is calculated as

Esamplegreen=fsamplegreenrgreen				 Equation A.3a

Esamplered=fsampleredrred				 Equation A.3b

The frequency output at the photodetector from each LED is scaled by the factor r which is a photodiode responsivity factor as given in the photodetector datasheet.  For the TSL230BRD package selected, rgreen=671.5 HzμWcm2 and rred=829.5 HzμWcm2.   After this conversion to irradiance (E), sample absorbances for at green and red wavelengths are calculated as:

Agreen=-logEsamplegreenEblankgreen			Equation A.4a

Ared=-logEsampleredEblankred			Equation A.4b

Plugging this in to Equation A.1 gives

H=-logEsamplegreenEblankgreen-logEsampleredEblankred				Equation A.5

Basic concepts in error propagation

Given a function

Z=F(a,b,c,…)

The variance of z, σZ, can be approximated by the variances in each function variable and the partial derivatives of the function with respect to each variable:

σZ2=σα2*δfδα2+σβ2*δfδβ2+σγ2*δfδγ2+…	       Equation A.6

Also, the uncertainty in Z, ∆Z, is defined as:

∆Z=σZ2			 	      Equation A.7

The signal-to-noise ratio (SNR) of the output of this system is defined as:

SNRZ=Z∆Z	   			        Equation A.8

The sensitivity of the output of this system to the change in an individual input under normal operating parameters (NOP) can then be defined as:

SinputSNR(Z)|NOP=δδinput(SNRZ)		     Equation A.9

Analytical calculation of error propagation in colorimeter system

The uncertainty in H, ∆H, is given by combining Equation A.6 and Equation A.7:

∆H=σH2=σa2*δfδa2+σb2*δfδb2+σc2*δfδc2+σd2*δfδd2     Equation A.10

Where placeholders are defined as follows to simplify symbolic calculations:

a≡Psamplegreen     e≡σa

b≡Pblankgreen     f≡σb

c≡Psamplered     g≡σc

d≡Pblankred     h≡σd

Symbolic calculations were performed using Mathematica 9 (Wolfram Research, Champaign, IL, USA).  The most general form of the analysis assumes that the light power for both green and red LEDs is free to vary; that is, Equation A.2 is not true.  The uncertainty ∆H is found to be:

a2b2(d2g2+c2h2)Log[ab]2+c2d2(b2e2+a2f2)Ln[cd]2a2b2c2d2Ln[cd]4

The signal-to-noise ratio (SNR) is defined as:

SNR(H)=H∆H

SNRH=Ln[ab]Ln[cd]a2b2(d2g2+c2h2)Ln[ab]2+c2d2(b2e2+a2f2)Ln[cd]2a2b2c2d2Ln[cd]4

The sensitivity of the output to a change in the emitted light intensity from the red LED is given by:

SdSNR(H)|NOP=δδd(SNRH)

SdSNR(H)|NOP=-Ln[ab]3(d2g2+c2h2-c2h2Ln[cd])c2d3Ln[cd]6(a2b2(d2g2+c2h2)Ln[ab]2+c2d2(b2e2+a2f2)Ln[cd]2a2b2c2d2Ln[cd]4)32

For a change in the emitted light intensity of the green LED, the analytical solution is:

SbSNR(H)|NOP=δδb(SNRH)

SbSNR(H)|NOP=-b2e2+a2f2-a2f2Ln[ab]a2b3Ln[cd]3(a2b2(d2g2+c2h2)Ln[ab]2+c2d2(b2e2+a2f2)Ln[cd]2a2b2c2d2Ln[cd]4)32

Error propagation under normal operating parameters

Assuming that each sample is perfectly made, the measurement uncertainties are dependent only upon the photodetector and LEDs.  From datasheet values, the light intensity of the green LED is typically 85 mcd but can vary by 71%, while that of the red LED is typically 2500 mcd but can vary by 60%.  At operating conditions, the datasheet values of the LEDs are ~127.5 mcd and ~187.5 mcd for the green and red, respectively.  The photodetector output at 10X sensitivity and 2X scaling is in the range of 10 kHz, with a tolerance of 10%.  The uncertainties σx were calculated from Equation A.6 and simplify to:

σx2x2=σα2α2+σβ2β2+σγ2γ2+…		        Equation A.11

The normal operating parameters given in Table A.1 were used to calculate the sensitivity from the red and green LED systems (last two table entries, respectively).

Table A.1. Normal operating parameters for colorimetric system. The values for the liquid sample are for 1 µM PNA probe, 250 nM complementary synthetic RNA, and 10 µM dye, and represent the average of three separate reads.

| Parameter | Equation placeholder | Typical experimental value, E (µW/cm2) |
| --- | --- | --- |
| Esamplegreen | a | 5.8163 |
| Eblankgreen | b | 10.7535 |
| Esamplered | c | 3.4968 |
| Eblankred | d | 10.7209 |
| σa | e | 5.665 |
| σb | f | 5.665 |
| σc | g | 0.2339 |
| σd | h | 0.2339 |
| SredSNR(H)¦NOP | SdSNR(H)¦NOP | 62.07E-6 |
| SgreenSNR(H)¦NOP | SbSNR(H)¦NOP | 0.0955 |

It is readily apparent that the green LED system, with an absolute sensitivity ~1500 times greater than that of the red LED system, has the most impact on the output of the colorimeter.  A more powerful green light source (such as laser diode) could be explored as a replacement to improve colorimeter performance.

Expressions used in Mathematica

H[a_,b_,c_,d_]≔-Log10[ab](-Log10[cd])

delH[a_,b_,c_,d_,e_,f_,g_,h_]≔Sqrt[e^2*(D[H[a,b,c,d],a])^2+f^2*(D[H[a,b,c,d],b])^2+g^2*(D[H[a,b,c,d],c])^2+h^2*(D[H[a,b,c,d],d])^2]

SNR[a_,b_,c_,d_,e_,f_,g_,h_]:=H[a,b,c,d]delH[a,b,c,d,e,f,g,h]

Smon[a_,b_,c_,d_,e_,f_,g_,h_]:=D[SNR[a,b,c,d,e,f,g,h],d]

Sagg[a_,b_,c_,d_,e_,f_,g_,h_]:=D[SNR[a,b,c,d,e,f,g,h],b]

##### CHARACTERIZATION OF DNA AND RNA EXTRACTED FROM POTATO WART SAMPLES

Table B.1. Genomic DNA extracted from potato wart samples.

| DNA sample | Target copy number/ng | Cq | Wart sample |
| --- | --- | --- | --- |
| 1 | 4.51E+03 | 23.47 | 2 X 900 µL Avondale wart 09/16/08 overnight rehydrated (R2 03/14/12), MasterPure Kit |
| 2 | 1.37E+05 | 18.66 | 2 X 160 µg frozen wart 03/14/12 |
| 3 | 1.88E+04 | 21.46 | 2 X 900 µL R2 Zymo |
| 4 | 4.07E+05 | 17.13 | 4 X 0.2 g Avondale frozen wart 09/17/09, mortar & pestle 03/14/12 |
| 5 | 5.29E+05 | 16.76 | 4 X 0.2 g 03/15/12 |
| 6 | 3.78E+04 | 20.48 | 6 X 1 mL F1 03/15/12 |
| 7 | 0 | NA | 3 X 0.2 g Arran Victory wart 03/16/12 |
| 8 | 9.68E+04 | 19.18 | 0.2 g Exploits wart |
| 9 | 2.64E+05 | 17.74 | 0.2 g Pink Pearl wart |

Table B.2. Genomic RNA extracted from potato wart samples.

| RNA sample | Well #s | Sample | Conc (ng/µL) | 260/230 | 260/280 | Target copy # (qPCR) | 30 ng/µL copy # |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A | A1-A5, B9-B12 | 9 x 20 µL R2 03/15 p. 92 | 41.07 | 0.10 | 1.55 | 0.868e3 | 0.521e6 |
| B | B1-B8, C11-C12 | 10 x 20 µL D2 03/15 p. 92 | 50.99 | 0.85 | 1.59 | 1.444e3 | 0.865e6 |
| C | C1-C10 | 10 x 20 µL D3 03/15 p. 92 | 182.01 | 1.29 | 1.89 | 1.697e4 | 10.2e6 |
| D | D1-D5, E12 | 6 x 20 µL D1 03/15 p. 92 | 104.21 | 1.35 | 1.83 | 4.177e3 | 2.51e6 |
| E | D6-D12 | 6 x 20 µL D3 03/15 p. 92 | 128.15 | 1.46 | 1.95 | 1.525e4 | 9.15e6 |
| F | E1 | 1 x 96 µL frozen wart 03/14 p. 89 | 797.55 | 1.31 | 1.79 | 6.903e4 | 41.41e6 |
| G | E2-E11 | 10 x 20 µL D1 03/15 p. 91 | 133.55 | 1.42 | 1.89 | 1.183e4 | 7.1e6 |
| H | F1-F2, G5-G12 | 10 x 20 µL D1 03/14 p. 89 | 51.77 | 1.18 | 1.78 | 2.253e4 | 13.52e6 |
| I | F3-F7 | 5 x 20 µL D1 03/14 p. 89 | 75.28 | 1.43 | 1.88 | 3.642e4 | 21.9e6 |
| J | F8-F12 | 5 x 20 µL D1 03/14 p. 89 | 83.14 | 1.22 | 1.81 | 2.12e4 | 12.72e6 |
| K | G1-G4, H2-H12 | 15 x 50 µL R2 03/13 p. 87 | 40.05 | 0.28 | 1.55 | 3.011e3 | 1.8e6 |
| L | Tube |  | 88.87 | 1.89 | 1.88 |  |  |

##### EFFECTS OF RNASTABLE® STORAGE MATRIX ON HYBRIDIZATION USING DiSC2(5)

This appendix details the experiments and data analysis associated with deriving hybridization signals for the DiSC2(5) colorimetric assay in the presence of rehydrated RNAstable®.

RNAstable® contains a pH indicator that turns pink in basic solutions.  This color corresponds to a concentration-dependent absorbance peak at ~560 nm, as shown in Figure B.1.  More importantly, the absorbance characteristics of RNAstable® directly affect the calculation of the DiSC2(5) hybridization signal due to non-negligible absorbance values at ~540 nm (Figure B.1b).  The solutions do not exhibit any significant absorbance at the dye monomer peak of 650 nm.

| a) | b) |
| --- | --- |

Figure B.1. Concentration-dependent RNAstable® absorbance characteristics. a) Absorbance spectra of RNAstable®, 30-90% in water. b) Concentration-dependent values of RNAstable® at the reagent absorbance peak (~560 nm) and at the DiSC2(5) hybridization peak (~540 nm).  These concentrations of RNAstable® were chosen to reflect the range of concentrations in the genomic RNA samples.

The effect of RNAstable® on the colorimetric hybridization signals was tested with both potato RNA and in vitro transcribed potato wart RNA (from genomic DNA sample 8).  PNA 033 was hybridized to 30 ng/µL RNA as in section 5.3.8, except that RNAstable® (30-90%) was mixed with the RNA samples prior to hybridization.  Hybridization signals were calculated as the ratio A541/A649 after the concentration-dependent RNAstable® absorbance values at these wavelengths were subtracted from the raw data.  These hybridization signals are shown in Figure B.2.

| a) | b) |
| --- | --- |

Figure B.2. Raw and corrected hybridization signals for potato and potato wart RNA.  a) Raw and b) % RNAstable®-corrected hybridization signals.  The corrected signals were calculated by first subtracting the RNAstable®-dependent absorbance values at 541 nm and 649 nm before calculating the ratio A541/A649.  The dashed line in b) is the highest signal from potato RNA at all concentrations.

The dashed line in Figure B.2b denotes the highest signal from potato RNA, which is the host organism for Synchytrium endobioticum and is the main contaminant in nucleic acids extracted from potato warts (from PCR and sequencing data).  This threshold was used to identify samples as either S. endobioticum-positive or negative.

##### ALEXANDRIUM CELL COUNTING PROTOCOL

Revised 05-Mar-12

Equipment & supplies needed

Sedgewick-Rafter counting chamber

Glass cover slip for counting chamber

1.7 mL polypropylene microcentrifuge tube

Tube rack

1000 µL pipette & 1000 µL pipette tips

10 µL pipette & 10 µL pipette tips

Lugol’s iodine solution

Light microscope

Alexandrium culture to be counted

Procedure (for each count)

- Take out 1 microcentrifuge tube and place into tube rack.

- Take out Sedgewick-Rafter chamber and place cover slip at an angle on the slide so that only 2 small inlets to the chamber are available (see below).

cover slip

- Set 1000 µL pipette as high as it will go (typically ~1050 µL).

- Get culture from incubator and swirl GENTLY to ensure all cells are mixed.

- Place a clean tip onto the 1000 µL pipette and pipette the cell culture into the microcentrifuge tube.  Dispose of tip into autoclave trash.

- Return culture to incubator.

- Place a clean tip onto the 10 µL pipette and pipette 10 µL of the Lugol’s solution into the microcentrifuge tube.  Close the tube and mix by vortexing gently.

- Use the 1000 µL pipette to transfer the Lugol’s-preserved culture into the counting chamber through one of the open areas, ensuring that no air bubbles are formed.

- Carefully slide the cover slip over the chamber.  Make sure that no air bubbles are trapped and that the chamber is completely filled.

- Count cells with the light microscope at 10X magnification.  Count AT LEAST 400 cells (if culture is very dense) or 100 squares (1 µL per square, 50 squares on long edge, 20 squares on short edge).  If the cell number is < 2000, count the whole chamber.

##### ALEXANDRIUM RNA EXTRACTION PROTOCOL USING CHEMICAL LYSIS

Revised 08-Mar-12

Equipment & supplies needed

25 mm filter assembly, autoclaved

2 blunt-ended tweezers, autoclaved

Millipore SVLP 5 µm pore size 25 mm diameter filter membrane

2 Sterile 3 mL disposable syringes with attached needle

2 mL polypropylene microcentrifuge tubes

Tube rack

1000 µL pipette & 1000 µL pipette tips

GuSCN lysis buffer

85°C heat block

Lab timer

Sterile 3 mL syringe with needle and Luer-Lok™ connector

0.45 µm pore size, 25 mm diameter sterile filter assembly with Luer-Lok™  connector

Alexandrium culture for extraction

Algal liquid waste container

GuSCN liquid waste container

Algal/GuSCN solid waste container

Sharps container

10% bleach

Procedure (for each extraction)

*** Preheat the heat block at least 1 hour before starting extraction***

- Clean work surface with 10% bleach.

- Take out 1 microcentrifuge tube and place into tube rack.  Use the 1000 µL pipette and clean tips to add 1.5 mL GuSCN lysis buffer to the tube (3 x 500 µL).  Recap tube and dispose of tips in the solid waste container.

- Remove the filter assembly from the autoclave bag and place on a clean surface.

- Use the tweezers to take one filter membrane from the filter container.  Check to make sure that only the filter has been picked up (no blue paper).

- Carefully place the filter onto the flat surface of the filter assembly.  Screw on the top of the assembly, making sure that the filter is centered and flat.

- Take cell culture from the incubator.

- Unwrap a clean 3 mL syringe with needle.

- Swirl the culture gently to mix the cells.

- Draw 1 mL of the culture into the syringe.

- Remove the needle from liquid and draw air into the syringe (5-10 mL of air).  Recap the culture flask.

- Keep the needle pointing up (to prevent liquid from running out of the syringe).  Remove the needle from the syringe by unscrewing it using the protective plastic cover.  Dispose of needle in the sharps container.

- Still keeping the syringe tip up, attach the filter assembly.  Make sure all the connections are tight to prevent the liquid from leaking.

- Place the syringe on top of the algal waste container and push the liquid through.  Push air (5-10 mL) through the filter as well.  If the syringe is too small to draw 10 mL of air through, remove the filter assembly, draw air through the syringe, then reconnect the filter assembly and push air through the membrane.  DO NOT TRY TO DRAW AIR THOUGH THE FILTER ASSEMBLY.

- Remove the filter assembly from the syringe.  Dispose of syringe in solid waste container.

- Open up filter assembly.  Use both tweezers to grasp the filter edges, then roll filter (cell side facing in) and place into microcentrifuge tube with GuSCN buffer.  Make sure tube is capped tightly.

- Shake tube to mix cells into the buffer.

- Place tube into heat block and start timer.

- After 2.5 minutes, remove tube from heat block and flick 6 times.  Replace into heat block.

- After 5 minutes, remove tube from heat block and flick 6 times.  Place into tube holder.

- Place a clean 2 mL tube into the tube rack.

- Unwrap a clean 3 mL syringe with needle.

- Draw all of the lysate out of the first 2 mL tube into the syringe.  Draw extra air into the syringe.  Dispose of tube in the solid waste container.

- Keep the syringe tip up and remove the needle.  Dispose of needle in the sharps container.

- Keep the syringe tip up and attach the 0.45 µm filter assembly to the syringe.  Make sure the connection is tight.

- Place the syringe over the clean 2 mL tube and push the lysate through.  Discard syringe and filter into the solid waste container.

- Lysate is ready to use.

##### PCR AMPLIFICATION OF 18S rRNA GENE FOR SUBSEQUENT IN VITRO TRANSCRIPTION

Revised 30-May-12

Goals

- PCR-amplify virtually the entire 18S rRNA gene

- Set up for the production of sense strand RNA by appending the T7 promoter on the forward primer

- Make and concentrate enough DNA to have 1 µg in 8 µL for the IVT reaction

Equipment & supplies needed

12.5 µL Phusion Hot Start Flex 2X Master Mix aliquoted in 0.2 mL nuclease-free polypropylene PCR tubes

Micropipettes & aerosol barrier pipette tips

Tube racks for 0.2 mL & 1.5 mL tubes

1.5 mL nuclease-free polypropylene microcentrifuge tubes

Thermocycler with heated lid (use the PTC-200)

PCR hood (make sure to sign up for the hood)

NanoDrop ND-1000 (sign up for the machine before running the experiment)

Thermo Fisher B1A gel box & voltage supply (sign up for running the gel)

Chemicals & reagents

Nuclease (RNase and DNase) decontamination solution

10 µM T7-PNS1 DNA (forward) primer

10 µM NS8 DNA (reverse) primer

Genomic DNA for PCR amplification

Molecular biology-grade water (MB-grade water)

1.8% IDNA agarose gel

SYBR Safe gel stain

6X gel loading dye (NEB)

1 kb DNA ladder (NEB #N3232S), pre-mixed with 6X gel loading dye

Make sure to wipe EVERYTHING down with nuclease decontamination solution before AND during the procedure

Wizard SV Gel & PCR Clean-Up System (Promega #A9281)

Procedure

- Take out 10 Phusion master mix aliquots from -20°C freezer, thaw in a nuclease-free 0.2 mL tube rack at room temperature.

- Take out 10 µM DNA primers from refrigerator and place into a nuclease-free tube rack.

- Take out genomic DNA from refrigerator and place into the tube rack with primers.

Make sure you know the concentration AND quality of the genomic DNA to be used.  Dilute your working DNA stock to 10 ng/µL in MB-grade water (NOT TE) as all the numbers in the table below assume this concentration.

- Clean out the PCR hood and pipettes.  Transfer primer, gDNA, and polymerase tubes to the PCR hood.

- Prepare and label 25 µL PCR samples:

| Sample | Phusion® 2X master mix (µL) | 10 µM T7-PNS1 (µL) | 10 µM NS8 (µL) | 10 ng/µL gDNA (µL) | MB-grade water (µL) |
| --- | --- | --- | --- | --- | --- |
| - DNA + primers | 12.5 | 1.25 | 1.25 | 0 | 10 |
| + DNA – primers | 12.5 | 0 | 0 | 5 | 7.5 |
| + DNA + primers (x8) | 12.5 | 1.25 | 1.25 | 5 | 5 |

- Check that the PTC-200 program 18S-PHU is set to the following, for 35 cycles:

| Step | Temperature (°C) | Time (mm:ss) |
| --- | --- | --- |
| Initial denaturation | 98 | 00:45 |
| Denaturation | 98 | 00:10 |
| Annealing | 70 | 00:30 |
| Extension | 72 | 01:15 |
| Final extension | 72 | 10:00 |

- Run the program (should take less than 2 hours).

- While waiting (during the last hour of the program), start making a 1.8% IDNA agarose gel with SYBR Safe according to lab protocol.  Let the gel solidify for 30-45 minutes at room temperature.

- Once the PCR program is done, pool the 8 identical PCR samples into 1 tube.

- Run all of the unpurified PCR products on the gel (6 µL product + 2 µL 6X gel loading dye).  Use 6 µL of the 1 kb ladder.  The expected band size is 1800 bp.

- IF the gel confirms the generation of correct products, purify the pooled sample with the Wizard SV Gel and PCR Clean-Up System according to manufacturer’s instructions.  Elute the DNA in 30 µL of MB-grade water.

- Check the DNA concentration and quality using the NanoDrop.  In vitro transcription requires 1 µg of template DNA in 8 µL, so the MINIMUM acceptable DNA concentration from the pooled sample is 125 ng/µL.

##### IN VITRO TRANSCRIPTION OF 18S rRNA GENE USING PCR-AMPLIFIED dsDNA TEMPLATES

Revised 03-Sep-12

Goals

- Transcribe the 18S rRNA gene into sense strand RNA

- Purify the RNA for subsequent hybridization reactions

- Verify the production of desired RNA transcript

Equipment & supplies needed

Styrofoam ice container and crushed ice

Micropipettes & aerosol barrier pipette tips

Tube racks for 0.2 mL & 1.5 mL tubes

1.5 mL nuclease-free polypropylene microcentrifuge tubes

Thermocycler with heated lid (use the PTC-200)

Microcentrifuge

NanoDrop ND-1000 (sign up for the machine before running the experiment)

FlashGel RNA cassette and dock (sign up for running the gel)

Chemicals & reagents

Nuclease (RNase and DNase) decontamination solution

T7 High Yield RNA Synthesis Kit (NEB, stored at -20°C)

RNAsecure (Ambion)

5M ammonium acetate (RNase-free, Ambion)

70% ethanol (RNase-free)

137.36 ng/µL PCR-amplified DNA for IVT (has T7 promoter sequence from forward primer)

Molecular biology-grade water (MB-grade water)

FlashGel RNA ladder (Lonza, stored at -80°C)

FlashGel formaldehyde sample buffer (Lonza)

Make sure to wipe EVERYTHING down with nuclease decontamination solution before AND during the procedure

Procedure

Part I. In Vitro Transcription

- Wipe down lab bench, equipment, and pipettes with nuclease decontamination solution.

- Take out T7 High Yield RNA Synthesis Kit components, thaw on ice.

- Take out PCR products from refrigerator, vortex and spin down.  Place tube(s) in tube rack.

Make sure you know the concentration AND quality of the DNA to be used.  Dilute your PCR products to 137.36 ng/µL in MB-grade water as all the numbers in the table below assume this concentration.

- Prepare 18 µL samples in 0.2 mL tubes as shown below (volumes are given for EACH sample):

| Component | Volume (µL) |
| --- | --- |
| 10X reaction buffer | 2 |
| ATP | 2 |
| GTP | 2 |
| UTP | 2 |
| CTP | 2 |
| PCR product with T7 promoter | 7.28 |
| RNAsecure | 0.72 |

- Heat the sample(s) at 60°C for 10 minutes in the thermocycler.

- Take out samples (keep on ice) and add 2 µL of the T7 RNA Polymerase mix.

- Return samples to the thermocycler and incubate at 37°C for 2 hours.

If necessary, reaction products may be stored at -80°C for up to 48 hoursByrne 2001 (I have not tested this)

Part II. Ammonium acetate precipitation of transcribed RNAEpicentre

- Add 1 volume 5M ammonium acetate to reaction product in 0.2 mL tube (20 µL per reaction), mix by pipetting.

- Transfer mixture to nuclease-free 1.5 mL microcentrifuge tubes, vortex to mix.

- Chill mixture on ice for 15 min.

- Centrifuge tubes at 10000 x g for 15 min.

- Carefully remove the supernatant with the 20 µL pipette.

The pellet should be white and easy to see at the bottom of the tube.

- Wash the pellet 5 times with 200 µL 70% ethanol.

Don’t suck up the pellet!!

- Dry the pellet at room temperature for ~30 min.

Make sure that the pellet is dry (no ethanol remaining)

- Resuspend the pellet in 50 µL of the RNA Storage Solution.

Part III. RNA quality check

- Check the quantity and quality of the transcribed RNA using the NanoDrop ND-1000.

Use the RNA Storage Solution as the blank.

- Assuming the RNA quality is good (acceptable A260/A230 and A260/A280), proceed with checking transcript size using the FlashGel RNA system:

- Dilute the RNA to 80 ng/µL (5 µL volume) in MB-grade water in 0.2 mL PCR tubes.

Maximum of 200 ng per band in a 5 µL load.

- Take out the FlashGel formaldehyde sample buffer from the -20°C freezer.

- Add 5 µL of the formaldehyde sample buffer to each 5 µL RNA sample (including ladder/marker) and mix well.

- Heat the mixture at 65°C for 5 minutes in the thermocycler.  Remove and keep on ice while loading the gel.

- While samples are heating, take out a FlashGel RNA cassette and remove the protective sticker.

- Flood the sample wells with MB-grade water, and pour off excess into the sink.

- Load 5 µL of each RNA sample into a separate well in the cassette.

- Run the cassette at 225 V for ~8 minutes.

It may be helpful to run a DNA ladder with the RNA samples as the RNA bands will disappear at ~4 minutes.

- Disconnect the FlashGel dock from the voltage source and cover the gel for 20-45 minutes.

- Photograph FlashGel.

##### PNA HYBRIDIZATION TO SYNTHETIC RNA SEQUENCES: QUALITY CONTROL

Revised 30-May-12

Goals

- Proficiency in setting up and performing colorimetric assay

- Assay reproducibility: standard deviation of triplicate samples done on the same day must be ≤ 5%

- Construct concentration curves for synthetic RNA samples with & without MNase treatment

Equipment & supplies needed

Styrofoam container filled with crushed ice (for RNA samples)

1.5 mL nuclease-free black microcentrifuge tube (Argos Technologies #T7100BK)

1.5 mL nuclease-free polypropylene microcentrifuge tubes

0.2 mL nuclease-free polypropylene PCR tubes (Molecular BioProducts #3412)

Micropipettes & aerosol barrier pipette tips

Tube racks for 0.2 mL & 1.5 mL tubes

Lab vortexer

Lab timer

25°C dry heat block

NanoDrop ND-1000 (sign up for the machine before running the experiment)

Chemicals & reagents

RNaseZap! or similar RNase decontamination solution (Ambion #AM9780/9782)

20 µM PNA mix in 10X micrococcal nuclease reaction buffer (10X MRB) (recipe below)

250 µM synthetic RNA

Molecular biology-grade water (MB-grade water)

Micrococcal nuclease (New England BioLabs #M0247S)

100 Kunitz units/µL micrococcal nuclease dilution buffer (MDB) (recipe below)

200 µM DiSC2(5) in methanol (Sigma-Aldrich #173754) (recipe below)

Methanol

Make sure to wipe EVERYTHING down with nuclease decontamination solution before AND during the procedure

Replace gloves frequently, especially after using the NanoDrop

After the experiment, use RNaseZap! to clean everything that is going back into the RNase-free cabinet

Procedure

I. Prepare reagents

- Take out synthetic RNA aliquot (250 µM stock concentration) from -80°C freezer, thaw in a tube rack on ice.

- Make RNA dilutions in MB-grade water, keep on ice for the duration of the experiment.

  - Make 1 mL 1.33 µM RNA working stock solution from 250 µM stock:

    - Take out RNase-free 1.5 mL polypropylene tube & label.

    - Fill tube with 1 mL MB-grade water.

    - Use a 10 or 20 µL pipette to remove 5.33 µL water from the tube.

    - Use the same pipette to add 5.33 µL of the 250 µM RNA stock into the tube.

    - Mix the tube by vortexing and store on ice until the end of the experiment.

  - Make 150 µL of specified RNA dilutions in RNase-free 0.2 mL tubes:

| Final (20 µL) concentration (nM) | Actual (15 µL) concentration (nM) | 1.33 µL RNA stock (µL) | MB-grade water (µL) |
| --- | --- | --- | --- |
| 0 | 0 | 0 | 150 |
| 100 | 133 | 15 | 135 |
| 250 | 333 | 37.5 | 112.5 |
| 500 | 667 | 75 | 75 |
| 750 | 1000 | 112.5 | 37.5 |
| 1000 | 1333 | 150 | 0 |

Keep samples in a tube rack on ice for the entire experiment.

- Make 20 µL 100 Ku/µL micrococcal nuclease:

  - Take out a 0.2 mL RNase-free tube & label.

  - Add 10 µL MB-grade water to tube.

  - Add 10 µL 200 Ku/µL micrococcal nuclease to tube & mix by vortexing.

MNase stock is viscous, pipette carefully and check tip to see that it has the correct volume

  - Keep tube on ice.

- Aliquot 20 µL of the micrococcal nuclease dilution buffer (MDB) into a 0.2 mL PCR tube, keep on ice.

- Make 40 µL of the BLANK solution for the NanoDrop:

  - Take out and label a 0.2 mL PCR tube.

  - Add 30 µL MB-grade water to tube.

  - Add 6 µL 10X blank mix.

  - Add 2 µL methanol and mix by vortexing.

II. Assay concentration curve experiment

Each experiment set consists of all the RNA concentrations (0, 100, 250, 500, 750, 1000 nM) with OR without MNase.  This means that 6 samples are read on the NanoDrop per set.  DO NOT exceed 6 samples per set because timing is crucial for these samples.  It generally takes about 1 minute (with practice) to pipette 1 µL of either the MNase/MDB or dye for all 6 samples, so there is no need to stop the timer at the 5 minute point to add MNase/MDB since at the 10 minute mark dye will be pipetted into the first tube of the set.

- Set out and label 6 RNase-free 0.2 mL PCR tubes into a tube rack.

- Add 3 µL PNA mix to each tube.

- Add 15 µL of MB-grade water (concentration 0) into the first tube.  Mix by GENTLE vortexing and start the timer.  Place tube into 25°C dry heat block.

- Repeat step 3 for EACH remaining concentration.  DO NOT add the RNA to all tubes and then start mixing/timing.

- AT the 5 minute mark, add 1 µL of either MDB or MNase (depending on experiment set) to the first tube.  Make sure that all of the MDB (MNase) has been expelled from the pipette tip.  Mix the tube gently by stirring with the pipette tip and/or by vortexing and return to the heat block.

*Start pipetting AT the 5 minute mark: this means make sure that your MDB/MNase tube is already vortexed and open in front of you, and that you have your pipette with a clean tip READY at 4:55.

- Repeat step 5 for each remaining tube.

- AT the 10 minute mark, add 1 µL of the 200 µM dye solution to the first tube.  Make sure to expel all of the dye into the tube with ONE push.  Mix the tube with gentle vortexing—DO NOT pipette to mix.  Transfer tube into a plastic tube rack.

Timing is crucial at this step since the MNase will continue to digest RNA if not stopped with the dye solution AT the 10 minute mark.  Again, this means that the dye tube and your pipette have to be ready at 9:55.

- Repeat step 7 for each remaining tube.

Don’t forget to close the dye tube!  The dye photobleaches.

- Take the samples (and the BLANK solution) to the NanoDrop immediately after dye has been added to the last tube.

Use a different pipette and a different box of pipette tips to use with the NanoDrop, and do not bring them back into the RNase-free area

- On the NanoDrop, use the UV-Vis program, and blank the machine using the BLANK solution (NOT water).

- Make sure that Max absorbance is 0.1 and that Hi Abs is off (unchecked).

- Record each sample sequentially on the machine.

Change gloves!!

- Repeat from step 1 for each experiment set (6 sets for each QC experiment: triplicate concentrations with or without MNase).

- Save & send the NanoDrop UV-Vis file:

C:\NanoDrop Data\Default\UV-Vis\UV-Vis YYYY MM DD v3.8

RECIPES

For each 20 µL reaction, the concentrations are:

| Component | Volume (µL) | Final concentration |
| --- | --- | --- |
| 20 µM PNA mix | 3 | 1 µM PNA, 1X MRB |
| RNA sample | 15 | Various |
| 100 Ku/µL MNase (MDB) | 1 | 100 Ku |
| 200 µM DiSC2(5) | 1 | 10 µM |

BLANK solution for NanoDrop:

| Component | Volume (µL) | Final concentration |
| --- | --- | --- |
| 10X MRB | 2 | 1X MRB |
| MB-grade water | 16 | NA |
| 100 Ku/µL MDB | 1 | NA |
| Methanol | 1 | NA |

20 µM PNA mix in 10X micrococcal nuclease buffer

For 300 µL, mix in 1.5 mL polypropylene nuclease-free microcentrifuge tube:

8 µL		250 µM PNA stock solution

92 µL		molecular biology-grade water

200 µL		10X micrococcal nuclease reaction buffer (New England BioLabs #M0247S)

100 Kunitz units/µL micrococcal nuclease dilution buffer (MDB)

* Used to account for the buffer that micrococcal nuclease is stored in

For 1 mL, mix in 1.5 mL polypropylene nuclease-free microcentrifuge tube:

734 µL		molecular biology-grade water

250 µL		glycerol

10 µL		1M Tris pH 8 (RNase-free, from Ambion)

5 µL		5M NaCl solution (RNase-free, from Ambion)

1 µL		0.5M EDTA pH 8 (RNase-free, from Ambion)

200 µM DiSC2(5)

For 500 µL, mix in 1.5 mL BLACK (light-safe) polypropylene nuclease-free microcentrifuge tube:

450 µL		methanol

50 µL		2 mM DiSC2(5) stock solution

# BIOGRAPHY OF THE AUTHOR

Janice Duy was born in Pennsylvania and was raised in the Philippines.  She graduated from St. Paul College, Pasig in 2000 and attended the University of the Philippines-Diliman.  When her family moved to Maine, she matriculated at the University of Maine and earned her Bachelor of Science degree in Electrical Engineering in 2005.  She continued her studies in the Functional Genomics doctoral program, now a track of the Graduate School of Biomedical Sciences and Engineering.  Janice is a candidate for the Doctor of Philosophy degree Interdisciplinary in Functional Genomics from the University of Maine in May 2013.