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Test Date:  8/11/06
Item Description :  Titanium housing for Peter Girguis
Item S/N:  PG 1 stamped on housing and cap
Tested For: MBARI at our own Pressure Test Facilities
Test Pressure:  6700 psi
 6700  psi / 1.5 S.F = 4470 psi     4470/1.487 psi/meter S.W. = 3000meters water depth.
Gauge:  Autoclave Pressure Transducer
Gauge S/N:  95104916
Gauge Range:  0-15000 psi

Calibration Date:  6/15/06 with a PSI-Tronixs digital gage calibrated by 


Precision Measurements  on 3/2/05,  Report # 1138295

The test articles were subjected to the implodable items alternative procedure stipulating 10 cycles to maximum allowed pressure.  A 1 minute hold at 0 and 6700 psi on cycles 
1-9 and a 60 minute hold at 6700 psi on the 10th cycle. 
 No evidence of leaks or damage was noted.  Test trace attached.
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